## Introduction
Appendix E: Checklists for Implementation of the SWS Profile
The checklist steps are arranged in approximate chronological order rather than by actor or task to make them more usable in the clinic.
The checklists are in excel to make data analysis and retrieval easier and to allow convenient conformance scoring
The final TAB contains a checklist to be used by a site doing self attestation of conformance or by an auditor to check conformance. It allows for scoring of conformance. A passing score has yet to be determined.
In the future, on-line checklist forms (such as in google docs) may be created for user convenience
To start move to the next TAB "Pre Acquisition", then to "Subjects and Data Acquisition" etc.

## Pre Acquisition
QIBA SHEAR WAVE SPEED PROFILE EXECUTION CHECKLIST V5.0 Scanner QA Pre & Post Installation
ITEM# | PARAMETER | REQUIREMENT | RESPONSIBILITY/ACTOR | COMPLETED BY [ 1= Radiologist; 2 = Technologist; 3 = Sonographer; 4= MFR; 5 = QA Manager; 6 = US Scanner ] | CONFORMS? [YES, NO, N/A] | NOTES/EXPLANATIONS
3 | Site Conformance Check
3.0.2.1 | Ultrasound Scanner | Confirm all US scanners conform to Profile | QA Manager
3.0.2.2 | MFR | Confirm manufacturer responsibilities for equipment performance and installation are met | QA Manager
3.0.2.3 | Technologist / Sonographer | Confirm each technologist / sonographer conforms to Profile with training, documented acquistion performance & proper SWS acquistions | QA Manager
3.0.2.4 | Radiologist | Confirm all Radiologists conform to Profile with patient interaction, acquistion performance, and reporting | QA Manager
3.1 | Staff Qualification
3.1.2.1 | Operator Training | Technologist trained and approved for SWS acquisition | Technologist / Sonographer or Radiologist
3.1.2.2 | Operator Qualification | Meets performance requirements on phantoms and subjects | Technologist / Sonographer or Radiologist |  |  | Phantom testing-wCV ≤ .05 and/or case review IQR/median ≤ 0.30
3.3 | Pre-delivery
3.3.2.1 | Acoustic Output (SWS Mode) | Manufacturer certifies maximum acoustic output levels meets FDA recommendation while operating SWS mode | MFR |  |  | Manufacturer specification and certification
3.3.2.2 | Acoustic Transmit Focusing | Manufacturer specifies and certifies SWS measurement and Imaging | MFR
3.3.2.3 | SWS Measurement Consistency | Manufacturer confirms SWS Measurement Consistency of the Ultrasound Scanner is within +/- 5% | MFR |  |  | See 4.2 Assessment Procedure: SWS Measurement Consistency
3.3.2.4 | Ultrasound Imaging Performance | Meets MFR specifications as published in scanner documentation | MFR |  |  | See 4.1 Assessment Procedure: Imaging Performance
3.3.2.5 | SWS Imaging Performance | Identification and display meets manufacturer specifications listed in Appendix D | MFR
3.3.2.6 | Software verification | Software version equals version specified in QIBA profile (Appendix D) | MFR
3.3.2.7 | Hardware & transducer MFR specified parameters | Ensure the equipment intended for use is listed in Appendix D as a compliant combination of System, Software Revision and Transducer. | MFR |  |  | Required Scanner components are present and found in Appendix D; scanner specific instructions
3.4 | Installation
3.4.2.1 | Hardware Damage | No physical damage | MFR
3.4.2.2 | Software verification | S/W version equals the version specified in products QIBA Conformance Statement or one listed in Appendix D | QA Manager
3.4.2.3 | SWS Measurement Concordance | Confirm SWS measurements are within +/- 5% | QA Manager |  |  | See 4.3 Assessment Procedure: SWS Measurement Concordance
3.5 | Periodic Quality Assurance (QA)
3.5.2.1 | US Imaging QA | U/S system QA checks and conform to quality criteria specified in AIUM guidelines | QA manager |  |  | On arrival then annually unless potential problem found during operations
3.5.2.2 | SWS Measurement Consistency & System QA Testing | Confirms that measurements of SWS on a QIBA elastic phantom using standard instrument settings and acquisition procedures annually, and after any software change are within ± 5% of the values of the Elastic SWS phantom specifications as determined by testing with a Verasonics system. If the system is already known to give results more than 5% different from the phantom values, the system should give values within ± 5% of the previously obtained results. (See Measurement Concordance Test Procedure Profile Section 3.4.1) | QA manager |  |  | Annual testing and testing after any software/hardware change.
3.5.2.3 | SWS Measurement Consistency & System QA Testing | US Scanner shall be capable of performing SWS measurements at reproducible instrument settings using manufacture specific standard procedures [appendix D] | US Scanner |  |  | Refere to Appendix D
3.5.2.3 | Operator Training & Qualification Testing | Confirm each operator trained on patient workflow & SWS acquistion and evaluated qualification criteria are met (the requirements are in 3.1. Staff Qualification) | QA manager |  |  | Training on initial operator selection
3.5.2.4 | US Imaging / SWS Phantom Characterization & Stability Testing | Shall confirm SWS Phantom Acoustic and Mechanical Properties at Independent Test Site using QIBA procedures after construction and if a weight change of >0.5% has occurred | QA Manager
Note on C15: On screen dialog message appears as manufacturer indicates i.e. "Failure to estimate a shear wave velocity is indicated by an "x.xx m/s" [GE]
Note on A28: Garra, Brian: Consider moving all QA activities to a separate tab since the QA activities do not fit in termporal order with the other components.

## Subjects & Data Acquisition
QIBA SHEAR WAVE SPEED PROFILE EXECUTION CHECKLIST V5.0 Patient & Acquisition
ITEM# | PARAMETER | REQUIREMENT | RESPONSIBILITY/ACTOR | COMPLETED BY [ 1= Radiologist; 2 = Technologist; 3 = Sonographer; ] | COMPLETED? [YES, NO, N/A] | NOTES/EXPLANATIONS
3.7 | Subject Selection
3.7.2.1 | Clinical Indication | Assess liver stiffness for liver pathology that may lead to increased organ stiffness and increased shear wave speed (for example liver fibrosis). | Radiologist |  |  | A valid research protocol or a clinical concern supported by the literature is needed
3.7.2.2 | Approach | Confirm intercostal approach is feasible | Radiologist or Technologist / Sonographer
3.7.2.3 | Body Wall Thickness | Confirm total body wall thickness 4cm or less at intercostal location | Radiologist or Technologist / Sonographer |  |  | Allows SWS measurement to be made at a maximum depth of 6.5 cm from the skin surface and at least 2 cm below the liver capsule.
3.7.2.4 | Intercostal Space | Intercostal spacing wide enough for transducer placement | Radiologist or Technologist / Sonographer |  |  | Evaluate intercostal space by manual palpation
3.7.2.5 | Breathing | Confirm patients ability and understanding breath hold instructions | Radiologist or Technologist / Sonographer |  |  | Refer to appendix D for manufacturers acquisition time
3.7.2.6 | Prior Surgery | Right lobe of liver present, no scarring or shadowing from surgery | Radiologist or Technologist / Sonographer |  |  | Confirm subject has right lobe of liver if history of previous hepatic surgery
3.7.2.7 | Informed Consent | Informed consent obtained | Radiologist or Technologist / Sonographer |  |  | IC and HIPAA forms per institutional policy (sample consent form language in Appendix G)
3.8 | Subject Handling
3.8.2.1 | Patient Instructions | Instructs patient to fast 4 hours prior to procedure | Radiologist or Technologist / Sonographer |  |  | Instruct patient on how to accomplish fasting and exceptions (e.g. oral medications, insulin); instruction sheet example Appendix G
3.8.2.2 | Fasting State | Query the patient whether successfully fasted 4 hours or not | Radiologist or Technologist / Sonographer |  |  | Offer to acquire data on a later date or later in the day if patient is not in a fasting state.
3.8.2.3 | Informed Consent | Informed consent confirmed | Radiologist or Technologist / Sonographer |  |  | IC and HIPAA forms per institutional policy. Sample consent form language in Appendix G.
3.8.2.4 | Patient Information | Provide information on SWE, specifically the acquisition process, # of breath holds, transducer pressure | Radiologist or Technologist / Sonographer |  |  | Refer to patient information-instructions sheet
3.9 | SWS Image Acquisition (SWEI) and Point SWS Measurement
3.9.2.1 | SWS Acquisition System | Measurements are performed on the same equipment/manufacturer when successive measurements are made on the same patient | Radiologist or Technologist / Sonographer |  |  | Obtain calibration values for successive measurements made on various machines for the same patient and forward along with test results for use during intepretation
3.9.2.2 | Patient Position | Patient positioned in supine or 30 degree left lateral decubitus | Radiologist or Technologist / Sonographer |  |  | Not full left lat decubitus which has been shown to produce a small but statistically significant increase in SWS
3.9.2.3 | Respiration | Patient has successfully practiced suspended tidal (normal or quiet) respiration and it is used during acquisition | Radiologist or Technologist / Sonographer |  |  | No other liver movement is observed during acquistion process
3.9.2.4 | Transducer Position | Position transducer at intercostal space wide enough to accommodate transducer and correct level to image/acquire from mid to upper right liver lobe Position transducer face long axis parallel to the intercostal space and check for correct positioning by inspecting the image for shadowing at the image edges. Position the transducer face in contact with the skin and parallel to the liver capsule so that the acoustic waves travel perpendicular to the capsule. | Radiologist or Technologist / Sonographer |  |  | Transducer positioned at intercostal space from the upper right liver lobe
3.9.2.5 | Transducer Pressure | Maintain contact with skin surface; lighter pressure during SWS acquisition | Radiologist or Technologist / Sonographer |  |  | May use more pressure to compress body wall when needed to enable ROI to be positioned in proper position in Liver
3.9.2.6 | Ultrasound Image Location Confirmation | Confirm absence of focal structures near image center, no acoustic shadowing from the ribs | Radiologist or Technologist / Sonographer |  |  | Keep the liver capsule parallel to the transducer face in both planes
3.9.2.7 | Measurement Region of Interest (ROI) Placement | ROI center positioned 2 cm deep to liver; <6.5 cm from transducer face | Radiologist or Technologist / Sonographer |  |  | Shall use the standard ROI size specified by the ultrasound vendor as the default for their system; for follow up study use the same depth as previous study
3.9.2.8 | Measurement Region of Interest (ROI) Placement- Position | ROI positioned near center of image, away from right/left margins and away from discrete structures | Radiologist or Technologist / Sonographer |  |  | Discrete structures refer to structures such as: such as liver margin, nodules, portal triads or hepatic veins
3.9.2.9 | Region of Interest (ROI) Placement- Size | ROI box diameter is 10mm or 6mm X 10mm | Radiologist or Technologist / Sonographer |  |  | See Appendix D. Refer to MFR specifications. If a choice is given choose an ROI equal to or larger than 10mm
3.9.2.10 | Follow-up Consistency | ROI placed consistent with baseline measurement depth for all follow up acquisitions | Radiologist or Technologist / Sonographer |  |  | Transducer position and ROI placement consistent for all measurements
3.9.2.11 | Number of Measurements | Number of measurements specified by Manufacturer (see Appendix D) or if not specified, 10 measurements. | Radiologist or Technologist / Sonographer |  |  | Please refer to manufacturer's specific instructions (Appendix D)
3.9.2.12 | Liver Movement | Acquire when no visible liver motion | Radiologist or Technologist / Sonographer |  |  | Acquistion repeated if breathing or motion observed
3.9.2.13 | SWS Imaging - Color Map | Consistent between exams and patients | Radiologist or Technologist / Sonographer |  |  | Use institutional and/or national standards. See manufacturer guidelines
3.9.2.14 | SWS Imaging - Color Transparency | Set to visualize color changes and grayscale anatomy | Radiologist or Technologist / Sonographer |  |  | Refer to manufacturer guidelines - Appendix D
3.9.2.15 | SWS Imaging - Frame Averaging | Set according to preference after manufacturer recommendation | Radiologist or Technologist / Sonographer |  |  | After the initial setting according to manufacturer recommendations
3.9.2.16 | SWS Imaging - Frame Rate/Color Box Size | Set box as large as possible for adequate color frame rate | Radiologist or Technologist / Sonographer |  |  | Refer to manufacturer guidelines - Appendix D
3.9.2.17 | SWS Imaging - Point Measurement ROI location | ROI located in most homogeneous region of SWS color map (Appendix D) | Radiologist or Technologist / Sonographer |  |  | Refer to manufacturer guidelines - Appendix D
3.9.2.18 | SWS Imaging - Point Measurement ROI size | Set per manufacturer specifications; minimum ROI size meets default 6mm X 10mm or diameter 10mm | Radiologist or Technologist / Sonographer |  |  | Refer to manufacturer guidelines - Appendix D
3.9.2.19 | SWS Imaging - Point Measurement Data Transfer | Successful Image and Data Transfer to PACS or other storage and confirmation | Radiologist or Technologist / Sonographer |  |  | Shall transfer SWS measurement objects to PACS or other storage and confirm successful storage.
Note on C4: possible diseases include Hepatitis C; Hepatitis B; NASH, ETOH
Note on C6: to allow positioning of ROI 2cm deep to liver capsule based on BMI
Note on C7: Wide intercostal space required. Evaluate intercostal space by manual palpation
Note on C8: Refer to appendix D for manufacturers acquisition time KA: times not listed
Note on C9: Confirm subject has right lobe of liver if positive for previous surgery
Note on C10: ICF and HIPAA forms per institutional policy
Note on C14: Query patient on fasting. Offer to acquire data on a later date or later in the day if patient is not in a fasting state.
Note on C15: ICF and HIPAA forms per institutional policy
Note on C22: Space must be wide enough to accommodate transducer & image upper Rt lobe (segments 5, 7, or 8)
Note on C24: • Clarity of visualization of normal liver parenchymal structures such as portal triads and hepatic veins on grayscale image. • No acoustic shadowing from the ribs. • Absence of grayscale artifacts.
Note on C25: if follow up study use the previous depth.
Note on C26: Away from discrete structures such as liver margin, nodules, portal triads or hepatic veins
Note on C27: Refer to Appendix D for MFR specifications. The standard ROI default size should conform to 6mm x 10mm or 10mm Diameter
Note on C28: Place ROI at constant depth for all acquisitions, and follow up acquisitions
Note on C29: MFR may specify more than 10 images. Please refer to vendor specific instructions (Appendix D)
Note on C30: Shall acquire only when there is no visible liver motion

## Quality Assurance
QIBA SHEAR WAVE SPEED PROFILE EXECUTION CHECKLIST V5.0 Processing & Analysis
ITEM# | PARAMETER | REQUIREMENT | RESPONSIBILITY/ACTOR | COMPLETED BY [ 1= Radiologist; 2 = Technologist; 3 = Sonographer; 4 = MFR; ] | COMPLETED? [YES, NO, N/A] | NOTES/EXPLANATIONS
3.11 | Image Quality Assurance (QA)
3.11.2.1 | Suboptimal SWS Acquistion | SWS estimates excluded if suboptimal due to motion or other factors | Radiologist or Technologist / Sonographer |  |  | See section 3.9.2 for acquisition-related exclusion criteria.
3.11.2.2 | User training on image display | Provide user training on image interpretation and Operator training on optimal placement of measurement ROI. | MFR
Note on C4: Exclude SWS numbers that were acquired suboptimally either by observation during acquisition, inspection of saved images,or pre/post liver images not matching. • liver movement during acquisition, • patient talking during acquisition, • transducer slippage during acquisition and • inadvertent shift of ROI to a deeper or shallower depth, • placement of the ROI near to a vessel or other discrete structure

## Profile Conformance-by section
| QIBA SHEAR WAVE SPEED PROFILE CHECKLIST V5.0 Conformance-by Section |  |  | Instructions: For each section, first fill in the answer (y or n) for each task. Then fill in the conformance score for the section in the next column. The overall site score as a percent of the maximum possible will then be calculated automatically.
SECTION CONFORMANCE CHECKLIST (enter: 1 if self attestation, 2 if independent review):
ITEM# | PARAMETER | Conforms Y/N | SECTION CONFORMANCE SCORE Full Conformance = 3 pts; Non-conformant one task = 2 pts; Non-conformant 2-3 tasks = 1 pts; Non-conformant 3+ OR ALL tasks = 0 pts | NOTES/EXPLANATIONS
3.1 | Staff Qualification
3.1.2.1 | Operator Training
3.1.2.2 | Operator Qualification
3.3 | Pre-delivery
3.3.2.1 | Acoustic Output (SWS Mode)
3.3.2.2 | Acoustic Transmit Focusing
3.3.2.3 | SWS Measurement Consistency
3.3.2.4 | Ultrasound Imaging Performance
3.3.2.5 | SWS Imaging Performance
3.3.2.6 | Software verification
3.3.2.7 | Hardware & transducer MFR specified parameters
3.4 | Installation
3.4.2.1 | Hardware Damage Check
3.4.2.2 | Software verification
3.4.2.3 | SWS Measurement Concordance
3.5 | Periodic Quality Assurance (QA)
3.5.2.1 | US Imaging QA
3.5.2.2 | SWS Measurement Consistency & System QA Testing Using SWS Phantom
3.5.2.3 | Operator training and qualification testing
3.5.2.4 | US Imaging and SWS Phantom Characterization and Stability Testing
3.7 | Subject Selection
3.7.2.1 | Clinical Indication
3.7.2.2 | Approach
3.7.2.3 | Body Wall Thickness
3.7.2.4 | Intercostal Space
3.7.2.5 | Breathing
3.7.2.6 | Prior Surgery
3.7.2.7 | Informed Consent
3.8 | Subject Handling
3.8.2.1 | Patient Instructions
3.8.2.2 | Fasting State
3.8.2.3 | Informed Consent
3.8.2.4 | Patient Information
3.9 | SWS Image Acquistion and Point SWS Measurement
3.9.2.1 | SWS Measurement Acquisition Device
3.9.2.2 | Patient Position
3.9.2.3 | Respiration
3.9.2.4 | Transducer Position
3.9.2.5 | Transducer Pressure
3.9.2.6 | Ultrasound Image Location Confirmation
3.9.2.7 | Measurement Region of Interest (ROI) Placement
3.9.2.8 | Follow-up Consistency
3.9.2.9 | Number of Measurements
3.9.2.10 | Liver Movement
3.9.2.11 | SWS Imaging-Color Map
3.9.2.12 | SWS Imaging-Color Transparency
3.9.2.13 | SWS Imaging-Frame Averaging
3.9.2.14 | SWS Imaging-Frame Rate/Color Box Size
3.9.2.15 | SWS Imaging-Point Measurement ROI location
3.9.2.16 | SWS Imaging-Point Measurement ROI size
3.9.2.17 | SWS Imaging-Point Measurement Data Transfer
3.11 | Image QA
3.11.2.1 | Suboptimal SWS Acquistion handling
3.11.2.2 | User training on image display
4.1 | Assessment Procedure: Imaging Performance
4.1.1 | Assessment: - Site has assessed imaging performance according to procedure 4.1.2 (scoring: yes = 3, other imaging assessment = 2, no assessment = 1)
4.1.2 | Corrective Action Taken if Performance Issue Identified (scoring: yes successful = 3, yes partially successful = 2, yes unsuccessful = 1, no corrective action)
4.1.3 | Assessment: SWS Performance - Site has assessed ACTOR performance according to ACTOR assessment procedure 4.1.2 (scoring: yes, periodic = 3, yes, once = 2, other assessment performed = 1, no assessment = 0)
4.1.4 | Corrective Action Taken if Performance Issue Identified (scoring: yes- successful, assessed actors as in 4.2.3 = 3, yes - successful, after other assessment = 2, yes - partially successful = 1, no corrective action = 0)
| CONFORMANCE SCORE (% OF MAX): |  | 0

## Profile Conformance by Actor
| QIBA SHEAR WAVE SPEED PROFILE CHECKLIST V5.0 Conformance by Actor |  |  |  | Instructions: For each Actor, first fill in the answer (y or n) for each task. Then fill in the conformance score for the actor in the next column. The overall site score as a percent of the maximum possible will then be calculated automatically.
ACTOR CONFORMANCE CHECKLIST (enter: 1 if self attestation, 2 if independent review):
Item # | ACTOR | ACTIVITY / PARAMETER | Conforms Y/N | ACTOR CONFORMANCE SCORE Full Conformance = 3 pts; Non-conformant one task = 2 pts; Non-conformant 2-3 tasks = 1 pt; Non-conformant 3+ OR ALL tasks= 0 pts | NOTES / EXPLANATIONS
3.5 | Ultrasound Scanner | Periodic Quality Assurance
3.5.2.1 | Scanner | SWS Measurement Consistency & System QA Testing Using SWS Phantom
3.3 | Manufacturer | Pre-Delivery
3.3.2.1 | Mfr | Acoustic Output (SWS Mode)
3.3.2.2 | Mfr | Acoustic Transmit Focusing
3.3.2.3 | Mfr | SWS Measurement Consistency
3.3.2.4 | Mfr | US Imaging Performance
3.3.2.5 | Mfr | SWS Imaging Performance
3.3.2.6 | Mfr | Software verification
3.3.2.7 | Mfr | Hardware and transducer
3.4 |  | Installation
3.4.2.1 | Mfr | Hardware Damage
3.1 | Radiologist | Site Qualitification
3.1.2.1 | Rad | Operator Training
3.1.2.2 | Rad | Operator Qualification
3.7 |  | Subject Selection
3.7.2.1 | Rad | Clinical Indication
3.7.2.2 | Rad | Approach
3.7.2.3 | Rad | Body Wall Thickness
3.7.2.4 | Rad | Intercostal space
3.7.2.5 | Rad | Breathing
3.7.2.6 | Rad | Prior Surgery
3.8 | Rad | Subject Handling
3.8.2.1 | Rad | Patient Instructions
3.8.2.2 | Rad | Fasting State
3.8.2.3 | Rad | Informed Consent
3.8.2.4 | Rad | Patient Information
3.9 |  | Image Data Acquisition
3.9.2.1 | Rad | SWS Acquisition
3.9.2.2 | Rad | Patient Position
3.9.2.3 | Rad | Respiration
3.9.2.4 | Rad | Transducer Position
3.9.2.5 | Rad | Transducer Pressure
3.9.2.6 | Rad | Ultrasound Image – location confirmation
3.9.2.7 | Rad | Measurement Region of Interest (ROI) Placement
3.9.2.8 | Rad | Follow-up Consistency
3.9.2.9 | Rad | Number of Measurements
3.9.2.10 | Rad | Liver Movement
3.9.2.11 | Rad | SWS Imaging Color Map
3.9.2.12 | Rad | SWS Imaging Color Transparency
3.9.2.13 | Rad | SWS Imaging Frame Averaging
3.9.2.14 | Rad | SWS Imaging Frame Rate/Color Box Size
3.9.2.15 | Rad | SWS Imaging Point Measurement ROI location
3.9.2.16 | Rad | SWS Imaging Point Measurement ROI size
3.9.2.17 | Rad | SWS Imaging Point Measurement Data Transfer
3.11 |  | Image QA
3.11.2.1 | Rad | Suboptimal SWS Acquisition
3.1 | Technologist / Sonographer | Site Qualification
3.1.2.1 | Sono Tech | Operator Training
3.1.2.2 | Sono Tech | Operator Qualification
3.7 |  | Subject Selection
3.7.2.2 | Sono Tech | Approach
3.7.2.3 | Sono Tech | Body Wall Thickness
3.7.2.4 | Sono Tech | Intercostal space
3.7.2.5 | Sono Tech | Breathing
3.7.2.6 | Sono Tech | Prior Surgery
3.8 |  | Subject Handling
3.8.2.1 | Sono Tech | Patient Instructions
3.8.2.2 | Sono Tech | Fasting State
3.8.2.3 | Sono Tech | Informed Consent
3.8.2.4 | Sono Tech | Patient Information
3.9 |  | Image Data Acquisition
3.9.2.1 | Sono Tech | SWS Acquisition
3.9.2.2 | Sono Tech | Patient Position
3.9.2.3 | Sono Tech | Respiration
3.9.2.4 | Sono Tech | Transducer Position
3.9.2.5 | Sono Tech | Transducer Pressure
3.9.2.6 | Sono Tech | Ultrasound Image – location confirmation
3.9.2.7 | Sono Tech | Measurement Region of Interest (ROI) Placement
3.9.2.8 | Sono Tech | Follow-up Consistency
3.9.2.9 | Sono Tech | Number of Measurements
3.9.2.10 | Sono Tech | Liver Movement
3.9.2.11 | Sono Tech | SWS Imaging Color Map
3.9.2.12 | Sono Tech | SWS Imaging Color Transparency
3.9.2.13 | Sono Tech | SWS Imaging Frame Averaging
3.9.2.14 | Sono Tech | SWS Imaging Frame Rate/Color Box Size
3.9.2.15 | Sono Tech | SWS Imaging Point Measurement ROI location
3.9.2.16 | Sono Tech | SWS Imaging Point Measurement ROI size
3.9.2.17 | Sono Tech | SWS Imaging Point Measurement Data Transfer
3.11 |  | Image QA
3.11.2.1 | Sono Tech | Suboptimal SWS Acquisition
3 | QA Manager | Site Conformance
| QA Mgr | Ultrasound Scanner Conformance
| QA Mgr | Manufacturer Conformance
| QA Mgr | Technologist/Sonographer Conformance
| QA Mgr | Radiologists Conformance
3.4 |  | Installation
| QA Mgr | Software Verification
| QA Mgr | SWS Measurement Concordance
3.5 |  | Periodic Quality Assurance (QA)
| QA Mgr | US Imaging QA
| QA Mgr | SWS Measurement Consistency & System QA Testing Using SWS Phantom
| QA Mgr | Operator training and qualtification testing
| QA Mgr | US Imaging & SWS Phantom Characterization & Stability Testing
ALL | SITE | SITE CONFORMANCE SCORE (% OF MAX) |  | 0