## Results and Resolutions
Commenter | Section | Line # | Priority | Issue | Proposal | Committee Discussion | Resolution (w Rationale if rejected) | Status
[8 empty] | 0 | TBD | To be decided
See http://qibawiki.rsna.org/index.php/Public_Comment_Process for more guidance on the comment resolution process | [7 empty] | 0 | OK | No action requested
[8 empty] | 0 | Discuss | Need to decide resolution
[8 empty] | 0 | Done | Profile update completed
(Esaote) Gian Matteo Morano | Appendix D | 1760 | M | Esaote (ultrasound manufacturer) would like to contribute to model specific insturctions and parameters (Appendix D); a model draft of 4 machines included. | Model specific instructions for Appendix D are included for consideration |  | Addition made
Riwa Kishimoto, MD, PhD | 2. Clinical Context & Claims | 215 | L | Clinical Context and Claims, is inter-operator variability (different technologists with the same scanner at the same point) regarded as claim 4 (longitudinal claim)? Our previous study showed a similar 95% limit of agreement for inter-operator reproducibility with Table 2-4, such as -21.3~23.8% for 4cm, -29.0~25.6% for 6cm and -31.7~30.8% for 8cm depth of right lobe of the liver in healthy volunteers (SWS:1.06-1.13m/s). We thought our 95% LOA were relatively large, but I think our results are consistent with this claim. Article Referenced: Kishimoto R, et al Eur Radiol. 2019 Nov;29(11):5999-6008. doi: 10.1007/s00330-019-06195-8. Intra- and inter-operator reproducibility of US point shear-wave elastography in various organs: evaluation in phantoms and healthy volunteers. | PROPOSE HAVING NANCY OBUCHOWSKI REVIEW AND COMMENT ON THIS. |  | The QIBA statistitian has reviewed your comment and provides the following response: "Yes, yes, either different technologist or different scanner or both (but from the same site)." The committee agrees that the interoperator variability is part of claim 4 as is inter-scanner variability. The committe agrees that the results you report as being in your article are consistent with claim 4. (no change was made to the profile as this was a comment, not a proposed change).
Paul Barbone | N/A |  | L | Recommend center frequency of SW pulse (and bandwidth) be reported with each SWS measurement. Individual platforms may have different center freq ie. 100 Hz, 300 Hz, it is useful to know if measured at the same center freq or not. | none listed | NOT SURE HOW TO IMPLEMENT IT WOULD GO INTO A SECTION ON “REPORTING RESULTS” I THINK BUT WE DO NOT HAVE SUCH A SECTION—FOR DISCUSSION | Committee response: The manufacturers currently do not provide this information and provision of this information may require submission to the FDA. We will discuss this further with mfrs for version 2.
Karen Alton | Appendix D | 1265 | L | "See Specifications in Profile Section 3.6, 3.8, 3.10" | Section 3.10 was removed. Propose change to "…Profile Sections 3.7, 3.8" |  | Correction made
(Canon) Todd Erpelding, PhD | Appendix D | 1250 | M | Canon “Model-Specific Instructions” are not up-to-date with the latest version. | Update Canon system and transducer information to include our latest releases (Aplio 450/550 and some new transducers). Addition line 1250: "Aplio a-series (a450/a550/a)" |  | Addition made
| Appendix D | 1256 | M | Add new software version | Addition line 1256: "Aplio a450/a550/a V1.0 or later" |  | Addition made
| Appendix D | 1260 | M | Add 1 machine and 2 transducers to the Canon Transducer Chart | Additon line 1260: Column 3 out of 5 "Aplio a450/a550/a", Row 3 out of 8: "PVI-574BX", Row 7 out of 8 "PVT-574BT" |  | Addition made
Andy Milkowski | Appendix D | 1563 | M | New Equipment | Addition line 1564: "ACUSON Redwood" |  | Addition made
| Appendix D | 1563 | M | New Equipment | Addition line 1565: "ACUSON Juniper" |  | Addition made
| Appendix D | 1568 | M | New Software | Addition line 1568: "ACUSON Redwood: VA10A or later" |  | Addition made
| Appendix D | 1568 | M | New Software | Addition line 1569: "ACUSON Juniper: VA10A or later" |  | Addition made
| 3.1 Pre-delivery | 377-378 | M | "Need additional clarity regarding frequency of pre-delivery testing" | "Clarify that manufacturer needs to perform this testing and document." |  | Clarifying sentence added
(No Name) collegue of Riwa Kishimoto, MD, PhD | Closed Issues | 105 | L | "There are many places without space between the number and the unit." "Page 7, Q. Give stiffness in m/sec or kPa? A. M/sec -> m/sec" | (I think he's suggesting that the Capital M should be a lowercase in the Answer line ??? -KA) |  | Correction made
David T. Fetzer, MD | Closed Issues | 150 (& 620) | L | "Tidal respiration" may be confusing | "Consider end inspiration or end expiration" | The point on confusion is well taken since end inspiration and end expiration are not the same as suspended tidal respiration. Although all physicians are trained to know what tidal respiration is, other workers may not be so I propose “suspended normal respiration (suspended tidal respiration)” | Suspended respiration is neither end inspiration or end expiration but simply stopping breathing during quiet breathing whenever the operator says stop. This gives a relatively reproducible degree of inspiration because 1. tidal respirations are not deep so there is not a large difference between the various stages of respiration. 2. tidal respirations are typically quite reproducible from breath to breath. 3. suspended respiration typically gives a degree of inspiration in between the extremes of end inspiration and end expiration further reducing variability without requiring extra patient training.
| Closed Issues - # Measurements | 150 | M | "I do not believe there is enough literature to support obtaining only 5 measurements." | "Focus on 10 measurements until more literature is available." | DISCUSS WITH SWS COMMITTEE. I BELIEVE WE SHOULD LEAVE IT AS IS. | Committee: suggestion that increased # of measurements may be doable if initial values are tight. But no tool to do this and no threshold has been established. Will re-examine this for next version. Will move back to 10 but say that mfr may specify a smaller number.
| Closed Issues - # Measurements | 150 | H | "The IQR/Median of 0.3 is if measurements are being reported in stiffness (kPa). If m/s, then this cut-off is actually 0.15." | none listed |  | Clarification wording has been added for each instance throughout the document. It should be noted that many references to IQR <0.3 are present in appendix D-manufacturer’s guidelines and these references have not been amended. Manufacturers will be notified of the need for clarification.
| Closed Issues - Max Liver Depth | 150 | M | "Some recommendations report a distance of 1.5 cm from the capsule, instead of 2.0 cm." | none listed |  | The committee will watch that discussion in case a change needs to be made.
| 2.1 Proposed Clinical Interpretation | 256-271 | M | "The SRU consensus panel has submitted an update for publication that changes the docus from staging fibrosis (eg. Metavir F0 or F4), to conveying a risk of clinically-significant fibrosis (compensated advanced chronic liver disease)." | none listed |  | Committee agrees with this response: To be able to release the profile in a field of constant change, content was frozen at the end of 2019. Once the SRU panel changes are published the SWS committee will determine whether to include them. The section 2.1 on Proposed Clinical Interpretation is provided only as a rough guide to possible use and is not the focus of the profile although we will try to keep in reasonably current as practice evolves.
| 3.3 Staff Qualification | 527 | L | "How was 20 chosen? What if a site has either a large number of sonographers or a small number of elastography requests that would make this requirement impractical. Also, what if a site is unable to obtain an elastic phantom?" | none listed |  | How was 20 chosen? This number was chosen in a somewhat arbitrary fashion but partly because of central limit theorem considerations. It is not based on a statistical argument but rather expert consensus. In section 4 when dealing with conformance, the N chosen has a statistical justification. "With 6 subjects and 10 replicates per subject per depth, and with claims stating wCV of 4% and 5% for depths of 4.5 and 7.0 for moderate SWS values, the maximum allowed wCVs are 3.3 and 4.1 for depths of 4.5 and 7.0, respectively." What if a site has either a large number of sonographers or a small number of elastography requests? In those cases a subset of sonographers can be qualified rather than all of them. In fact not all sonographers will be qualified even after training. These criteria are meant to establish who should be doing the measurements. Gradually over time more sonographers will accumulate enough cases to establish whether they are qualified or not. What if a site is unable to obtain an elastic phantom? Appropriate phantoms will be available from CIRS initially at reasonable cost so this should not be an issue. Alternatively a small clinic may borrow a phantom from a clinic who has one already.
| 3.5 Subject Selection | 574 | M | "For an FDA-approved clinical technique, why would informed consent be needed?" | none listed |  | The FDA by law does not regulate clinical techniques as these are considered to be the practice of medicine. The assumption is that medical boards and medical societies will regulate the practice of medicine. For research protocols IRB approval will likely require a consent. Most institutions have requirements for consent that must be followed if a procedure is performed on their premises. The FDA has the authority to override local IRB’s in certain limited situations but this does not include procedural matters such as obtaining informed consent.
| 3.6 Subject Handling | N/A | M | "Would it be prudent for an actor to verify other potential confounders such as presence of markedly elevated AST/ALT, evidence of biliary obstruction, and evidence of right heart failute/passive hepatic congestion?" | none listed |  | I propose the following response. The goal of the profile is to provide a means by which accurate SWS measurements may be made. The confounders mentioned affect SWS and must be considered when using SWS clinically as a surrogate for liver fibrosis. The clinical use and interpretation of the SWS measurements is constantly changing and is beyond the scope of this profile.
| 3.7.2 Transducer Position | 723 | M | "Why not include hepatic segment 6? In some patterns of hypertrophy/atrophy, 6 is quite prominent, and is easily imaged, esp when the patient is 30 degrees semi-decub." |  | Will review other standards re: segment 6 inclusion for future versions. Users should be guided by machine specific mfr recommendations. Accordingly the reference to liver segments is removed (the parenthetical). | The committee will review other standards with respect to inclusion of segment 6 for future versions of the profile. Since listing segments is not very helpful for someone when acquiring SWS data, and since the actual acquisition will be guided by the manufacturer’s recommendations, the reference to liver segments has been removed from this version of the profile.
(AAPM) Cynthia McCollough, PhD, FAAPM, FACR, AFIMBE; Response replies to Richard J. Martin, JD | General Remark | N/A | N/A | "Essentially, the liver is not as homogenous as we had hoped. Considering the current recommendation of sampling just one area, AAPM asks what do the committee members think about the limited sampling issue and how do the users know which area is the best to measure?" |  |  | The focus of a QIBA profile is to provide a method for accurately estimating a parameter from an image or using an imaging system. In the profile we discuss how to make the measurement accurately. The exact location of the region from which measurements are made and how many different regions should be measured are beyond the scope of the QIBA profile and depend on the disease process causing the fibrosis among other factors. This will not be covered in this QIBA profile but will likely appear in practice of ultrasound documents published by the WFUMB and AIUM. This is discussed in Section 3.9.1 SWS Image Acquisition (SWEI) and Point SWS Measurement.
| General Remark | N/A | N/A | Have any members been struggling with significant underestimation in stiffness with US shear wave elastography (SWE) compared to magnetic resonance elastography (MRE)? Part of the value of imaging methods of assessing liver fibrosis compared to biopsy approaches is that imaging methods allow evaluation of the entire liver. Sampling only one area of the US image discards important information about heterogeneity in the liver. |  |  | The differences between SWE and MRE measurements may be due to the markedly different shear wave frequencies used by the two modalities and the fact that the liver is not purely elastic. The discrepancy in values may be addressed in the next version of the profile which deals with viscoelastic tissue measurements instead of presuming that the tissue is purely elastic. Shear wave elastographic imaging records heterogeneity of liver stiffness as well as MRE. As stated in the profile, selection of which area in a heterogeneous area of liver stiffness depends on the needs of the clinician and the underlying cause of the fibrosis and is not part of the QIBA profile as discussed above. See the topic: Point Shear Wave Speed Measurements from Shear Wave Images in section 3.7. SWS Image Acquisition (SWEI) and Point SWS Measurement
| General Remark | N/A | N/A | Discussion of pediatric patients is limited. AAPM asks whether pediatric patients have been adequately considered in this Profile. |  |  | This profile was originally created for liver fibrosis due to hepatitis C. It may be extended to NASH but it is not a general profile and not intended for use in Pediatrics (though it may work and people are free to use it for that purpose). See the section 2.1 on proposed clinical interpretation. A later revision of the profile that contains specifics for pediatrics may appear as might a different version just for pediatric use.
| Closed Issues | 105,(see also P20 L381 et seq., P21 L415 et seq.) | N/A | The specifications of the phantom are not consistent. Attenuation is given as 0.6 or 0.5 dB/cm/MHz with a standard deviation of 0.1 or 0.2. In addition, speed of sound value and tolerance range are also not consistent. |  |  | Attenuation was changed to be 0.5 dB/cm/MHz in all cases and tolerance to .1 for consistency. SOS tolerance on procurement is plus minus 20 but for later checking is plus minus 30. There is nothing inconsistent about relaxing the tolerance for follow up testing. Changed the SOS tolerance to ± 30 (2%) for consistency
| Closed Issues | 105 (See also P26 L557, P30 L659, P32 L725) |  | There are several places talking about measurements being 2 cm below the liver capsule and within 6.5 cm from the probe surface. It would be helpful to define this depth: Is it the center/top/bottom of the ROI within the point SWE and 2D SWE image? A definition similar to that given on page 37 would be helpful. | I EXPECT THAT THE UPPER EDGE OF THE ROI SHOULD BE NO CLOSER THAN 2CM BUT ELSEWHERE WHEN DISCUSSING DEPTH WE USE THE CENTER OF THE ROI…..SHOULD BE JUST USE THE CENTER OR SWITCH TO THE UPPER EDGE OR USE BOTH? |  | Committee: use center of ROI but specify that if artifacts are present on 2D they should be avoided.
| 2. Clinical Content | 161 |  | The phrase "moderate but significant liver fibrosis" should be changed to “moderate to significant.” | No, the correct reading is moderate BUT significant. Sentence left as is. |  | No, the correct reading is moderate BUT significant. Sentence left as is. Verified as complete.
| 2. Clinical Content & Claims | 228-230 and 246-247 |  | Measurements in viscoelastic medium can be quite different for scanners from different vendors. The cut points in the table of page 15 for System A and System B seems to differ by 20% already (also see different cut points given on page 51 and page 56 for two vendors), which seems to suggest that the scannerdifference for in vivo measurement can be as high as 20%. However, some reproducibility coefficients in tables 2-3 and 2-4 are below 20% for the claim of using different ultrasound scanners. Therefore, some clarifications may be needed to address this gap. | The current profile claims only cover measurements of purely elastic materials. A future version will extend claims to visco-elastic materials. |  | The current profile claims only cover measurements in purely elastic materials as visco-elastic phantoms are not generally available and further evaluation of performance in visco-elastic phantoms/materials is needed to create valid claims. A statement explaining all this has been inserted at the end of the executive summary just prior to section 2.
| 2.1 Proposed Clinical Interpretation | 258 |  | Replace missing period mark. |  |  | Correction made
| 2.1 Proposed Clinical Interpretation | 279 |  | Closing parenthesis is missing and the line should be re-aligned. |  |  | Will revisit after reformatting to the latest profile configuration.
| 2.2 Discussion | 294 |  | Should remove the words 'or not'. The phrase is used two times. |  |  | The change has been made.
| 3.1.1 Pre-delivery | 395 |  | Do the phantom specifications make sense for a pediatric patient as well (attenuation & backscatter)? |  |  | The phantom specifications are for “healthy” liver. Patients may have wide variation in the parameters but producing phantoms with a wide range of various acoustic parameters would be costly and testing on all those phantoms would be unnecessarily burdensome for clinical sites given our current knowledge. If problems with SWS estimates appear during the clinical confirmation phase of profile testing then additional phantoms may be required. Clinical confirmation is currently beyond the scope of QIBA due to the need for a costly multicenter trial.
| 3.1.1 Pre-delivery | 434 |  | Are there any calibration specifications for a scale that would weigh the phantom? |  |  | The committee has not approved any specification for the type of scale to use. Any scale with the appropriate capacity to weigh the phantom should be sufficiently accurate. Since a phantom is typically around 3kg in mass (weight) and the specifications call for detecting a change of 0.5% or .005 which equals 15 g, a scale with plus minus 2g accuracy is more than adequate. I will put approximately plus minus 1g and try to get committee approval, if i cannot get it, the profile will go out without a specific specification.
| 3.1.1. Pre-delivery | 442 |  | “The specifications of the phantom are not consistent. Attenuation is given as 0.6 or 0.5 dB/cm/MHz with a standard deviation of 0.1 or 0.2. In addition, speed of sound value and tolerance range are also not consistent.” | The correct values is: .5 dB/cm/MHz ± 0.1. Will make all listings consistent. |  | All listings are verified as being correct.
| 3.1.1 Pre-delivery | 466 |  | It may be helpful to add discussions of the impact of extreme temperature during shipping of calibration phantoms and any precautions to take during shipping. | COMMITTEE, SHOULD WE INCLUDE A STATEMENT ON SHIPPING? IF SO HOW SHOULD THAT READ AND WHERE SHOULD WE PUT IT? |  | Temp of acquisition should be recorded in case of phantom temp dependence. Acquisition temp already specified. Shipping should be according manufuacturers recommendations. Added as new section titled SWS Phantom Temp Sensitivity and shipping Considerations has been added at the end of 3.1.1.
| 3.1.2 Phantom Testing | 473 |  | The link is broken; cannot evaluate the AIUM QA manual for Gray Scale US Scanners. | contact AIUM | AIUM response: only available to members | Removed link pending another solution—footnote stays.
| 3.1.3 Specification | 482 |  | Are the MFR specifications and certification available to the user? Only available upon request? | MANUFACTURERS: IS THIS INFORMATION AVAILABLE IN THE FDA REQUIRED LABELING? OR ONLY ON REQUEST? |  | Variable based on manufacturer. Will appear in the manual. User may consult the manual. Added a footnote containing this information
| 3.5.1 Subject Selection | 557 |  | >=2cm distance from the liver capsule seems to be different compared to most existing literatures which state ~1.5-2cm. Are there any data to support this decision? In addition, how to handle the inconsistency with the vendor recommendations? For example, >=1cm from Canon (page 48). |  |  | The SWS committee believes that 2cm is safer from a reproducibility standpoint than 1.5cm and still falls within the literature specifications. This section is a discussion to explain how body wall thickness may affect patient selection. For actual acquisition, the profile specifies that the manufacturer recommendations should be followed.
| 3.5.1 Subject Selection | 557 |  | Is the maximum 6.5cm recommendation corresponding to the lower edge of the acquisition ROI? What about patients with ascites? Because of the lower attenuation, many of these cases appear to have sufficient signal at 6.5cm or slightly deeper, even though they are more susceptible to motion artifacts. |  |  | The 6.5cm maximum refers to placement of the lower (deeper) edge of the ROI. The 2cm minimum depth refers to placement of the upper edge of the ROI. This has been clarified in the profile. Depths greater than 6.5cm have not been tested by the committee, so we cannot guarantee that the claims will be met. This issue will be re-examined for the next version of the profile. We have also added a phrase pointing out the consequences of not placing the ROI deeper than 2cm from the liver capsule in a patient with a body wall thickness of >4cm.
| 3.7.1 SWS Image Acquisition | 611-614 |  | The efficacy of phantom calibration may depend on whether the phantom is viscoelastic or elastic. Elastic phantom may not be able to provide a good calibration among different scanners for in vivo liver measurements. Some comments about phantom properties may be helpful here. |  | Committee suggestion: add elastic vs visco statement line 191 | The current profile is based on our experiments with elastic phantoms only, but our data from viscoelastic phantoms seem to indicate that the claims will be similar for visco-elastic phantoms and tissues. A statement clarifying why this version of the profile is based on elastic phantoms only has been added to the end of the executive summary at the beginning of the profile. The statement reads: “Strictly speaking, the claims in this version of the profile only apply to purely elastic materials and phantoms. This is because visco-elastic phantoms are generally not available for sites to verify the profile claims and the committee must further verify the profile claims for a clinically relevant range of visco-elastic materials. Claims for visco-elastic phantoms and tissues will appear in the next version of the profile.”
| 3.7.1 Transducer Pressure | 651 |  | Newer data seem to suggest that stronger pressure is fine for intercostal acquisition, which actually decreases the acquisition failure rate without changing the quantification results (Byenfeldt, et al. 2019 Ultrasound in Med & Biol). One manufacturer suggests using this approach. Obviously, it would be different for subcostal acquisition. |  |  | The committee feels that there are not yet enough data to recommend higher pressure and how much higher that might be. This wilil be re-evaluated for the next version of the profile.
| 3.7.2 Specification | 725 |  | Under parameter/number of measurements, "Shall make a minimum of 5 measurements should be made" should be changed to “Shall make a minimum of 5 measurements.” |  |  | After public comment/SWS Comm review changed back to 10 measurements except when mfr suggests another number.
| 4.1.2 Image Performance | 772-773 |  | The link is broken. |  |  | The footnote listing has been retained but the broken link has been removed. A link to the reference library has been tentatively inserted pending an answer from AIUM regarding a stable link and note has been made that one must login to the AIUM site to obtain the document.
| 4.2.1 Site Assessment Tools & Tests | 809 |  | Are there any pediatric-specific phantoms that are required during this evaluation? It states only that two phantoms will be used. |  |  | Pediatric specific phantoms should not be required since existing phantoms test system performance for probes typically used for pediatric imaging as well as those for adults. A phrase briefly clarifying what the two phantoms are has been added.
| 4.2.1.2.B | 945 |  | There needs to be a closing parenthesis to end the sentence. |  |  | Parenthesis has been added.
| Appendix D | 1300 et seq. |  | For GE LE9 users, it is extremely important to maintain the elastography gain at 55. It is inherently tied to a quality threshold. It's an easy change for the operators to make, which could actually alter the quantitative measurements. Similar requirement exists on S8 with a different gain value. |  |  | This issue has been referred to GE who are the only ones who can make changes to their acquisition protocol in appendix D. If they provide an updated acquisition protocol it will be incorporated at the next revision of the profile.
Note on A7: Alton, Karen *: QST Hospital, National Institute of Radiological Sciences 4-9-1, Anagawa, Chiba 262-0033, Japan
Note on A18: Alton, Karen *: QST Hospital, National Institute of Radiological Sciences 4-9-1, Anagawa, Chiba 262-0033, Japan
Note on A28: Alton, Karen *: Cynthia McCollough President, AAPM Brooks-Hollern Professor Professor of Medical Physics and Biomedical Engineering Director, CT Clinical Innovation Center and X-ray Imaging Core Department of Radiology, Mayo Clinic