## Altogether
Parameter | Actor | Requirement | Coronary? | Carotid? | Conform? | Opinion?
Use of intravenous contrast | Imaging Physician | Shall prescribe a contrast protocol to achieve appropriate lumen conspicuity relative to wall tissues. | Yes | Yes
Use of intravenous contrast | Technologist | Shall use the prescribed intravenous contrast protocol. | Yes | Yes
Artifact Sources | Technologist | Shall remove or position potential sources of artifacts (specifically including breast shields, metal-containing clothing, EKG leads, and other metal equipment) such that they will not degrade the reconstructed CT image. | Yes | Yes
Breath hold | Technologist | Shall instruct the subject in proper breath-hold and start image acquisition shortly after full inspiration, taking into account the lag time between full inspiration and diaphragmatic relaxation. | Yes | No
Table Height & Centering | Technologist | Shall adjust the table height for the mid-axillary plane to pass through the isocenter. Shall center the thorax shall be centered in the AP and L/R directions according to the following: table height shall be adjusted for the mid axillary plane to pass through the isocenter and the sagittal laser line shall pass through the sternum from suprasternal notch to xiphoid process. | Yes | No
Nitrates | Technologist | Shall administer nitrates as prescribed, 5-7 minutes after nitro is administered. | Yes | No
In-plane Spatial Resolution | Acquisition Device | Shall validate that the protocol achieves an f50 value that is greater than 0.35 line pairs per mm for both air and soft tissue edges. | Yes | Yes
In-plane Spatial Resolution | Acquisition Device | See section 4.1. Assessment Procedure: In-plane Spatial Resolution | Yes | Yes
Pixel noise | Acquisition Device | Shall validate that the protocol achieves a standard deviation that is < 30HU. See 4.2. Assessment Procedure: Pixel noise | Yes | Yes
Acquisition Protocol | Acquisition Device | Shall be capable of making validated protocols (designed and validated by the manufacturer and/or by the site) available to the technologist at scan time. | Yes | Yes
Acquisition Protocol | Physicist | Shall prepare a protocol to meet the specifications in this table. | Yes | Yes
Acquisition Protocol | Physicist | Shall ensure technologists have been trained on the requirements of this profile. | Yes | Yes
Acquisition Protocol | Technologist | Shall select a protocol that has been previously prepared and validated for this purpose. | Yes | Yes
Total Collimation Width | Imaging Physician | Shall set to Greater than or equal to 18mm. | Yes | No
Nominal Tomographic Section Thickness (T) | Physicist | Shall set to Less than or equal to 0.75mm. | Yes | No
Temporal Resolution | Acquisition Device | Shall achieve an effective rotation time of less than or equal to 400ms. | Yes | No
Total Collimation Width | Physicist | Shall set to Greater than or equal to 16mm. | No | Yes
Nominal Tomographic Section Thickness (T) | Physicist | Shall set to Less than or equal to 1.0mm. | No | Yes
Reconstruction Protocol | Physicist | Shall prepare a protocol to meet the specifications in this table. Shall ensure technologists have been trained on the requirements of this profile. | Yes | Yes
Reconstruction Protocol | Reconstruction Software | Shall be capable of performing reconstructions and producing images with all the parameters set as specified "Protocol Design Specification". | Yes | Yes
Reconstruction Protocol | Technologist | Shall select a protocol that has been previously prepared and validated for this purpose. | Yes | Yes
ECG Gating | Technologist | Shall use prospective ECG gating and consider iterative reconstruction to allow for the lowest possible radiation exposure. If the heart rate is too high, retrospective ECG gating may be required to obtain optimal images. | Yes | No
Reconstructed Image Thickness | Physicist | Shall be less than 1mm. | Yes | Yes
Reconstructed Image Thickness | Technologist | Shall be less than 1mm if not set in the protocol. | Yes | Yes
Reconstructed Image Interval | Physicist | Shall set to less than or equal to the Reconstructed Image Thickness (i.e. no gap, may have overlap). | Yes | Yes
Reconstructed Image Interval | Technologist | Shall set to less than or equal to the Reconstructed Image Thickness (i.e. no gap, may have overlap) and consistent with baseline. | Yes | Yes
Reconstructed In-plane Voxel Size | Physicist | Shall set to less than or equal to 0.625mm | Yes | Yes
In-plane Spatial Resolution | Physicist | Shall validate that the protocol achieves an f50 value that is Greater than 0.35 mm-1 for both air and soft tissue edges. See section 4.1. Assessment Procedure: In-plane Spatial Resolution | Yes | Yes
Pixel noise | Physicist | Shall validate that the protocol achieves a standard deviation that is < 30HU. See section 4.2. Assessment Procedure: Pixel noise | Yes | Yes
Image Header | Reconstruction Software | Shall record in the DICOM image header the actual values for the tags listed in the DICOM Tag column "Protocol Design Specification" as well as the model-specific Reconstruction Software parameters utilized to achieve conformance. | Yes | Yes
Reconstruction Field of View | Technologist | Shall ensure the Field of View spans at least the full extent of the thoracic cavity, but not substantially greater than that. | Yes | Yes
Image Header | Reconstruction Software | Shall record in the DICOM image header the actual values for the tags listed in the DICOM Tag column "Protocol Design Specification" as well as the model-specific Reconstruction Software parameters utilized to achieve conformance. | Yes | Yes
Patient Motion Artifacts | Imaging Physician | Shall confirm the images containing the lesion are free from artifact due to motion. | Yes | Yes
Physiological motion artifact (particularly cardiac) | Imaging Physician | Shall confirm the images containing the lesion are free from artifact due to motion based on visual review for blurred anatomic features. | Yes | Yes
Artifacts | Imaging Physician | Shall confirm the images containing the lesion are free from artifacts due to dense objects, anatomic positioning (e.g., arms down at sides), or equipment issues (e.g., ring artifacts). | Yes | Yes
Contrast Enhancement | Imaging Physician | Shall confirm that the intravascular level of contrast enhancement, if any, is appropriate for evaluating the lesion. | Yes | Yes
Patient Positioning Consistency | Imaging Physician | Shall confirm that any lesion deformation due to patient positioning is consistent with baseline (e.g. lesions may deform differently if the patient is supine in one scan and prone in another). | Yes | Yes
Scan Plane Consistency | Imaging Physician | Shall confirm that the anatomical slice orientation (due to gantry tilt or patient head/neck repositioning) is consistent with baseline. | Yes | Yes
Field of View | Imaging Physician | Shall confirm that the image field of view (FOV) resulting from acquisition and reconstruction settings appears consistent with baseline. | Yes | Yes
Pacemaker leads, stents | Imaging Physician | Shall confirm that anatomy assessed does not contain metal artifacts. | Yes | Yes
Vessel structure | Image Analysis Tool | Shall be validated to achieve bias and linearity (expressed as intercept, slope, and quadratic term) within the values shown in the following table. See 4.3. Assessment Procedure: Vessel Structure Bias and Linearity, noting that the full 95% confidence intervals (not only the point estimates) shall meet or exceed the indicated specifications when tested over range as given in Claims section: Lumen Area (mm2) Bias: ±2, Intercept: ±1.0, Slope: 1±.1, Quadratic term: ±.1 Wall Area (mm2) Bias: 2, Intercept: ±10, Slope: 1±.1, Quadratic term: ±.1 Maximum Wall Thickness (mm) Bias: ±1, Intercept: ±1, Slope: 1±.1, Quadratic term: ±.1 Plaque Burden (ratio) Bias: ±0.1, Intercept: ±.1, Slope: 1±.1, Quadratic term: ±.1 | Yes | Yes
Tissue Composition | Image Analysis Tool | Shall be validated to achieve bias and linearity (expressed as intercept, slope, and quadratic term) within the values shown in the following table. See 4.4. Assessment Procedure: Tissue Characteristics Bias and Linearity, noting that the full 95% confidence intervals (not only the point estimates) shall meet or exceed the indicated specifications when tested over range as given in Claims section: Calcified Area (mm2) Bias: ±1.5, Intercept: ±2, Slope: 1±.5, Quadratic term: ±.1 LRNC Area (mm2) Bias: ±3, Intercept: ±3.5, Slope: 1±.8, Quadratic term: ±.3 | Yes | Yes
Reader variability | Image Analysis Tool | Shall be validated to achieve Intra-reader wSD and Inter-reader wSD less than the values shown in the following table. See 4.5. Assessment Procedure: Reader / Image Analysis Tool Variability, noting that the full 95% confidence intervals (not only the point estimates) shall meet or exceed the indicated specifications when tested over range as given in Claims section. Lumen Area (mm2) Intra-reader wSD: 2.5, Inter-reader wSD: 5.0 Wall Area (mm2) Intra-reader wSD: 2.5, Inter-reader wSD: 5.0 Maximum Wall Thickness (mm) Intra-reader wSD: 0.75, Inter-reader wSD: 1.0 Plaque Burden (ratio) Intra-reader wSD: 0.1, Inter-reader wSD: 0.1 Calcified Area (mm2) Intra-reader wSD: 1.0, Inter-reader wSD: 1.5 LRNC Area (mm2) Intra-reader wSD: 1.0, Inter-reader wSD: 1.5 | Yes | Yes
Basis of cross-sectional area results | Image Analysis Tool | Shall base cross-sectional area results on obliquely-resliced orthogonal to centerline at spacing less than or equal to 0.5mm | Yes | Yes
Basis of volume results | Image Analysis Tool | Shall base volume results on three-dimensional object definitions (specifically excluding methods such as determining cross-sectional areas and multiplying by the slice thickness, or other approximations) | Yes | Yes
Confidence interval | Image Analysis Tool | Shall be able to display to the Imaging Physician, for each measurand, the range of plausible values for the given measurement stated in terms of the completed validation for the tool as a 95% interval. | Yes | Yes
Result Verification | Imaging Physician | Shall review & approve segmentations produced by the Image Analysis Tool. | Yes | Yes
Multiple Lesions | Image Analysis Tool | Shall allow multiple lesions to be measured. | Yes | Yes
Multiple Lesions | Image Analysis Tool | Shall either correlate each measured lesion across encounters or support the Imaging Physician to unambiguously correlate them. | Yes | Yes
Multiple encounters | Imaging Physician | Shall re-process the first encounter if it was processed by a different Image Analysis Tool or Imaging Physician. | Yes | Yes
Multiple encounters | Image Analysis Tool | Shall be able to present the reader with both encounters side-by-side for comparison when processing the second encounter. | Yes | Yes
Multiple encounters | Image Analysis Tool | Shall be able to re-process the first encounter (e.g. if it was processed by a different Image Analysis Tool or Imaging Physician). | Yes | Yes

## Imaging Physician
Parameter | Actor | Requirement | Coronary? | Carotid? | Conform? | Opinion?
Use of intravenous contrast | Imaging Physician | Shall prescribe a contrast protocol to achieve appropriate lumen conspicuity relative to wall tissues. | Yes | Yes
Total Collimation Width | Imaging Physician | Shall set to Greater than or equal to 18mm. | Yes | No
Patient Motion Artifacts | Imaging Physician | Shall confirm the images containing the lesion are free from artifact due to motion. | Yes | Yes
Physiological motion artifact (particularly cardiac) | Imaging Physician | Shall confirm the images containing the lesion are free from artifact due to motion based on visual review for blurred anatomic features. | Yes | Yes
Artifacts | Imaging Physician | Shall confirm the images containing the lesion are free from artifacts due to dense objects, anatomic positioning (e.g., arms down at sides), or equipment issues (e.g., ring artifacts). | Yes | Yes
Contrast Enhancement | Imaging Physician | Shall confirm that the intravascular level of contrast enhancement, if any, is appropriate for evaluating the lesion. | Yes | Yes
Patient Positioning Consistency | Imaging Physician | Shall confirm that any lesion deformation due to patient positioning is consistent with baseline (e.g. lesions may deform differently if the patient is supine in one scan and prone in another). | Yes | Yes
Scan Plane Consistency | Imaging Physician | Shall confirm that the anatomical slice orientation (due to gantry tilt or patient head/neck repositioning) is consistent with baseline. | Yes | Yes
Field of View | Imaging Physician | Shall confirm that the image field of view (FOV) resulting from acquisition and reconstruction settings appears consistent with baseline. | Yes | Yes
Pacemaker leads, stents | Imaging Physician | Shall confirm that anatomy assessed does not contain metal artifacts. | Yes | Yes
Result Verification | Imaging Physician | Shall review & approve segmentations produced by the Image Analysis Tool. | Yes | Yes
Multiple encounters | Imaging Physician | Shall re-process the first encounter if it was processed by a different Image Analysis Tool or Imaging Physician. | Yes | Yes

## Technologist
Parameter | Actor | Requirement | Coronary? | Carotid? | Conform? | Opinion?
Use of intravenous contrast | Technologist | Shall use the prescribed intravenous contrast protocol. | Yes | Yes
Artifact Sources | Technologist | Shall remove or position potential sources of artifacts (specifically including breast shields, metal-containing clothing, EKG leads, and other metal equipment) such that they will not degrade the reconstructed CT image. | Yes | Yes
Breath hold | Technologist | Shall instruct the subject in proper breath-hold and start image acquisition shortly after full inspiration, taking into account the lag time between full inspiration and diaphragmatic relaxation. | Yes | No
Table Height & Centering | Technologist | Shall adjust the table height for the mid-axillary plane to pass through the isocenter. Shall center the thorax shall be centered in the AP and L/R directions according to the following: table height shall be adjusted for the mid axillary plane to pass through the isocenter and the sagittal laser line shall pass through the sternum from suprasternal notch to xiphoid process. | Yes | No
Nitrates | Technologist | Shall administer nitrates as prescribed, 5-7 minutes after nitro is administered. | Yes | No
Acquisition Protocol | Technologist | Shall select a protocol that has been previously prepared and validated for this purpose. | Yes | Yes
Reconstruction Protocol | Technologist | Shall select a protocol that has been previously prepared and validated for this purpose. | Yes | Yes
ECG Gating | Technologist | Shall use prospective ECG gating and consider iterative reconstruction to allow for the lowest possible radiation exposure. If the heart rate is too high, retrospective ECG gating may be required to obtain optimal images. | Yes | No
Reconstructed Image Thickness | Technologist | Shall be less than 1mm if not set in the protocol. | Yes | Yes
Reconstructed Image Interval | Technologist | Shall set to less than or equal to the Reconstructed Image Thickness (i.e. no gap, may have overlap) and consistent with baseline. | Yes | Yes
Reconstruction Field of View | Technologist | Shall ensure the Field of View spans at least the full extent of the thoracic cavity, but not substantially greater than that. | Yes | Yes

## Physicist
Parameter | Actor | Requirement | Coronary? | Carotid? | Conform? | Opinion?
Acquisition Protocol | Physicist | Shall prepare a protocol to meet the specifications in this table. | Yes | Yes
Acquisition Protocol | Physicist | Shall ensure technologists have been trained on the requirements of this profile. | Yes | Yes
Nominal Tomographic Section Thickness (T) | Physicist | Shall set to Less than or equal to 0.75mm. | Yes | No
Total Collimation Width | Physicist | Shall set to Greater than or equal to 16mm. | No | Yes
Nominal Tomographic Section Thickness (T) | Physicist | Shall set to Less than or equal to 1.0mm. | No | Yes
Reconstruction Protocol | Physicist | Shall prepare a protocol to meet the specifications in this table. Shall ensure technologists have been trained on the requirements of this profile. | Yes | Yes
Reconstructed Image Thickness | Physicist | Shall be less than 1mm. | Yes | Yes
Reconstructed Image Interval | Physicist | Shall set to less than or equal to the Reconstructed Image Thickness (i.e. no gap, may have overlap). | Yes | Yes
Reconstructed In-plane Voxel Size | Physicist | Shall set to less than or equal to 0.625mm | Yes | Yes
In-plane Spatial Resolution | Physicist | Shall validate that the protocol achieves an f50 value that is Greater than 0.35 mm-1 for both air and soft tissue edges. See section 4.1. Assessment Procedure: In-plane Spatial Resolution | Yes | Yes
Pixel noise | Physicist | Shall validate that the protocol achieves a standard deviation that is < 30HU. See section 4.2. Assessment Procedure: Pixel noise | Yes | Yes

## Acquisition Device
Parameter | Actor | Requirement | Coronary? | Carotid? | Conform? | Opinion?
In-plane Spatial Resolution | Acquisition Device | Shall validate that the protocol achieves an f50 value that is greater than 0.35 line pairs per mm for both air and soft tissue edges. | Yes | Yes
In-plane Spatial Resolution | Acquisition Device | See section 4.1. Assessment Procedure: In-plane Spatial Resolution | Yes | Yes
Pixel noise | Acquisition Device | Shall validate that the protocol achieves a standard deviation that is < 30HU. See 4.2. Assessment Procedure: Pixel noise | Yes | Yes
Acquisition Protocol | Acquisition Device | Shall be capable of making validated protocols (designed and validated by the manufacturer and/or by the site) available to the technologist at scan time. | Yes | Yes
Temporal Resolution | Acquisition Device | Shall achieve an effective rotation time of less than or equal to 400ms. | Yes | No

## Reconstruction Software
Parameter | Actor | Requirement | Coronary? | Carotid? | Conform? | Opinion?
Reconstruction Protocol | Reconstruction Software | Shall be capable of performing reconstructions and producing images with all the parameters set as specified "Protocol Design Specification". | Yes | Yes
Image Header | Reconstruction Software | Shall record in the DICOM image header the actual values for the tags listed in the DICOM Tag column "Protocol Design Specification" as well as the model-specific Reconstruction Software parameters utilized to achieve conformance. | Yes | Yes
Image Header | Reconstruction Software | Shall record in the DICOM image header the actual values for the tags listed in the DICOM Tag column "Protocol Design Specification" as well as the model-specific Reconstruction Software parameters utilized to achieve conformance. | Yes | Yes

## Image Analysis Tool
Parameter | Actor | Requirement | Coronary? | Carotid? | Conform? | Opinion?
Vessel structure | Image Analysis Tool | Shall be validated to achieve bias and linearity (expressed as intercept, slope, and quadratic term) within the values shown in the following table. See 4.3. Assessment Procedure: Vessel Structure Bias and Linearity, noting that the full 95% confidence intervals (not only the point estimates) shall meet or exceed the indicated specifications when tested over range as given in Claims section: Lumen Area (mm2) Bias: ±2, Intercept: ±1.0, Slope: 1±.1, Quadratic term: ±.1 Wall Area (mm2) Bias: 2, Intercept: ±10, Slope: 1±.1, Quadratic term: ±.1 Maximum Wall Thickness (mm) Bias: ±1, Intercept: ±1, Slope: 1±.1, Quadratic term: ±.1 Plaque Burden (ratio) Bias: ±0.1, Intercept: ±.1, Slope: 1±.1, Quadratic term: ±.1 | Yes | Yes
Tissue Composition | Image Analysis Tool | Shall be validated to achieve bias and linearity (expressed as intercept, slope, and quadratic term) within the values shown in the following table. See 4.4. Assessment Procedure: Tissue Characteristics Bias and Linearity, noting that the full 95% confidence intervals (not only the point estimates) shall meet or exceed the indicated specifications when tested over range as given in Claims section: Calcified Area (mm2) Bias: ±1.5, Intercept: ±2, Slope: 1±.5, Quadratic term: ±.1 LRNC Area (mm2) Bias: ±3, Intercept: ±3.5, Slope: 1±.8, Quadratic term: ±.3 | Yes | Yes
Reader variability | Image Analysis Tool | Shall be validated to achieve Intra-reader wSD and Inter-reader wSD less than the values shown in the following table. See 4.5. Assessment Procedure: Reader / Image Analysis Tool Variability, noting that the full 95% confidence intervals (not only the point estimates) shall meet or exceed the indicated specifications when tested over range as given in Claims section. Lumen Area (mm2) Intra-reader wSD: 2.5, Inter-reader wSD: 5.0 Wall Area (mm2) Intra-reader wSD: 2.5, Inter-reader wSD: 5.0 Maximum Wall Thickness (mm) Intra-reader wSD: 0.75, Inter-reader wSD: 1.0 Plaque Burden (ratio) Intra-reader wSD: 0.1, Inter-reader wSD: 0.1 Calcified Area (mm2) Intra-reader wSD: 1.0, Inter-reader wSD: 1.5 LRNC Area (mm2) Intra-reader wSD: 1.0, Inter-reader wSD: 1.5 | Yes | Yes
Basis of cross-sectional area results | Image Analysis Tool | Shall base cross-sectional area results on obliquely-resliced orthogonal to centerline at spacing less than or equal to 0.5mm | Yes | Yes
Basis of volume results | Image Analysis Tool | Shall base volume results on three-dimensional object definitions (specifically excluding methods such as determining cross-sectional areas and multiplying by the slice thickness, or other approximations) | Yes | Yes
Confidence interval | Image Analysis Tool | Shall be able to display to the Imaging Physician, for each measurand, the range of plausible values for the given measurement stated in terms of the completed validation for the tool as a 95% interval. | Yes | Yes
Multiple Lesions | Image Analysis Tool | Shall allow multiple lesions to be measured. | Yes | Yes
Multiple Lesions | Image Analysis Tool | Shall either correlate each measured lesion across encounters or support the Imaging Physician to unambiguously correlate them. | Yes | Yes
Multiple encounters | Image Analysis Tool | Shall be able to present the reader with both encounters side-by-side for comparison when processing the second encounter. | Yes | Yes
Multiple encounters | Image Analysis Tool | Shall be able to re-process the first encounter (e.g. if it was processed by a different Image Analysis Tool or Imaging Physician). | Yes | Yes

## pick lists
Conform
Do not conform
Routinely performed
Feasible, will do to conform
Feasible, but not going to do it
Not feasible (explain why)