if Corrected Image (0x0028,0x0051) contains ATTN and DECAY and Decay Correction (0x0054,0x1102) is START {

if Units (0x0054,0x1001) are BQML {

if Series Date (0x0008,0x0021) and Time (0x0008,0x0031) are not after Acquisition Date (0x0008,0x0022) and Time (0x0008,0x0032)  {

scan Date and Time = Series Date and Time

}

else {	// may be post-processed series

scan Date and Time = GE private scan Date and Time (0x0009,0x100d,“GEMS_PETD_01”)

}

start Time = Radiopharmaceutical Start Time (0x0018,0x1072) in Radiopharmaceutical Information Sequence (0x0054,0x0016)

// start Date is not explicit … assume same as Series Date; but consider spanning midnight

decay Time = scan Time – start Time // seconds

half Life = Radionuclide Half Life (0x0018,0x1075) in Radiopharmaceutical Information Sequence (0x0054,0x0016)	// seconds

injected Dose = Radionuclide Total Dose (0x0018,0x1074) in Radiopharmaceutical Information Sequence (0x0054,0x0016)	// Bq

decayed Dose = injected Dose * pow (2, -decay Time / half Life)

weight = Patient's Weight (0x0010,0x1030)	// in kg

SUVbwScaleFactor = (weight * 1000 / decayed Dose)

}

else if Units (0x0054,0x1001) are CNTS {

SUVbwScaleFactor = Philips private scale factor (0x7053,0x1000,“ Philips PET Private Group”)

}

}

SUVbw = stored pixel value in Pixel Data (0x7FE0,0x0010) * Rescale Slope (0x0028,0x1053) * SUVbwScaleFactor	// g/ml