Mapping of Brain Regions involved in Motor Performance using Blood Oxygenation Level Dependent (BOLD) functional MRI as a Presurgical Assessment Tool.

Version 1.0

1 January 2012

Table of Contents

I. Executive Summary	3

II. Clinical Context and Claims	3

Indications, and Clinical Utility,	Error! Bookmark not defined.

Claim:  Measures and Endpoints	3

A. Eloquent Cortex Characterization

B. Characterization of Lesion and Cluster Margin

C. Relationship to Lesion

III. Profile Details	3

1. Subject Handling	5

2. Imaging Data Acquisition	5

3. Imaging Data Reconstruction	5

4. Image Analysis	Error! Bookmark not defined.

IV. Compliance	5

Acquisition Devices	5

Reconstruction Software	5

Software Analysis Tool	5

Performing Site	5

References	5

Appendices	5

Acknowledgements and Attributions	5

Background Information	5

Conventions and Definitions	5

Model-specific Instructions and Parameters	5

# I. Executive Summary

# Executive Summary

This Profile has been developed to provide a systematic approach for optimizing fMRI brain mapping for treatment planning prior to surgery or invasive treatment intervention. Whereas the primary purpose of this Profile development is for individual patient care, application of the best practice guidelines it creates has application to clinical trials as well.

fMRI can be used clinically as a biomarker for functionally eloquent brain tissue that might be at risk of damage from invasive procedures used to treat brain cancer or other focal pathologies (ref).  The clinical utility and professional acceptance of fMRI as a biomarker is dependent on the reproducibility and validity of brain activation patterns - the primary measure produced by fMRI exams and from which secondary quantitative measures are derived (ref). Current methodology is quite variable at all stages from exam administration, data acquisition, analysis and report of results, and can best be described by a model of integration across multiple data acquisition systems, MR and data analysis platforms.  To address reproducibility we take into account the degree to which variability in methodological approach (e.g., patient training, data acquisition methods, data analysis approaches and devices employed) impact the accuracy and specificity of readout measures derived. The current priority of the QIBA fMRI Technical Committee is to characterize the current state of the art and to identify sources of variability in methodology which contribute significantly to variance and negatively affect quantitative measures derived. If we can reduce the variability associated with methodological approach we can improve reproducibility and the quantitative value of fMRI as a biomarker.

Our initial studies provide quantitative measures of fMRI reproducibility that will be used in the statement of claims included as part of a QIBA fMRI profile currently under development. The same results will be provided to the scientific community at large in order to fill a critical gap in existing knowledge about fMRI reproducibility as assessed with quantitative measures that are particularly relevant for clinical use in pre-surgical planning

This QIBA Profile is expected to provide specifications that may be adopted by users as well as equipment developers (hardware and software devices) to meet targeted levels of clinical performance in identified settings. This profile makes claims about the precision with which hemodynamic response in eloquent cortex can be measured and displayed under a set of defined image acquisition, processing, and analysis conditions.

The intended audience of this document is:

- Technical staff of vendors planning to participate in the QIBA initiative

- Practicing clinicians at healthcare institutions considering appropriate specifications for acquiring equipment

- Experts involved in quantitative medical image analysis

- Anyone interested in the technical and clinical aspects of medical imaging

# II. Clinical Context and Claims

fMRI is used as a tool for pre-treatment planning in individual patients with brain lesions, including tumors, vascular malformation and epileptogenic foci. The presenting symptoms and location of the affected brain tissue determine the particular region or regions of the brain to be mapped and the behavioral paradigm(s) selected (e.g. motor task, language task). The change in BOLD signal (relative to a control condition) provides information about the brain region(s) involved in task performance and about the proximity of this eloquent cortex to brain site(s) to be treated. Endpoints that will influence treatment planning include risk assessment (impact of treatment on functioning cortex, e.g. surgical or radiation induced damage) and predictive value estimation (will damage to eloquent tissue result in a deficit). The goal of this profile is to specify the procedures and quantitative parameters under which BOLD fMRI is an accurate and reliable predictor of brain function, that is, as a valid imaging biomarker for medically meaningful changes in brain activity elicited by a particular task.

The clinical context sets out the utilities and endpoints for presurgical mapping cases and then proceeds to identify targeted levels of quality for named measurement read-outs that may be used in the relevant clinical indications.

## Claims characterizing reproducibility of BOLD response

- On a test-retest basis, fMRI can be performed reproducibly to a level such that the center of mass of activation of a focus of interest is within 5mm of itself, with at least 90% overlap of the activation clusters.

- On a test-retest basis, fMRI can be performed reproducibly to a level such that the relative magnitude of activation in homologous regions across hemispheres should be within 10%.

## Claims characterizing risk assessment (predictive value?)

- Quantitative measures of “risk” to eloquent brain structures… distance metrics… etc.

Compliance Levels for Measurement Read-outs

| Measurement or Categoric Result | Performance Levels Achieved under Bull's Eye Conditions |
| --- | --- |
| center of mass of activation of a focus of interest | If Activities are Performed at Acceptable level within 5mm of itself, with at least 90% overlap of the activation clusters If Activities are Performed at Target Level If Activities are Performed at Ideal Level |
| relative magnitude of activation in homologous regions across hemispheres | If Activities are Performed at Acceptable level within 10% If Activities are Performed at Target Level If Activities are Performed at Ideal Level |
| Quantitative measures of “risk” to eloquent brain structures… distance metrics… etc | If Activities are Performed at Acceptable level If Activities are Performed at Target Level If Activities are Performed at Ideal Level |

Utilities and Endpoints for Clinical Trials

**Describe one or more utilities or endpoints this Imaging Protocol could serve in a Clinical Trial. (e.g. to determine eligibility of potential subjects in the clinical trial; to triage eligible subjects into cohorts based on stage or severity of disease; to assess response to treatment; to establish the presence of progression for determining TTP, PFS, etc.; to monitor for adverse events; to establish a database for the development, optimization, and validation of imaging biomarkers, etc.)

# III. Profile Details

A technical description of tests for the biomarker, identifying measurement activities and read-outs, is provided:

The following sections provide details for the various components shall do in order to be in compliance:

Section 1, Subject Handling, is practiced by a Performing Site.

Section 2, Imaging Data Acquisition, is practiced by a Performing Site using an Acquisition Device.

Section 3, Imaging Data Reconstruction, is practiced by a Performing Site using Reconstruction Software.

Section 4, Image Analysis, is practiced by a Performing Site using one or more Software Analysis Tools.

The requirements included herein are intended to establish a baseline level of capabilities. Providing higher performance or advanced capabilities is both allowed and encouraged and the profile is not intended to be limiting in any way with respect to capabilities. The intention is not to dictate implementation details.

## 1. Subject Handling

1.1 Timing Relative to Index Intervention Activity

fMRI BOLD scanning for language, motor, and visual is performed and corresponding color maps generated prior to any interventional procedures.

1.2 Scheduling Ancillary Testing

If associated biopsy/resection (Neurosurgery) is expected to be performed during the same visit as the imaging procedure, it shall be described in the Trial Calendar.

1.3 Subject Preparation Prior to Arrival

Local standard of care shall be followed for MRI without contrast.  Assess the patient’s understanding and performance capability, alertness, cognitive ability, and behavioral capabilities in performing the necessary task(s) to complete the study. If necessary, train the patient with an identical or similar task using powerpoint or video presentations. If necessary, modify the paradigms to meet the capabilities of the patient and note limitations in patient record.

1.4 Subject Preparation Upon Arrival

1.5 Paradigm Library Specification

| Parameter | Compliance Levels |
| --- | --- |
| Language | Acceptable Target Ideal |
| Parameter | Compliance Levels |
| Motor | Acceptable Target Ideal |
| Parameter | Compliance Levels |
| Visual | Acceptable Target Ideal |
| Parameter | Compliance Levels |
| Memory | Acceptable Target Ideal |

1.6 Subject Ability Assessments

1.6.1 General Assessment

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

1.6.2 Language Assessment

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

1.6.3 Motor Assessment

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

1.6.4 Visual Function Assessment

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

1.6.5 Memory Assessment

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

1.7 Task Training prior to scanning

1.7.1. Language Task

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

1.7.2 Motor Task

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

1.7.3 Visual Task

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

1.7.4 Memory Task

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

1.8 Performance measurement of the tasks

| Parameter (Task) | Compliance Levels |
| --- | --- |
| Language | Acceptable We need to include some acceptable performance measures Target Ideal |
| Motor | Acceptable We need to include some acceptable performance measures Target Ideal |
| Visual | Acceptable We need to include some acceptable performance measures Target Ideal |
| Memory | Acceptable We need to include some acceptable performance measures Target Ideal |

1.9 Subject Instruction prior to scanning

Provide information to the patient regarding the flow of the exam (e.g. order of the tasks, what can be expected in terms of time for each paradigm administered). If the patient has never been in the MR, review what can be expected in terms of noise, discomfort, etc.

A quick review of the task is recommended to be sure that the patient is still familiar with what they will see or hear, and what they are asked to do during the task.

1.10 Subject Positioning in the scanner

Positioning of the patient should be consistent with local MRI head positioning procedures. There may be additional requirements to properly adjust the fMRI stimulus presentation devices (e.g.,adjustment of goggle or a mirror on or above the head coil to correctly adjust for visual acuity and ensure the entire visual field is visible. If necessary, additional stimulus response devices such as a MR compatible mouse or a trackball should be positioned such that the patient is able to operate the device without much hindrance. It is advisable to use foam padding to reduce head motion, and use head phones with foam ear plugs to reduce interference from scanner noise.

| Parameter | Compliance Levels |
| --- | --- |
| Subject Positioning | Acceptable Information Needed Target Ideal |

## 2. Imaging Data Acquisition

MRI scans for fMRI analysis will be performed on qualified equipment. It is recommended to use a field strength of 1.5 Tesla and higher with echo planar capabilities. Once the patient is positioned inside the scanner within the head coil it is good practice for the MRI technician to provide brief instruction to the patient about the task(s), as a reminder of what they are expected to do. It is recommended to have appropriate personnel present during the scan to meet CPT code requirements. The MRI scan starts with a localizer followed by T1 or T2 scans to cover the whole brain following the local imaging protocol. Following the anatomical image acquisition, a shim scan is followed by an fMRI BOLD series .The fMRI scan duration is defined by the paradigm design and the MR protocol should be configured and named with the same description used to name the stimulus paradigm.

Precise synchronization of stimulus presentation and image acquisition is highly recommended. The system to control the stimulus presentation can be a standalone workstation or PC with software for presenting stimulus paradigms or it can be software which is integrated within the MR console. As previously mentioned, the visual presentation of the stimulus can be displayed via MR compatible systems such as  goggle-based system, or a projector-based system (LCD Monitor, or Projector). The audio stimulus can be presented using audio delivery systems provided with the MR or a third party systems which are MR compatible and specifically designed for presenting stimuli in the MR environment. Monitoring task performance (direct observation of eye movement, finger/hand/foot movement etc) as well as recording patient responses (button box or other devices to monitor patient performance) is highly recommended. Frequent communication between the patient and technician between scan series to assess comfort and attention, and to provide intermittent instruction is required.

2.1 Technician Instructions

2.2 Stimulus Presentation specifications

2.2.1 Scan Synchronization/Triggering requirements

It is recommended to perform BOLD fMRI imaging with precise synchronization of the start of the scan and the start of the stimulus presentation using trigger pulses.

| Parameter | Compliance Levels |
| --- | --- |
| Timing / Triggers | Acceptable Shall use a standard time. Target Scanner triggers presentation software or vice versa. |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Actual Timing and Triggers shall be recorded. |

2.2.2 Visual Stimulus specifications

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.2.3 Auditory stimulus specifications

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.2.4 Memory stimulus specifications

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.3 Subject instruction during Acquisition

Patients are instructed to lay still, breathe normally and follow and pay utmost attention to the instructions on the stimulus presentation screen during the entire data acquisition.

2.4 Structural MRI scanning specifications

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.5 fMRI-EPI scanning specifications

2.5.1 Fmri Task scans

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.5.2 Breath hold scans (NVU)

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.6 Required Visualization/Monitoring

Visual monitoring of the patients during the performance of the task is recommended. This may aid in evaluating compliance of certain fMRI tasks such as motor tasks. It is also recommended to conduct an interview after the scan for patient compliance.

2.6.1 Monitor Visual Task performance

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.6.1 Monitor button box responses

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.6.3 Monitor Respiration

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.6.4 Monitor Eye movement

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.6.5 Monitor real time motion

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.6.6 Monitor Real time data analysis

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.7 Data Content & Structure

The BOLD T2* images are reconstructed on the scanner as individual images or as mosaics.  An fMRI series will typically consist of several measurement periods. Each individual measurement period will have a set of images corresponding to the anatomical coverage specified by the user (typically whole brain).  The total imaging time to acquire an fMRI series will depend on the repetition time (TR) and the number of measurement periods acquired throughout the series.

The following parameters describe what the acquired images shall contain/cover.

| Parameter | Compliance Levels |
| --- | --- |
| Anatomic Coverage | Acceptable Coverage of Area of interest Target Whole Brain |
| Field of View | Acceptable Coverage of Area of interest Target Whole brain |
| Scan Duration Motor Task; Language Task; | Acceptable 2 min Target xxxxx Ideal 3 min |
| Scan Plane (Image Orientation) | Acceptable Transverse or Axial Target Transverse or Axial |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Number of Measurement Periods; Actual Anatomic Coverage, Field of View, Scan Duration, and Scan Plane shall be recorded. |

2.8 Data Quality Requirements

It is highly recommended that the sites perform some or all of these quality assurances on their devices for improved and consistent data quality. They include routine SNR and fSNR measurements to test for signal and image quality, routine checks on the fMRI specific equipments such as the response buttons, projector, goggles, audio etc prior to the scan. Motion artifacts can significantly impede the quantitative fMRI outcome measures. Hence it is important to use head restrainers such as foam pads and provide reminders to the patient’s before the scan to reduce the motion inside the scanner while performing the test.

2.8.1 MRI Signal specifications (SNR, fSNR)

| Parameter | Compliance Levels |
| --- | --- |
| 1.Motion Artifact | Acceptable Minimal artifact (Potentially corrected prospectively) Target No artifact |
|  | Acceptable |
| 2. Spatial Resolution | Acceptable Target Ideal |
| 3. Noise | Acceptable |
| xxxxx | Acceptable Target |
| xxxxx | Acceptable Target Ideal |

2.9 Post Scan Requirements

2.9.1 Patient assessment

Info Needed

2.9.2 Real-time results evaluation

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

2.10 Staff requirements for fMRI

Info Needed

## 3. Post Acquisition Processing

The Post-acquisition processing and statistical analysis can be performed on the scanner provided software or on a standalone workstation. A variety of software’s and algorithms are available for this purpose (xxxxxxx). This section provides guidelines and recommends the following steps to obtain high quality and reliable color maps of the fMRI data.  The BOLD data is typically corrected for a low frequency signal drift, spatial smoothed to improve SNR, artifacts identified (manual or automatic) and corrected, corrected for slice timing (if event related design is used), motion identified and corrected, and coregistered with a T1 or T2 structural data. The registration parameters are typically saved in a file for later Q/A check to access the patient’s motion. A variety of statistical tools (GLM, non GLM, Xxxxx) methods can be used to perform statistical analysis to create color functional maps. These maps are later overlayed onto the structural data or 3D maps created for better visualization by the end users. Individual maps pertaining to different paradigms are created. These maps can be saved in DICOM, generic formats on the scanner or off the scanner.

It is highly recommend to generate a technical report that includes the summary of the imaging procedure, patient performance of the task, qualitative and quantitative summary of the head motion, subjective assessments of artifacts and outliers, assessment of the data alignment (functional vs structural), pre, during and post scan evaluation of the patient, and neurovascular uncoupling of the patient.

These parameters describe general characteristics of the reconstruction among other specifications in the post acquisition processing:

3.1 Image processing computer specifications

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

3.2 Image processing software specifications

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

3.3 BOLD Analysis pipeline specifications

3.3.1 	Anatomical Image Segmentation

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

3.3.2	Low frequency drift correction

| Parameter | Compliance Levels |
| --- | --- |
| Drift Specs | Acceptable Target Ideal |

3.3.3	Spatial filtering

| Parameter | Compliance Levels |
| --- | --- |
| Filter Specs | Acceptable Target Ideal |

3.3.4	Artifact Identification/Removal

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

    - Slice timing correction

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Target Ideal |

    - Motion correction

| Parameter | Compliance Levels |
| --- | --- |
| Motion Correction Parameters | Acceptable Target Ideal |

3.3.7	Coregistration

| Parameter | Compliance Levels |
| --- | --- |
| Color Map Specs | Acceptable Target Ideal |

    - Statistical map generation

| Parameter | Compliance Levels |
| --- | --- |
| Color Map Specs | Acceptable Target Ideal |

    - Statistical map threshold

| Parameter | Compliance Levels |
| --- | --- |
| Statistical Thresholds | Acceptable Target Ideal |

3.3.10	Functional map generation

| xxxxx | Acceptable Target |
| --- | --- |
| xxxxx | Acceptable Target Ideal |

3.4 Color Map Specifications

| Parameter | Compliance Levels |
| --- | --- |
| Color Map Specs | Acceptable Target Ideal |

3.5 Quality Control Requirements

    - Registration parameters

| Parameter | Compliance Levels |
| --- | --- |
| Color Map Specs | Acceptable Target Ideal |

3.6 Technical Report Specifications

Info Needed

3.6.1	Color Images

3.6.2	Imaging procedure summary

3.6.3 	Patient motion summary

    - Artifact summary

    - Patient evaluation summary

3.6.6 	NVU summary

## Image Analysis & Clinical Interpretation

4.1 Clinical Report Content

4.2. Clinical history review

4.3. Color overlay characterization

4.3.1 Review of color overlays

4.3.2 Gyral/Sulcal location

4.3.3 Gyral/Sulcal distance

4.3.4 Lesion/Cluster margin characterization

4.3.5 Characterization of Artifacts

4.3.6 Confidence Index

4.3.7 Documentation for CPT requirements

4.4 Clinical Report Structure

4.4.1 Demographic information

4.4.2 Indications

4.4.3 Techniques

4.4.4 Structural imaging findings

4.4.5 Functional imaging findings

4.4.6 Annotated images

4.4.7 Clinical Impression

We need input to this section from a clinician. What you see below is the CT Template. We could use this as a guide.

Each lesion shall be characterized as described in this section.  Lesions of interest include: a) brain tumors; b) XXXXXXXX small to medium pulmonary Masses surrounded by air and/or with adjacent normal and abnormal (non-neoplastic) anatomic structures; c) large pulmonary masses surrounded by air and/or with adjacent normal and abnormal (non-neoplastic) anatomic structures and/or confluent with mediastinum, chest wall, and diaphragm.

Procedures for segmenting or excluding tissue types and fluid, blood, necrotic debris within a mass are not described by this protocol, but may be implemented when technically feasible, in addition to measuring the entire volume within the outer tumor margin.

4.1 Methods to Be Used

Each lesion shall be characterized by determining the boundary of the lesion (referred to as segmentation), then computing the volume of the segmented lesion.  Segmentation may be performed automatically by a software algorithm, manually by a human observer, or semi-automatically by an algorithm working with human guidance/intervention.  The volume of the segmented region is then computed automatically.

| Parameter | Compliance Levels |
| --- | --- |
| Common Lesion Selection | Acceptable No requirement Target The software shall allow a common set of lesions to be designated for measurement, which are then subsequently measured by all readers Ideal The software shall detect and measure all measurable lesions automatically without the need for human intervention or multiple readers |
| Lesion Volume | Acceptable Shall be calculated as the sum of all the voxels within the boundaries of a discrete tumor mass on all the tomographic slices on which it is visible, regardless of its irregular shape. Target Shall be calculated without regard to spatial sampling loss (i.e., accounting by some means of interpolation for volume averaging due to non-isotropic voxel reconstruction and finite sampling). |
| Change Assessment Workflow | Acceptable Shall be performed as “locked sequential read”. |
| Multiple Lesions | Acceptable The software shall allow multiple lesions to be measured, and each measured lesion to be associated with a human-readable identifier that can be used for correlation across time points |
| Sum of Target Lesion Volumes | Acceptable A value computed by adding up all of the target lesion volumes calculated using Acceptable approach above shall be computed. Target A value computed by adding up all of the target lesion volumes calculated using Target approach above shall be computed. Ideal A value computed by adding up all of the target lesion volumes calculated using Ideal approach above shall be computed. |

For semi-automated or automated segmentation, the analysis software shall segment (based on a starting seed point/stroke/ROI) various types of tumors on CT images. The following further requirements are placed on image analysis software:

| Parameter | Compliance Levels |
| --- | --- |
| Boundary segmentation | Acceptable With many (> 50%) lesions requiring reader correction Target With few (< 10%) lesions requiring reader correction Ideal Fully automatically without reader correction |
| Automatically computed read-outs | Acceptable Automatic computation of volume of the segmented tumor shall be provided. Target Error margins for each measurement Provide a HU-histogram of the segmented voxels shall be provided. |
| Image Header Recording | Acceptable Software shall record in (and reload for review from) lesion segmentation boundary and volumetric measurement as well as metadata about reader identity, date and time and purpose of measurement. Target Software shall record in (and reload for review from) lesion segmentation boundary and volumetric measurement as well as metadata in standard formats including one or more of the following output formats: DICOM Presentation State, DICOM Structured Report; DICOM RT Structure Set; DICOM raster or surface segmentation. Ideal Software shall record in (and reload for review from) ALL of the Target formats. |

4.2 Required Characteristics of Resulting Data

It is expected that automated boundary detection algorithms will place segmentation edges with greater precision, accuracy and speed than an operator can draw by hand with a pointing device. The performance of the algorithms will, however, depend on the characteristics of the lesions may be challenged by complex tumors. Operator assisted semi-automatic segmentation shall produce at least the same level of intra- and inter-rater reliability for the volume measurements of each target lesion as manual segmentation.

| Parameter | Compliance Levels |
| --- | --- |
| Read-outs as described in Methods section | Acceptable Same precision, accuracy and speed than an operator can draw by hand with a pointing device shall be demonstrated. Target Greater precision, accuracy and speed than an operator can draw by hand with a pointing device shall be demonstrated. |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Annotation and Markup metadata | Acceptable Actual model-specific Analysis Software set-up and configuration parameters utilized to achieve compliance with these metrics shall be recorded. |

# 5. Storage and Distribution

5.1	Data Storage specification

| Parameter | Compliance Levels |
| --- | --- |
| Storage Specs | Acceptable Target Ideal |

  - Data sharing specification

| Parameter | Compliance Levels |
| --- | --- |
| Sharing Specs | Acceptable Target Ideal |

# IV. Compliance

## Acquisition Devices

Compliance to specifications as set out in the Image Acquisition section above.  Additionally, compliant Acquisition Devices shall provide means to record the information identified in the Subject Handling section as means to document compliance of the Performing Site to the specifications noted there.

## Reconstruction Software

Compliance to specifications as set out in the Image Reconstruction section above.  Additionally, compliant Reconstruction Software shall propagate the information collected at the prior Subject Handling and Imaging Acquisition stages and extend it with those items noted in the Reconstruction section.  See the compliance procedure notes associated with Acquisition Devices above for procedural assistance to identify Model Specific Parameters for Reconstruction Software.

## Software Analysis Tool

Compliance to specifications as set out in the Image Analysis section above.  Additionally, compliant Software Analysis Tools shall propagate the information collected at the prior Subject Handling, Imaging Acquisition, and Imaging Reconstruction stages and extend it with those items noted in the Analysis section

## Performing Site

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large patient population under consideration.  For imaging it is important to consider the availability of:

- appropriate imaging equipment and quality control processes,

- appropriate injector equipment and contrast media,

- experienced CT technologists for the imaging procedure, and

- processes that assure imaging protocol compliant image generation at the correct point in time.

A protocol specific calibration and QA program shall be designed consistent with the goals of the clinical trial. This program shall include (a) elements to verify that sites are performing the specified protocol correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address issues relating to both radiation dose and image quality (which may include issues relating to water calibration, uniformity, noise, spatial resolution -in the axial plane-, reconstructed slice thickness z-axis resolution, contrast scale, CT number calibration and others). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.

# References

[] Moertel CG, Hanley JA. The effect of measuring error on the results of therapeutic trials in advanced disease. Disease 1976; 38: 388-394.

[2] Quivey JM, Castro JR, Chen GT, Moss A, Marks WM. Computerized tomography in the quantitative assessment of tumour response. Br J Disease Suppl 1980; 4:30-34.

[3] Munzenrider JE, Pilepich M, Rene-Ferrero JB, Tchakarova I, Carter BL. Use of body scanner in radiotherapy treatment planning. Disease 1977; 40:170-179.

[4] Wormanns, D., Kohl, G., Klotz, E., Marheine, A., Beyer, F., Heindel, W., and Diederich, S. Volumetric measurements of pulmonary nodules at multi-row detector CT: In vivo reproducibility. Eur Radiol 14: 86–92, 2004.

[5] Kostis WJ, Yankelevitz DF, Reeves AP, Fluture SC, Henschke CI, Small Pulmonary Nodules: Reproducibility of Three-dimensional Volumetric Measurement and Estimation of Time to Follow-up CT, Radiology, Volume 231 Number 2, 2004.

[6] Revel M-P, Lefort C, Bissery A, Bienvenu M, Aycard L, Chatellier G, Frija G, Pulmonary Nodules: Preliminary Experience with Three-dimensional Evaluation, Radiology  May 2004.

[7] Marten K, Auer F, Schmidt S, Kohl G, Rummeny EJ, Engelke C, Inadequacy of manual measurements compared to automated CT volumetry in assessment of treatment response of pulmonary metastases using RECIST criteria, Eur Radiol (2006) 16: 781–790.

[8] Goodman, L.R., Gulsun, M., Washington, L., Nagy, P.G., and Piacsek, K.L. Inherent variability of CT lung nodule measurements in vivo using semiautomated volumetric measurements. AJR Am J Roentgenol 186: 989–994, 2006.

[9] Gietema HA, Schaefer-Prokop CM, Mali W, Groenewegen G, Prokop M, Pulmonary Nodules: InterscanVariability of Semiautomated Volume Measurements with Multisection CT— Influence of Inspiration Level, Nodule Size, and Segmentation Performance, Radiology: Volume 245: Number 3 December 2007.

[10] Wang Y, van Klaveren RJ, van der Zaag–Loonen HJ, de Bock GH, Gietema HA, Xu DM, Leusveld ALM, de Koning HJ, Scholten ET, Verschakelen J, Prokop M, Oudkerk M, Effect of Nodule Characteristics on Variability of Semiautomated Volume Measurements in Pulmonary Nodules Detected in a Lung Cancer Screening Program, Radiology: Volume 248: Number 2—August 2008.

[11] Zhao, B., Schwartz, L.H., and Larson, S.M. Imaging surrogates of tumor response to therapy: anatomic and functional biomarkers. J Nucl Med 50: 239–249, 2009.

[12] Hein, P.A., Romano, V.C., Rogalla, P., Klessen, C., Lembcke, A., Dicken, V., Bornemann, L., and Bauknecht, H.C. Linear and volume measurements of pulmonary nodules at different CT dose levels: Intrascan and interscan analysis. Rofo 181: 24–31, 2009.

[13] Mozley PD, Schwartz LH, Bendtsen C, Zhao B, Petrick N, Buckler AJ. Change in lung tumor volume as a biomarker of treatment response: A critical review of the evidence. Annals Oncology; doi:10.1093/annonc/mdq051, March 2010.

[14] Petrou M, Quint LE, Nan B, Baker LH. Pulmonary nodule volumetric measurement variability as a function of CT slice thickness and nodule morphology. Am J Radiol 2007; 188:306-312.

[15] Bogot NR, Kazerooni EA, Kelly AM, Quint LE, Desjardins B, Nan B. Interobserver and intraobserver variability in the assessment of pulmonary nodule size on CT using film and computer display methods. Acad Radiol 2005; 12:948–956.

[16] Erasmus JJ, Gladish GW, Broemeling L, et al. Interobserver and intraobserver variability in measurement of non-small-cell carcinoma lung lesions: Implications for assessment of tumor response. J Clin Oncol 2003; 21:2574–2582.

[17] Winer-Muram HT, Jennings SG, Meyer CA, et al. Effect of varying CT section width on volumetric measurement of lung tumors and application of compensatory equations. Radiology 2003; 229:184-194.

[18] Buckler AJ, Mozley PD, Schwartz L, et al. Volumetric CT in lung disease: An example for the qualification of imaging as a biomarker. Acad Radiol 2010; 17:107-115.

[19] AMERICAN COLLEGE OF RADIOLOGY IMAGING NETWORK, ACRIN 6678, FDG-PET/CT as a Predictive Marker of Tumor Response and Patient Outcome: Prospective Validation in Non-small Cell Lung Cancer, August 13, 2010.

[20] Miller AB, Hoogstraten B, Staquet M, Winkler A. Reporting results of cancer treatment. Cancer 1981;47:207-214.

[21] Eisenhauer EA, Therasse P, Bogaerts J, et al. New response evaluation criteria in solid tumors: Revised RECIST guideline (version 1.1). Eur J Cancer 2009;45:228-247.

[22] McNitt-Gray MF. AAPM/RSNA Physics Tutorial for Residents: Topics in CT. Radiation dose in CT. Radiographics 2002;22:1541-1553.

[23] Xie L, O'Sullivan J, Williamson J, Politte D, Whiting B,  TU‐FF‐A4‐02: Impact of Sinogram Modeling Inaccuracies On Image Quality in X‐Ray CT Imaging Using the Alternating Minimization Algorithm, Med. Phys. 34, 2571 (2007); doi:10.1118/1.2761438.

[24] Moertel CG, Hanley JA. The effect of measuring error on the results of therapeutic trials in advanced cancer. Cancer 38:388-94, 1976.

[25] Lavin PT, Flowerdew G: Studies in variation associated with the measurement of solid tumors. Cancer  46:1286-1290, 1980.

[26] Eisenhauera EA, Therasseb P, Bogaertsc J, et a. New response evaluation criteria in solid tumours: Revised RECIST guideline (version 1.1).  Eur J Cancer 2009; 45: 228-247.

# Appendices

## Acknowledgements and Attributions

This imaging protocol is proffered by the Radiological Society of North America (RSNA) Quantitative Imaging Biomarker Alliance (QIBA) Volumetric Computed Tomography (v-CT) Technical Committee. The v-CT technical committee is composed of scientists representing the imaging device manufacturers, image analysis software developers, image analysis laboratories, biopharmaceutical industry, academia, government research organizations, professional societies, and regulatory agencies, among others. All work is classified as pre-competitive. A more detailed description of the v-CT group and its work can be found at the following web link: http://qibawiki.rsna.org/index.php?title=Volumetric_CT.

The Volumetric CT Technical Committee (in alphabetical order):

- Athelogou, M. Definiens AG

- Avila, R. Kitware, Inc.

- Beaumont, H. Median Technologies

- Borradaile, K. Core Lab Partners

- Buckler, A. BBMSC

- Clunie, D. Core Lab Partners

- Cole, P. Imagepace

- Dorfman, G. Weill Cornell Medical College

- Fenimore, C. Nat Inst Standards & Technology

- Ford, R. Princeton Radiology Associates.

- Garg, K. University of Colorado

- Gottlieb, R. Roswell Park Cancer Center

- Gustafson, D. Intio, Inc.

- Hayes, W. Bristol Myers Squibb

- Hillman, B. Metrix, Inc.

- Judy, P. Brigham and Women’s Hospital

- Kim, HG. University of California Los Angeles

- Kohl, G. Siemens AG

- Lehner, O. Definiens AG

- Lu, J. Nat Inst Standards & Technology

- McNitt-Gray, M. University California Los Angeles

- Mozley, PD. Merck & Co Inc.

- Mulshine, JL. Rush

- Nicholson, D. Definiens AG

- O'Donnell, K. Toshiba

- O'Neal, M. Core Lab Partners

- Petrick, N. US Food and Drug Administration

- Reeves, A. Cornell University

- Richard, S. Duke University

- Rong, Y. Perceptive Informatics, Inc.

- Schwartz, LH. Columbia University

- Saiprasad, G. University of Maryland

- Samei, E. Duke University

- Siegel, E. University of Maryland

- Sullivan, DC. RSNA Science Advisor and Duke University

- Thorn, M. Siemens AG

- Yankellivitz, D. Mt. Sinai School of Medicine

- Yoshida, H. Harvard MGH

- Zhao, B. Columbia University

The Volumetric CT Technical Committee is deeply grateful for the support and technical assistance provided by the staff of the Radiological Society of North America.

## Background Information

QIBA

The Quantitative Imaging Biomarker Alliance (QIBA) is an initiative to promote the use of standards to reduce variability and improve performance of quantitative imaging in medicine. QIBA provides a forum for volunteer committees of care providers, medical physicists, imaging innovators in the device and software industry, pharmaceutical companies, and other stakeholders in several clinical and operational domains to reach consensus on standards-based solutions to critical quantification issues. QIBA publishes the specifications they produce (called QIBA profiles), first to gather public comment and then for field test by vendors and users.

QIBA envisions providing a process for developers to test their implementations of QIBA profiles through a compliance mechanism. After a committee determines that a profile has undergone sufficient successful testing and deployment in real-world care settings, it is released for use.  Purchasers can specify conformance with appropriate QIBA profiles as a requirement in requests for proposal. Vendors who have successfully implemented QIBA profiles in their products can publish conformance statements (called QIBA Compliance Statements) represented as an appendix called “Model-specific Parameters.”  General information about QIBA, including its governance structure, sponsorship, member organizations and work process, is available at http://qibawiki.rsna.org/index.php?title=Main_Page.

CT Volumetry for Cancer Response Assessment

Anatomic imaging using computed tomography (CT) has been historically used to assess tumor burden and to determine tumor response (or progression) to treatment based on uni-dimensional or bi-dimensional measurements. The original WHO response criteria were based on bi-dimensional measurements of the tumor and defined response as a decrease of the sum of the product of the longest perpendicular diameters of measured lesions by at least 50%. The rationale for using a 50% threshold value for definition of response was based on data evaluating the reproducibility of measurements of tumor size by palpation and on planar chest x-rays [24][25]. The more recent RECIST criteria introduced by the National Cancer Institute (NCI) and the European Organisation for Research and Treatment of Cancer (EORTC) standardized imaging techniques for anatomic response assessment by specifying minimum size thresholds for measurable lesions and considered other imaging modalities beyond CT. As well, the RECIST criteria replace longest bi-directional diameters with longest uni-dimensional diameter as the representation of a measured lesion [26]. RECIST defines response as a 30% decrease of the largest diameter of the tumor. For a spherical lesion, this is equivalent to a 50% decrease of the product of two diameters. Current response criteria were designed to ensure a standardized classification of tumor shrinkage after completion of therapy. They have not been developed on the basis of clinical trials correlating tumor shrinkage with patient outcome.

Technological advances in signal processing and the engineering of multi-detector row computed tomography (MDCT) devices have resulted in the ability to acquire high-resolution images rapidly, resulting in volumetric scanning of anatomic regions in a single breath-hold. Volume measurements may be a more sensitive technique for detecting longitudinal changes in tumor masses than reliance on linear tumor diameters as defined by RECIST. Comparative analyses in the context of real clinical trial data have found volume measurements to be more reliable and often more sensitive to longitudinal changes in response than the use of diameters in RECIST. As a result of this increased detection sensitivity and reliability, volume measurements may improve the predictability of clinical outcomes during therapy compared with RECIST. Volume measurements could also benefit patients who need alternative treatments when their diseases stops responding to their current regimens.

The rationale for volumetric approaches to accessing assessing longitudinal changes in tumor burden is multi-factorial. First, most cancers may grow and regress irregularly in three dimensions. Measurements obtained in the transverse plane fail to account for growth or regression in the longitudinal axis, whereas volumetric measurements incorporate changes in all dimensions. Secondly, changes in volume are less subject to either reader error or inter-scan variations. For example, partial response using the RECIST criteria requires a greater than 30% decrease in tumor diameter, which corresponds to greater than 50% reduction in volume of tumor. If one assumes a 21 mm diameter lesion (of 4850 mm3 volume), partial response would result require that the tumor shrink to a in a diameter of less than 158 mm, but which would correspond to a decrease in volume all the way down to 17702145 mm3. The much greater absolute magnitude of volumetric changes is potentially less prone to measurement error than changes in diameter, particularly if the lesions are irregularly shaped or spiculated. As a result of the observed increased sensitivity and reproducibility, volume measurements may be more suited than uni-dimensional measurements to identify early changes in patients undergoing treatment.

## Conventions and Definitions

Acquisition vs. Analysis vs. Interpretation: This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be computational steps that transform the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge. (The transformation of knowledge into wisdom is beyond the scope of this document.)

Bulls-eye Compliance Levels Acquisition parameter values and some other requirements in this protocol are specified using a “bulls-eye” approach. Three rings are considered from widest to narrowest with the following semantics:

ACCEPTABLE: failing to meet this specification will result in data that is likely unacceptable for the intended use of this protocol.

TARGET: meeting this specification is considered to be achievable with reasonable effort and equipment and is expected to provide better results than meeting the ACCEPTABLE specification.

IDEAL: meeting this specification may require unusual effort or equipment, but is expected to provide better results than meeting the TARGET.

An ACCEPTABLE value will always be provided for each parameter. When there is no reason to expect better results (e.g. in terms of higher image quality, greater consistency, lower dose, etc.), TARGET and IDEAL values are not provided.

Some protocols may need sites that perform at higher compliance levels do so consistently, so sites may be requested to declare their “level of compliance”. If a site declares they will operate at the TARGET level, they must achieve the TARGET specification whenever it is provided and the ACCEPTABLE specification when a TARGET specification is not provided. Similarly, if they declare IDEAL, they must achieve the IDEAL specification whenever it is provided, the TARGET specification where no IDEAL level is specified, and the ACCEPTABLE level for the rest.

Other Definitions:

Image Analysis, Image Review, and/or Read: Procedures and processes that culminate in the generation of imaging outcome measures, such tumor response criteria. Reviews can be performed for eligibility, safety or efficacy. The review paradigm may be context specific and dependent on the specific aims of a trial, the imaging technologies in play, and the stage of drug development, among other parameters.

Image Header: The Image Header is that part of the file or dataset containing the image other than the pixel data itself

Imaging Phantoms: Devices used for periodic testing and standardization of image acquisition. This testing must be site specific and equipment specific and conducted prior to the beginning of a trial (baseline), periodically during the trial and at the end of the trial.

Intra-Rater Variability is the variability in the interpretation of a set of images by the same reader after an adequate period of time inserted to reduce recall bias.

Inter-Rater Variability is the variability in the interpretation of a set of images by the different readers.

A Time Point is a discrete period during the course of a clinical trial when groups of imaging exams or clinical exams are scheduled as defined in the study protocol.

## Model-specific Instructions and Parameters

Compliance with a profile involves meeting a variety of requirements of which operating by these parameters is just one. To determine if a product (and a specific model/version of that product) is compliant, please refer to the Compliance section above.

Sites using models listed here are encouraged to consider using these parameters for both simplicity and consistency. Sites using models not listed here may be able to devise their own settings that result in data meeting the requirements but this is outside the formal scope of QIBA compliance.

In some cases, parameter sets may be available as an electronic file for direct implementation on the imaging platform.

Table G.1: Acquisition Device Model-specific Parameters Demonstrated to Achieve Compliance

IMPORTANT NOTE with respect to this example table: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  These settings were determined by the team in the 1C study as an example of how it could be done but more strict attention to all parameters identified in the Profile are necessary in order for a company to claim any particular model is compliant.  That said, we appreciate the good will and help that the vendors represented here have provided in this early phase of QIBA.

| Acquisition Device | Product Setting to Achieve Compliance Levels |
| --- | --- |
| GE Discovery HD750 sct3 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.625 Gantry Rotation Time in seconds 1 mA 120 Pitch 0.984 Scan FoV Large Body (500mm) |
| Philips Brilliance 16 IDT mx8000 | kVp 120 Number of Data Channels (N) 16 Width of Each Data Channel (T, in mm) 0.75 Gantry Rotation Time in seconds 0.75 Effective mAs 50 Pitch 1.0 Scan FoV 500 |
| Philips Brilliance 64 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.625 Gantry Rotation Time in seconds 0.5 Effective mAs 70 Pitch 0.798 Scan FoV 500 |
| Siemens Sensation 64 | kVp 120 Collimation (on Operator Console) 64 x 0.6 (Z-flying focal spot) Gantry Rotation Time in seconds 0.5 Effective mAs 100 Pitch 1.0 Scan FoV 500 |
| Toshiba Aquilion 64 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.5 Gantry Rotation Time in seconds 0.5 mA TBD Pitch .828 Scan FoV Medium and Large |

Table G.2: Reconstruction Software Model-specific Parameters Demonstrated to Achieve Compliance

IMPORTANT NOTE: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  These settings were determined by the team in the 1C study as an example of how it could be done but more strict attention to all parameters identified in the Profile are necessary in order for a company to claim any particular model is compliant.  That said, we appreciate the good will and help that the vendors represented here have provided in this early phase of QIBA.

| Reconstruction Software | Product Setting to Achieve Compliance Levels |
| --- | --- |
| GE Discovery HD750 sct3 | Reconstructed Slice Width, mm 1.25 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel STD |
| Philips Brilliance 16 IDT mx8000 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm (contiguous) Display FOV, mm 350 Recon kernel B |
| Philips Brilliance 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm (contiguous) Display FOV, mm 350 Recon kernel B |
| Siemens Sensation 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel B30 |
| Toshiba Aquilion 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm Display FOV, mm TBD Recon kernel FC12 |

Table G.3: Image Analysis Software Model-specific Parameters Demonstrated to Achieve Compliance

IMPORTANT NOTE: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  In particular, the following example table only has placeholders for these example products which need to be replaced with product model-specific settings in order to claim compliance.

| Image Analysis Software | Product Setting to Achieve Compliance Levels |
| --- | --- |
| Siemens LunCARE | a <settings to achieve…> b <settings to achieve…> c <settings to achieve…> d <settings to achieve…> |
| GE Lung VCAR | e <settings to achieve…> f <settings to achieve…> g <settings to achieve…> h <settings to achieve…> |
| R2 ImageChecker CT Lung System | i <settings to achieve…> j <settings to achieve…> k <settings to achieve…> l <settings to achieve…> |
| Definiens (name specific product) | m <settings to achieve…> n <settings to achieve…> o <settings to achieve…> p <settings to achieve…> |
| Median (name specific product) | q <settings to achieve…> r <settings to achieve…> s <settings to achieve…> t <settings to achieve…> |
| Intio (name specific product) | u <settings to achieve…> v <settings to achieve…> w <settings to achieve…> x <settings to achieve…> |