Protocol Template

|  | QIBA VolCT Lung | ACRIN 6678 | NLST |  |
| --- | --- | --- | --- | --- |
|  | Ideal | Target | Acceptable | Ideal | Target | Acceptable | Ideal | Target | Acceptable |  |
| Image |  |  |  |  |  |  |  |  |  |  |
| Slice Width | <=1mm | 1.0-2.5mm | <=5mm |  | 1.0-1.5mm | 1.0-1.5mm |  | 2.0-3.2mm | --- |  |
| Slice Interval | Contig. or 20% overlap | Contig. or 20% overlap | Contiguous |  | 0-20% overlap | 0-20% overlap |  | 10-35% overlap | --- |  |
| Isotropic Voxels | Yes | Yes | Not |  | --- | Not |  | --- | --- |  |
| Field of View:Voxel Size | Rib-to-rib: 0.55-.75mm | Outer Thorax: 0.7-.8mm | Compl. Thorax: <=1.0mm |  | 0.55-.75mm | 0.75mm |  | --- | --- |  |
| Scan Plane | 0 azimuth | 0 azimuth | ?? |  | --- | --- |  | --- | --- |  |
| Motion Artifact | None | None | “Minimal” ?? |  | None | None |  | --- | --- |  |
| Noise Level | “minimal”? | “low”? | “predictable”? |  | --- | --- |  | --- | --- |  |
| Noise Level (note Slice Width) | <40 HU std. dev. in 20cm water phantom | <40 HU std. dev. in 20cm water phantom | <40 HU std. dev. in 20cm water phantom |  | --- | --- |  | --- | --- |  |
| Spatial Resolution | 7-8 lp/cm | 6-8 lp/cm | <=8 lp/cm |  | --- | --- |  | --- | --- |  |
| Technique |  |  |  |  |  |  |  |  |  |  |
| KVP | 120 | 110-130 | 110-140 |  | 120 | --- |  | 120 | --- |  |
| Effective mAs (med.) | 80-120 | 60-200 | 40-350 |  | 95-135 | --- |  | 20-30 | --- |  |
| Effective mAs (large) |  |  |  |  | 220-310 | --- |  | 40-60 | --- |  |
| CTDIvol mGy |  |  |  |  | --- | --- |  | 1.9-6.2 | --- |  |
| Rotation Speed |  |  |  |  | 0.5-0.6 sec | --- |  | 0.5 sec | 0.8 sec |  |
| Collimation Width | 20-40mm | 10-80mm | 5-160mm |  | ?? | --- |  | ??? | --- |  |
| # of Channels | >=64 | >=16 | >=1 |  | 8/16/40/64 | --- |  | 4/8/16/32/64 | --- |  |
| Mode |  |  |  |  | --- | --- |  | --- | --- |  |
| Table Speed mm/sec. |  |  |  |  | --- | --- |  | 16-96 | --- |  |
| Recon Kernel Type | Slight. Enhance | Standard to Enhancing | Soft to Overenhancing |  | --- | --- |  | --- | --- |  |
| Recon Kernel Name |  |  |  |  | STD/B/B30/FC10 | --- |  | STD/B30/B/C/FC10 | --- |  |
| Scanner Model |  |  |  |  |  |  |  |  |  |  |
| Patient Prep. |  |  |  |  |  |  |  |  |  |  |
| Contrast |  |  |  |  | None | None |  |  | --- |  |
| Creatinine Clearance |  |  |  |  | --- | --- |  |  | --- |  |
| Bladder State |  | Empty | Any |  | --- | --- |  |  | --- |  |
| Positioning |  | Supine/ArmsUp/HeadFirst | Any (Consistent w prior scan) |  | --- | --- |  |  | --- |  |
| Breath Hold |  | Single Breath Hold at Full Inspiration |  |  | --- | --- |  |  | --- |  |

Acceptable – if the data does not meet these values it is disqualified

Target – sites should strive to meet these values

Ideal – while it is unreasonable to expect all participating sites to meet these values, results would be improved if they can be achieved.

NLST - <Purpose>