QIBA DCE-MRI Committee

Discussion on Phantom Development

Wednesday, June 18, 2008

1pm-2pm CDT

Call Summary (Final recommendation listed at bottom)

In attendance:

Michael Buonocore, MD, PhD (Moderator)

Jeffrey Evelhoch, PhD

Luna Hilaire, PhD

David Purdy, PhD

Mitchell Schnall, MD, PhD

Stephen Zammit, PhD (Ohio State?)

Tracy Schmidt, MS (RSNA)

Joe Koudelik (RSNA)

Modified ADNI phantom developed for IRAT study discussed

Adapt IRAT phantom for this QIBA project

Review of phantom features

- IRAT phantom design provides

  - 8 3-cm spheres in separate locations with different contrast concentrations

  - Signal levels in spheres reduced by 1/3rd (using D20) to match DCE-MRI studies

  - Signal levels will be dependent on contrast agent concentration

- Generates contrast curves at eight separate locations

- Geometric distortion correction achieved using small spheres (analysis by Image Owl)

- Background signal used for correction of phased-array RF coil sensitivity

  - Background fluid shows RF coil sensitivity

  - Images acquired with Body coil can be used to derive RF coil sensitivity

- Provides dielectric loading consistent with in-vivo abdominal imaging

  - Fluid-filled rectangular block used to provide additional load to RF transmit coil

- Phantom setup eliminates variability of surface coil placement around phantom

  - In human imaging, RF coil placement is large source of signal variability

  - RF coil supported by wood frame, phantom lock ensures minimal variability

- Gibbs artifact present in 3 cm spheres in phantom

  - 3 cm minimum diameter acceptable for quantitative measurements

Four sites in current IRAT study will run DCE-MRI protocol:

- Ohio State University

- Memorial Sloan-Kettering

- Johns Hopkins

- NIH

Loading Discussion

- Benefits of gel vs. liquid in phantoms

  - Gels don’t leak, but, difficult to handle and not stable, not liked by Phantom Labs

  - Solutions made by Phantom Labs reduce variability in measurements

  - Stability of CuSO4 (Copper Sulfate) in phantom solutions

  - Redox state of CuSO4 changes if impure water used

  - CuSO4 change of state possible within 1 year

  - Lack of hydration issues

  - CuSO4 vs. signal levels

  - With Phantom Laboratory solutions, no problems with CuSO4 solutions reported

What other phantoms are available to compare?

- VirtualScopics phantoms

  - Separate geometric distortion phantom and T1 phantom

  - 100 different contrast agent concentrations

    - limited to one plane (single slice evaluation)

    - T1 measurements, but no comparison in different physical regions

    - Cost unknown - “Information content per dollar” should be considered

      - Spatial variation is minimal when RF body coil is used for reception

Features not provided by current IRAT phantom

- Not gel-based, fluid does not model tissue as well as gel

- No lesion modeling (i.e. regions with characteristics of normal and abnormal tissue)

- No physical perfusion in phantom

- Does not match shape of abdomen (possibly more accepted if shaped like abdomen)

- T2 values not controlled (controlling both T1and T2 difficult, T2 less important)

- No test of proton density variation

Use of D20 to reduce SNR to match in-vivo scans, but sacrifices SNR

- Current phantom uses D20 to reduce signal intensity by 1/3rd in 3 cm spheres.

- Use of D20 insures that system receiver gains of in-vitro and in-vivo scans are the same

- For QIBA phantom, recommended that D2O from phantom be replaced with H20

Other Comments

- Lack of test-retest data by RSNA’09 not relevant at this time

- One phantom, using one set of solutions produced by one company, would make a well controlled study.

Weekly Update Conference Calls

Weekly calls scheduled for Wednesday, 1-2pm CDT until satisfied with phantom design

Final recommendations:

1. Adopt modified IRAT phantom for the QIBA study

2. For QIBA phantom, H20/D20 mixture in 3 cm spheres will be replaced with pure H20