## Instructions
IMPORTANT: Do not edit this template spreadsheet.
FIRST: Use File->Make a copy... to copy this template into a new file for your comments
UPDATE the file name (upper left corner of your new sheet window) which says "Copy of ..." to include your Profile Name and remove Copy of
NOW you can edit your copied document.
See http://qibawiki.rsna.org/index.php/Public_Comment_Process for more guidance on the comment resolution process
Checklist Resolutions Legend
The "Checklist Resolutions" tab is for collating the checklist responses submitted by all the Technical Confirmation sites | Site Conformance with the requirement during Technical Feasibility testing
| Y | Conformed
The "Comment Resolutions" tab is for collating comments submitted by all the Technical Confirmation sites | N | Did not conform
Both tabs support Biomarker Committee assessment of the feedback and potential modification of the Profile | Site Opinion of the requirement
| R | Do routinely
| FW | Feasible, will do to conform | If site reports didn't conform but feasible, ask for details on why not done
TODO Incorporate QIBA User Feedback Template content. | FN | Feasible, but will not do to conform | Ask for details on why not
| NF | Not feasible | Ask for details on why not

## Checklist Resolutions
| Checklist items (See Comment Resolutions tab for feedback comments) | Gottsegen National Cardiovascular Center |  | Walter Reed Medical Center |  | Medical University of South Carolina |  | University of Cagliari Medical |  | Committee Discussion | Resolution | Status
|  | Conform | Opinion | Conform | Opinion | Conform | Opinion | Conform | Opinion |  |  | TBD | To be discussed
[12 empty] | OK | No action requested
[12 empty] | Discuss | Need to decide resolution
Parameter | Requirement | [10 empty] | TODO | Resolution decided
Acquisition Device Checklist | [11 empty] | Done | Profile update completed
In-plane Spatial Resolution | Shall validate that the protocol achieves an f50 value that is greater than 0.35 line pairs per mm for both air and soft tissue edges. See section 4.1. Assessment Procedure: In-plane Spatial Resolution | Y | R | Y | FW | Y | R | Y | R | held specific discussions regarding how to make these routinely possible | Dr. Samei provided greater guidance edited into text, and indicated that specified tool wiill be made easier to use | Done
Pixel noise | Shall validate that the protocol achieves a standard deviation that is < 30HU. See 4.2. Assessment Procedure: Pixel noise | Y | R | Y | FW | Y | R | Y | R | held specific discussions regarding how to make these routinely possible | Dr. Samei provided greater guidance edited into text, and indicated that specified tool wiill be made easier to use | Done
Acquisition Protocol | Shall be capable of making validated protocols (designed and validated by the manufacturer and/or by the site) available to the technologist at scan time. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Temporal Resolution | Shall achieve an effective rotation time of less than or equal to 400ms. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Image Analysis Tool Checklist
Basis of cross-sectional area results | Shall base cross-sectional area results on obliquely-resliced orthogonal to centerline at spacing less than or equal to 0.5mm | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
Basis of volume results | Shall base volume results on three-dimensional object definitions (specifically excluding methods such as determining cross-sectional areas and multiplying by the slice thickness, or other approximations) | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
Confidence interval | Shall be able to display to the Imaging Physician, for each measurand, the range of plausible values for the given measurement stated in terms of the completed validation for the tool as a 95% interval. | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
High-Risk Plaque | Shall be validated to achieve bias and linearity (expressed as intercept, slope, and quadratic term) within the values shown in the following table. See 4.4. Assessment Procedure: Tissue Characteristics Bias and Linearity, noting that the full 95% confidence intervals (not only the point estimates) shall meet or exceed the specifications given in the Claims section. | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
Multiple encounters | Shall be able to present the reader with both encounters side-by-side for comparison when processing the second encounter. | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
Multiple encounters | Shall be able to re-process the first encounter (e.g. if it was processed by a different Image Analysis Tool or Imaging Physician). | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
Multiple Lesions | Shall allow multiple lesions to be measured. | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
Multiple Lesions | Shall either correlate each measured lesion across encounters or support the Imaging Physician to unambiguously correlate them. | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
Reader variability | Shall be validated to achieve Intra-reader wSD and Inter-reader wSD less than the values shown in the following table. See 4.5. Assessment Procedure: Reader / Image Analysis Tool Variability, noting that the full 95% confidence intervals (not only the point estimates) shall meet or exceed the indicated specifications when tested over range as given in Claims section. Lumen Area (mm2) Intra-reader wSD: 2.5, Inter-reader wSD: 5.0 Wall Area (mm2) Intra-reader wSD: 2.5, Inter-reader wSD: 5.0 Maximum Wall Thickness (mm) Intra-reader wSD: 0.75, Inter-reader wSD: 1.0 Plaque Burden (ratio) Intra-reader wSD: 0.1, Inter-reader wSD: 0.1 Calcified Area (mm2) Intra-reader wSD: 1.0, Inter-reader wSD: 1.5 LRNC Area (mm2) Intra-reader wSD: 1.0, Inter-reader wSD: 1.5 | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
Tissue Composition | Shall be validated to achieve bias and linearity (expressed as intercept, slope, and quadratic term) within the values shown in the following table. See 4.4. Assessment Procedure: Tissue Characteristics Bias and Linearity, noting that the full 95% confidence intervals (not only the point estimates) shall meet or exceed the indicated specifications when tested over range as given in Claims section: Calcified Area (mm2) Bias: ±1.5, Intercept: ±2, Slope: 1±.5, Quadratic term: ±.1 LRNC Area (mm2) Bias: ±3, Intercept: ±3.5, Slope: 1±.8, Quadratic term: ±.3 | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
Vessel structure | Shall be validated to achieve bias and linearity (expressed as intercept, slope, and quadratic term) within the values shown in the following table. See 4.3. Assessment Procedure: Vessel Structure Bias and Linearity, noting that the full 95% confidence intervals (not only the point estimates) shall meet or exceed the indicated specifications when tested over range as given in Claims section: Lumen Area (mm2) Bias: ±2, Intercept: ±1.0, Slope: 1±.1, Quadratic term: ±.1 Wall Area (mm2) Bias: 2, Intercept: ±10, Slope: 1±.1, Quadratic term: ±.1 Maximum Wall Thickness (mm) Bias: ±1, Intercept: ±1, Slope: 1±.1, Quadratic term: ±.1 Plaque Burden (ratio) Bias: ±0.1, Intercept: ±.1, Slope: 1±.1, Quadratic term: ±.1 | Y | FW | Y | FW | Y | FW | Y | FW |  |  | OK
Imaging Physician Checklist
Artifacts | Shall confirm the images containing the lesion are free from artifacts due to dense objects, anatomic positioning (e.g., arms down at sides), or equipment issues (e.g., ring artifacts). | Y | R | Y | R | Y | R | Y | R |  |  | OK
Contrast Enhancement | Shall confirm that the intravascular level of contrast enhancement, if any, is appropriate for evaluating the lesion. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Field of View | Shall confirm that the image field of view (FOV) resulting from acquisition and reconstruction settings appears consistent with baseline. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Multiple encounters | Shall re-process the first encounter if it was processed by a different Image Analysis Tool or Imaging Physician. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Pacemaker leads, stents | Shall confirm that anatomy assessed does not contain metal artifacts. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Patient Motion Artifacts | Shall confirm the images containing the lesion are free from artifact due to motion. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Patient Positioning Consistency | Shall confirm that any lesion deformation due to patient positioning is consistent with baseline (e.g. lesions may deform differently if the patient is supine in one scan and prone in another). | Y | R | Y | R | Y | R | Y | R |  |  | OK
Physiological motion artifact (particularly cardiac) | Shall confirm the images containing the lesion are free from artifact due to motion based on visual review for blurred anatomic features. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Result Verification | Shall review & approve segmentations produced by the Image Analysis Tool. | Y | R | Y | R | Y | R | Y | FW |  |  | OK
Scan Plane Consistency | Shall confirm that the anatomical slice orientation (due to gantry tilt or patient head/neck repositioning) is consistent with baseline. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Total Collimation Width | Shall set to Greater than or equal to 18mm. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Use of intravenous contrast | Shall prescribe a contrast protocol to achieve appropriate lumen conspicuity relative to wall tissues. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Physicist Checklist
Acquisition Protocol | Shall prepare a protocol to meet the specifications in this table. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Acquisition Protocol | Shall ensure technologists have been trained on the requirements of this profile. | Y | R | Y | R | Y | FW | Y | R |  |  | OK
In-plane Spatial Resolution | Shall validate that the protocol achieves an f50 value that is Greater than 0.35 mm-1 for both air and soft tissue edges. See section 4.1. Assessment Procedure: In-plane Spatial Resolution | Y | R | Y | R | Y | R | Y | R |  |  | OK
Nominal Tomographic Section Thickness (T) | Shall set to Less than or equal to 0.75mm. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Nominal Tomographic Section Thickness (T) | Shall set to Less than or equal to 1.0mm. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Pixel noise | Shall validate that the protocol achieves a standard deviation that is < 30HU. See section 4.2. Assessment Procedure: Pixel noise | Y | R | Y | R | Y | R | Y | R |  |  | OK
Reconstructed Image Interval | Shall set to less than or equal to the Reconstructed Image Thickness (i.e. no gap, may have overlap). | Y | R | Y | R | Y | R | Y | R |  |  | OK
Reconstructed Image Thickness | Shall be less than 1mm. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Reconstructed In-plane Voxel Size | Shall set to less than or equal to 0.625mm | Y | R | Y | R | Y | R | Y | R |  |  | OK
Reconstruction Protocol | Shall prepare a protocol to meet the specifications in this table. Shall ensure technologists have been trained on the requirements of this profile. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Total Collimation Width | Shall set to Greater than or equal to 16mm. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Reconstruction Software Checklist
Image Header | Shall record in the DICOM image header the actual values for the tags listed in the DICOM Tag column "Protocol Design Specification" as well as the model-specific Reconstruction Software parameters utilized to achieve conformance. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Image Header | Shall record in the DICOM image header the actual values for the tags listed in the DICOM Tag column "Protocol Design Specification" as well as the model-specific Reconstruction Software parameters utilized to achieve conformance. | Y | R | Y | R | FW | R | Y | R |  |  | OK
Reconstruction Protocol | Shall be capable of performing reconstructions and producing images with all the parameters set as specified "Protocol Design Specification". | Y | R | Y | R | Y | R | Y | R |  |  | OK
Technologist Checklist
Acquisition Protocol | Shall select a protocol that has been previously prepared and validated for this purpose. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Artifact Sources | Shall remove or position potential sources of artifacts (specifically including breast shields, metal-containing clothing, EKG leads, and other metal equipment) such that they will not degrade the reconstructed CT image. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Breath hold | Shall instruct the subject in proper breath-hold and start image acquisition shortly after full inspiration, taking into account the lag time between full inspiration and diaphragmatic relaxation. | Y | R | Y | R | Y | R | Y | R |  |  | OK
ECG Gating | Shall use prospective ECG gating and consider iterative reconstruction to allow for the lowest possible radiation exposure. If the heart rate is too high, retrospective ECG gating may be required to obtain optimal images. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Nitrates | Shall administer nitrates as prescribed, 5-7 minutes after nitro is administered. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Reconstructed Image Interval | Shall set to less than or equal to the Reconstructed Image Thickness (i.e. no gap, may have overlap) and consistent with baseline. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Reconstructed Image Thickness | Shall be less than 1mm if not set in the protocol. | Y | R | Y | R | Y | R | Y | R |  |  | OK
Reconstruction Field of View | Shall ensure the Field of View spans at least the full extent of the thoracic cavity, but not substantially greater than that. | Y | R | Y | FW | Y | R | N | FN | In some scanners FOV>25cm increases the radiation dose (Philips iCT), and also would result in in-plane voxel dimensions not conform with the objective | Made specification more specific with respect to quantitation but with footnote regarding Profile scope. | Done
Reconstruction Protocol | Shall select a protocol that has been previously prepared and validated for this purpose. | [10 empty] | OK | bar: Y to FN
Table Height & Centering | Shall adjust the table height for the mid-axillary plane to pass through the isocenter. Shall center the thorax shall be centered in the AP and L/R directions according to the following: table height shall be adjusted for the mid axillary plane to pass through the isocenter and the sagittal laser line shall pass through the sternum from suprasternal notch to xiphoid process. | [10 empty] | OK | bar: Y to FN
Use of intravenous contrast | Shall use the prescribed intravenous contrast protocol. | [10 empty] | OK | bar: Y to FN