## Abdomen
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Bilgili, Mirace YK | Reproductibility of apparent diffusion coefficients measurements in diffusion-weighted MRI of the abdomen with different b values | Eur. J. of Radiol. |  |  | 2011 | http://www.sciencedirect.com/science/article/pii/S0720048X11005869
Braithwaite, AC et al | Short- and midterm reproducibility of apparent diffusion coefficient measurements at 3.0-T diffusion-weighted imaging of the abdomen | Radiology | 250 | 459-465 | 2009 | http://radiology.rsna.org/content/250/2/459.long | 41
Rosenkrantz, M et al | Diffusion‐weighted imaging of the abdomen at 3.0 Tesla: Image quality and apparent diffusion coefficient reproducibility compared with 1.5 Tesla | J. Mag. Reson. Imag. | 33 | 128-135 | 2011 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22395/full | 13
Kwee, TC et al | Comparison and reproducibility of ADC measurements in breathhold, respiratory triggered, and free‐breathing diffusion‐weighted MR imaging of the liver | J. Mag. Reson. Imag. | 28 | 1141-1148 | 2008 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21569/full | 70
Taouli, B et al | Diffusion-weighted MR Imaging of the Liver | Radiology | 254 | 47-66 | 2010 | http://radiology.rsna.org/content/254/1/47.short | 106
Mo, YH et al | Hepatic ADC value correlates with cirrhotic severity of patients with biliary atresia | Eur. J. of Radiol. | 80 | e253-e257 | 2011 | http://www.sciencedirect.com/science/article/pii/S0720048X10005474 | 1
DeVries, A et al | Tumor microcirculation and diffusion predict therapy outcome for primary rectal carcinoma | Int. J. of Rad. Oncol. | 56 | 958-965 | 2003 | http://www.sciencedirect.com/science/article/pii/S0360301603002086 | 116
Mannelli, l et al | Assessment of Tumor Necrosis of Hepatocellular Carcinoma After Chemoembolization: Diffusion-Weighted and Contrast-Enhanced MRI With Histopathologic Correlation of the Explanted Liver | Am. J. of Roent. | 193 | 1044-1052 | 2009 | http://www.ajronline.org/content/193/4/1044.short | 35
Miquel, ME et al | In vitro and in vivo repeatability of abdominal diffusion-weighted MRI | Brit. J. of Radiol. |  |  | 2012 | http://bjr.birjournals.org/content/early/2012/06/06/bjr.32269440.abstract
Eccles, CL et al | Change in diffusion weighted MRI during liver cancer radiotherapy: Preliminary observations | Acta Oncologica | 48 | 1034-1043 | 2009 | http://informahealthcare.com/doi/abs/10.1080/02841860903099972 | 22
Giannarini G, et al | Potential and Limitations of Diffusion-Weighted Magnetic Resonance Imaging in Kidney, Prostate, and Bladder Cancer Including Pelvic Lymph Node Staging: A Critical Analysis of the Literature | Eur. Urology | 61 | 326-340 | 2012 | http://www.sciencedirect.com/science/article/pii/S0302283811010268 | 7
Labregts, DMJ et al | Value of ADC measurements for nodal staging after chemoradiation in locally advanced rectal cancer—a per lesion validation study | Eur. Radiology | 21 | 265-273 | 2011 | http://www.ncbi.nlm.nih.gov/pmc/articles/PMC3034880/ | 13
Colagrande, S et al | MR‐diffusion weighted imaging of healthy liver parenchyma: Repeatability and reproducibility of apparent diffusion coefficient measurement | J. Mag. Reson. Imag. | 31 | 912-920 | 2010 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22117/full | 8
Kim, SY et al | Reproducibility of measurement of apparent diffusion coefficients of malignant hepatic tumors: Effect of DWI techniques and calculation methods | J. Mag. Reson. Imag. |  |  | 2012 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.23744/full
Kim, SY et al | Malignant Hepatic Tumors: Short-term Reproducibility of Apparent Diffusion Coefficients with Breath-hold and Respiratory-triggered Diffusion-weighted MR Imaging | Radiology | 255 | 815-823 | 2010 | http://radiology.rsna.org/content/255/3/815.short | 14
Yoshikawa, T et al | ADC Measurement of Abdominal Organs and Lesions Using Parallel Imaging Technique | Am. J. Roent. | 187 | 1521-1530 | 2006 | http://www.ajronline.org/content/187/6/1521.short | 145

## Bone
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Lee, KC et al | A Feasibility Study Evaluating the Functional Diffusion Map as a Predictive Imaging Biomarker for Detection of Treatment Response in a Patient with Metastatic Prostate Cancer to the Bone | Neoplasia | 9 | 1003-1011 | 2007 | http://www.ncbi.nlm.nih.gov/pmc/articles/PMC2134897/ | 50 | Prostate cancer and metastatic disease to the bone. | This study presents the feasibility of functional Diffusion Map (fDM) measurements in osseous lesions over time and shows that changes in fDM values were consistent with therapeutic response. | 1 | 3-T | Achieva model; Philips Medical Systems, Bothell, WA | 8 averages for low (b = 0 sec/mm2) | 2 | 0, 800. | single | Apparent diffusion coefficient (ADC) maps were calculated in the routine manner given by the logarithm of the ratio of low-b and high-b images, then scaled by the inverse of b-value difference. | fat-suppressed, single-shot, spinecho, echo-planar imaging | quadrature body coil for transmission and a six-channel cardiac coil for reception | none | yes | 5-mm | 350 mm right/left by 302 mm anterior/posterior | 200 X 172 | 8 averages for low and 16 averages for high | 58 | 2000 | SENSE factor = 2 | no
Uhl, M et al | Osteosarcoma: Preliminary Results of In Vivo Assessment of Tumor Necrosis After Chemotherapy With Diffusion- and Perfusion-Weighted Magnetic Resonance Imaging | Invest. Radiol. | 41 | 618-623 | 2006 | http://journals.lww.com/investigativeradiology/Abstract/2006/08000/Osteosarcoma__Preliminary_Results_of_In_Vivo.3.aspx | 62 | osteosarcoma of the limbs | We sought to evaluate diffusion and perfusion weighted 1.5 T magnetic resonance imaging (MRI) in detecting tumor necrosis with histologic correlation after preoperative chemotherapy. | 8 | 1.5-T | Symphony, Siemens Medical Solutions, Erlangen, Germany | 1 | 2 | 0, 700 | single | The apparent diffusion coefficient (ADC) was calculated from the diffusion-weighted images according to the following equation: ADC = (ln (S0/S1))/(b1 - b0) | echoplanar imaging-sequence | flexible phased-array surface coil | none | yes | 5 mm | NA | 128 x 96 | NA | 80 | 5000 | no | no
Oka, K et al | The value of diffusion-weighted imaging for monitoring the chemotherapeutic response of osteosarcoma: a comparison between average apparent diffusion coefficient and minimum apparent diffusion coefficient | Skeletal Radiology | 39 | 141-146 | 2010 | http://www.springerlink.com/content/p1707m2v3106202x/ | 16 | osteosarcomas (subtypes: 16 osteoblastic, 4 chondroblastic, 1 telangiectatic, and 1 small cell type). | The objective of this study was to evaluate whether the average apparent diffusion coefficient (ADC) or the minimum ADC is more useful for evaluating the chemotherapeutic response of osteosarcoma. | 22 | 1.5 T | Magnetom Vision; Symphony, Siemens, Erlangen, Germany | 1 | 2 | 0 and 1,000 s/mm2 | DW images were acquired in motion-probing gradient (MPG) pulses that were applied sequentially along three directions (x, y, and z axes) | ADC values were calculated using the following equation: ADC = (ln S0×S900)/(b900−b0) | spin echo echo-planar imaging sequence | four-element phased-array torso coil and an endorectal coil. | none | no | 5 mm | 220×220 mm | 128×64 | n/a | 105 | 4504 | no | no
Baunin, C et al | Value of diffusion-weighted images in differentiating mid-course responders to chemotherapy for osteosarcoma compared to the histological response: preliminary results | Skeletal Radiology |  |  | 2012 | http://www.springerlink.com/content/v6868589k98148x0/ |  | long-bone osteosarcoma | To assess the accuracy of DW-MRI in evaluating response to chemotherapy in the treatment of osteosarcoma, more specifically at mid-course of treatment. | 15 | NA | Philips (Best, The Netherlands) | 1 | 2 | 0 and 900 | three gradient directions. | ADC values were calculated using the following equation: ADC = (ln S0×S900)/(b900−b0) | SPIR coronal DW turbo-spin echo | Body coil: sagittal and coronal; Surface coil: centered on the bone tumor | none | yes | 20 mm | NA | matrix 112×89 | NA | 138 | 1500 | no | no
Messiou, C et al | Optimising diffusion weighted MRI for imaging metastatic and myeloma bone disease and assessing reproducibility | Eur. Radiol. | 21 | 1713-1718 | 2011 | http://www.springerlink.com/content/k12nv2k4314u751m/ | 2 | metastatic involvement of spine and pelvis secondary to carcinoma of the prostate, breast or myeloma | To establish normal bone marrow values of apparent diffusion coefficient (ADC) over an age range, compare them with metastatic and myelomatous involvement, to establish reproducibility and to optimise b values. | 84 (41 healthy volunteers and 43 patients with metastatic involvement of spine and pelvis secondary to carcinoma of the prostate, breast or myeloma) | 1.5 T | Avanto, Siemens, Forchheim, Germany | 1 | 6 | 0, 50, 100, 250, 500 and 750 smm−2; Additionally: In a single patient from cohort 3 DWI was repeated to investigate the feasibility of introducing a higher b value: 0, 50, 100, 250, 500 and 750 and 1,400 smm−2 100, 250, 500 and 750 and 1,400 smm−2 | 3 orthogonal directions | standard ADC calculation | Axial single shot twice refocused diffusion weighted spin echo planar images | external coil array | none | yes | 5 mm, | 350 mm | unknown | 4 | 80 | 3000 | no | no
Reischauer C, et al | Bone Metastases from Prostate Cancer: Assessing Treatment Response by Using Diffusion-weighted Imaging and Functional Diffusion Maps—Initial Observations | Radiology | 257 | 523-531 | 2010 | http://radiology.rsna.org/content/257/2/523.short | 22 | prostate cancer with pelvic bone metastases | To prospectively investigate and monitor the response to antiandrogen treatment of bone metastases in patients with prostate cancer by using diffusion-weighted (DW) magnetic resonance (MR) imaging with the apparent diffusion coefficient (ADC) and functional diffusion maps (DMs). | 9 | 1.5-T | Achieva, release 11.2.9; Philips Healthcare, Best, the Netherlands | 1 | 5 | 0, 200, 400, 600, 800 sec/mm2 | axial imaging of the pelvis; right to left direction and anteroposterior direction (see FOV comments) | monoexponential fit for all b factors | axial DW imaging | four-element sensitivity-encoding body coil (Philips Healthcare) | none | no | 6 mm | 400mm (right to left direction); 256 mm (anteroposterior direction) | 200x128 | 8 | 64 | 4687 | no | no
Nakanishi, K et al | Whole-body MRI for Detecting Metastatic Bone Tumor: Diagnostic Value of Diffusion-weighted Images | Mag. Reson. Med. Sci. | 6 | 147-155 | 2007 | https://www.jstage.jst.go.jp/article/mrms/6/3/6_3_147/_article | 91 | breast cancer, prostrate cancer, thyroid cancer, liposarcoma, leiomyosarcoma, and extraskeletal Ewing sarcoma | We assessed the diagnostic value of whole body magnetic resonance (MR) imaging (WB-MRI) using diffusion-weighted images (DWI) for detecting bone metastasis and compared it with that of skeletal scintigraphy (SS). | 30 | 1.5 T | Master Philips Medical Systems, Best, The Netherlands | 1 | 2 | 0 and 600 s/mm2 | axial slices were acquired from the lower neck to the bottom of the pelvis using a 2-station approach…the radial directions of maximal intensity projection (MIP) images of each station were reconstructed and displayed as anteroposterior and left lateral views. | NA | single shot short T1 inversion recovery-echo planar imaging (STIR-EPI) sequences | quadrature body coil | none | yes | 5 mm with 1 mm overlap of axial view during free breathing | 450 mm | 112 x 112 | ? | 59 | 6243 | no | no
Messiou, C et al | Assessing response in bone metastases in prostate cancer with diffusion weighted MRI | Eur. Radiol. | 21 | 2169-2177 | 2011 | http://www.springerlink.com/content/t5v6247243v512v7/ | 3 | prostate cancer and metastatic bone disease | The purpose of this study therefore was to determine whether changes in ADC in bone metastases secondary to carcinoma of the prostate are significantly different in responders compared with progressors on chemotherapy in order to establish its utility as a response biomarker in bone. | 38 | 1.5 T | Siemens Avanto, Siemens, Forchheim, Germany | 1 | 6 | 0, 50,100, 250, 500, 750 smm−2 | 3 orthogonal directions | Monoexponential calculation | Axial single shot twice refocused diffusion-weighted spin echo planar | external coil array | none | Yes | 5 mm | 350 mm | unknown | 4 | 80 | 3000 | no | no
Hayashida, Y et al | Monitoring therapeutic responses of primary bone tumors by diffusion-weighted image: initial results | Eur. Radiol. | 12 | 2637-2643 | 2006 | http://www.springerlink.com/content/gw3x6469771j0711/ | 76 | osteogenic and Ewing sarcomas | The purpose of our study was to investigate whether quantitative diffusion-weighted images (DWI) were useful for monitoring the therapeutic response of primary bone tumors. | 18 | 1.5 T | Symphony, Siemens Erlangen, Germany | 1 | 2 | 0 and 1000 s/mm2 | motion-probing gradient (MPG) pulses applied sequentially along three (x, y, and z axes) directions | monoexponential fit | axial or sagittal plane using a spin-echo, echo planar imaging (EPI) sequence | All studies were performed using a circular polarized (CP) extremity, CP flex large phased-array, or CP body-array coil | none | yes | 5 mm | 220 mm | 128 x 64 | 1 | 105 | 4504 | no | no
Messiou C, et al | Optimising diffusion weighted MRI for imaging metastatic and myeloma bone disease and assessing reproducibility. | European radiology | 21 | 1713-8 | 2011 | http://www.ncbi.nlm.nih.gov/pubmed?term=21472473 |  | metastatic and myeloma bone disease | To establish normal bone marrow values of apparent diffusion coefficient (ADC) over an age range, compare them with metastatic and myelomatous involvement, to establish reproducibility and to optimise b values. | 84 | 1.5 T | Avanto, Siemens, Forchheim, Germany | 1 | 6 | 0, 50, 100, 250, 500 and 750 smm−2 | 3 orthogonal | monoexponential fit | Axial single shot twice refocused diffusion weighted spin echo planar images | external coil array | none | yes | 5 mm | 350 mm | n/a | 4 | 80 | 3000 | no | no
Eiber M et al | Whole-body MRI including diffusion-weighted imaging (DWI) for patients with recurring prostate cancer: technical feasibility and assessment of lesion conspicuity in DWI. | Journal of magnetic resonance imaging - JMRI | 33 | 1160-70 | 2011 | http://www.ncbi.nlm.nih.gov/pubmed?term=21509875 |  | prostate cancer | To evaluate the principal methodological aspects of whole-body magnetic resonance imaging (MRI) including diffusion-weighted imaging (DWI) with background suppression using a time-optimized protocol for restaging of prostate cancer patients in a technical feasibility study. | 17 | 1.5 T | Magnetom Avanto, Siemens Medical Solutions, Erlangen, Germany | 0 | 2 | 50 and 500 s/ mm2 | three orthogonal directions (X, Y, and Z). | Trace images were synthesized for each b-value | Axial single- shot spin-echo echo-planar DWI-sequences including background suppression with spectral selection attenuated inversion recovery (SPAIR) | Two 6-element body matrix coils placed anteriorly were used in conjunction with two posterior spine clusters (three channels each)...In addition the 12-element head matrix coil, 4-element neck matrix coil, 6-element peripheral angio matrix coil were used with automated table motion | none | yes | 5 mm | NA | 148 x 168 | na | 78 | 7,500 (neck); 1,900 (thorax) | GRAPPA 2 | no
Sommer G, et al | Signal characteristics of focal bone marrow lesions in patients with multiple myeloma using whole body T1w-TSE, T2w-STIR and diffusion-weighted imaging with background suppression. | European radiology | 21 | 857-62 | 2011 | http://www.ncbi.nlm.nih.gov/pubmed?term=20853109 |  | multiple myeloma | This study analyses the diagnostic potential of Diffusion-Weighted Imaging with Background Suppression (DWIBS) in the detection of focal bone marrow lesions from multiple myeloma. The signal and contrast properties of DWIBS are evaluated in correlation with the serum concentration of M-component (MC) and compared with established T1- and T2-weighted sequences. | 103 consecutive studies of 81 patients | 1.5 T | Magnetom Avanto, Siemens Healthcare, Erlangen, Germany | 1 | 2 | 0, 800 | na | monoexponential fit | DWIBS (DWI with Background Suppression) with ss-EPI | a dedicated 18-channel coil array system (Total Imaging Matrix) | none | yes | 5 mm | 50×50 cm2 | 384×269 | na | 11 | 682 | na | no
Moon Won-Jin, et al | Diffusion-weighted imaging with sensitivity encoding (SENSE) for detecting cranial bone marrow metastases: comparison with T1-weighted images. | Korean journal of radiology | 8 | 185-91 | 2007 | http://www.ncbi.nlm.nih.gov/pubmed?term=17554184 |  | cranial bone marrow (CBM) metastases | to determine whether diffusion-weighted imaging (DWI) with sensitivity encoding (SENSE) could detect bone marrow involvement in patients with cranial bone marrow (CBM) metastases. | 13 | 1.5-T | Intera; Philips Medical Systems, Best, the Netherlands | 1 | 2 | 0 and 1,000 sec/mm2 | applying six diffusion gradients in different directions for each slice | ADC maps were also obtained at a workstation | Single-shot spin-echo echo-planar DWI sequences | six-channel SENSE head coil | none | yes | 5 mm with a 1mm intersection gap | 23 cm | 256x128 | averaging six images with different diffusion gradients. | 66 | 3200 | SENSE | no
Nemeth A J, et al | Improved detection of skull metastasis with diffusion-weighted MR imaging. |  | 28 | 1088-92 | 2007 | http://www.ncbi.nlm.nih.gov/pubmed?term=17569964 |  | extracranial primary malignancy | to determine if diffusion-weighted MR imaging (DWI) could improve the detection of skull metastasis in patients with primary carcinomas that metastasized to bone compared with conventional MR imaging. | 75 | 1.5 T | NA | 1 | six high-b values and a single low-b value | high b value was 1000 s/mm2. low b value was 0 seconds/mm2 | 23 axial sections | na | single-shot echo-planar imaging | NA | none | no | 5 mm with a 1mm gap | 22 cm x 22 cm | 128x128 | 5 signal intensity averages | 90 | 5000 | na | no
Raya J G, et al | Feasibility of a RARE-based sequence for quantitative diffusion-weighted MRI of the spine | Eur. Radiol. | 17 | 2872-79 | 2007 | http://www.springerlink.com/content/24l80567u0318532/ | 9 | normal and pathologic spinal bone marrow |  | 20N/16P | 1.5 T | Siemens, Magnetom Symphony Maestro |  | 4 | 50 ,250, 500, 750 | [13 empty] | mean CoV 2.1%, temporal (10 patients), 6.4% spatial (12 patients)

## Brain
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Ellingson, BM et al | Validation of functional diffusion maps (fDMs) as a biomarker for human glioma cellularity | J. Mag. Reson. Imag | 31 | 538-548 | 2010 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22068/full | 32 | Glioma & normal appearing brain | Path cellularity | 69 | 1.5 | GE | 1 | 2 | 0, 1000 | 3 | Monoexp. 2-pt | DW SS EPI & Propellor | Head | No | Implicit | 5 | 240 | 128 | 1 | 102.2 | 8000 | no | Test Retest from 1week to 1yr. | 95% CI 0.25 to 0.75 x 10-3mm2/s dependent on tissue
Rana, AK et al | Apparent diffusion coefficient (ADC) measurements may be more reliable and reproducible than lesion volume on diffusion-weighted images from patients with acute ischaemic stroke–implications for study design | Magn. Reson. Imag | 21 | 617-624 | 2003 | http://www.sciencedirect.com/science/article/pii/S0730725X03000870 | 18 | Acute Isch & Stroke | Lesion ADC & volume | 15; 38 datasets | 2??? | GE | 1 | 2 | 0, 700 | 3 | Monoexp. 2-pt | DW SS EPI | Head | No | Implicit | 6 | [6 empty] | Not TestRetest. 38 scans considered independent. Intra- Inter-reader CoV | Intra-CoV = 0.2+/- 12%; Inter CoV = 7.7+/-19%
Jain R, et al | Imaging response criteria for recurrent gliomas treated with bevacizumab: Role of diffusion weighted imaging as an imaging biomarker | J. Neuro-oncology | 96 | 423-431 | 2010 | http://www.springerlink.com/content/4938685676217400/ | 28 | Glioma | ADC change vs Disease progression |  | 3 | GE | 1 | 2 | 0, 1000 | 25 | Mean eigen & FA from tensor | DTI SS EPI | Head | No | Implicit | 2.6 | unk | 128x128 | 1 | unk | unk | unk | no | Serial (>6wks) ADC change in Tx patients; Progressors had decreased ADC; non progressors had no change
Ozturk, A et al | Regional Differences in Diffusion Tensor Imaging Measurements: Assessment of Intrarater and Interrater Variability | Am. J. Neurorad. | 29 | 1124-1127 | 2008 | http://www.ajnr.org/content/29/6/1124.short | 17 | ClosedHeadInj/Normal |  | 7 / 5 controls | 3 | Philips | 1 | 2 | 0, 700 | 32 | Mean eigen & FA from tensor | DTI SS EPI | Head 8ch | No | Implicit | 2.2 | 212 | 96x96 |  | 84 | 7897 |  | SingleExam TestRetest??? 2xSD range 6.6% to 25%
Miller, J and Sorensen, G | Imaging bio- markers predictive of disease/therapy outcome: ischemic stroke and drug development | Imag. Drug Disc. And Early Clinical Trials | 62 | 319-356 | 2005 | http://www.springerlink.com/content/m740m57n4102442r/ | 5 | Stroke | ADC change w/ flow [ref97] | NA - review | [10 empty] | Implicit | [8 empty] | Gen Lin Models of multiple parameters
Tofts, PS and Collins, DJ | Multicentre imaging measurements for oncology and in the brain | Brit. J. Radiology | 84 | S213-S226 | 2011 | http://bjr.birjournals.org/content/84/Special_Issue_2/S213.short | 1 | Onc | General discussion of sources and degree of variance | NA - review | [10 empty] | Implicit | [7 empty] | CV's table 3 | Indirect refs to multiple sources.
Hou, DJ et al | Diffusion-Weighted Magnetic Resonance Imaging Improves Outcome Prediction in Adult Traumatic Brain Injury | J. Neurotrauma | 24 | 1558-1569 | 2007 | http://online.liebertpub.com/doi/abs/10.1089/neu.2007.0339?2 | 38 | TBI / Normal | GCS severity | 37 pats/35 normal |  | Siemens |  | 2 | 500, 1000 | 3 | Monoexp. 2-pt | DW SS EPI | Head CP | No | Implicit | 4 to 5 | UNK (200-220cm) | unk |  | 110 | 4000 | no | no; ADC differences from controls; dichotomized GOS score
Ellingson, BM et al | Volumetric analysis of functional diffusion maps is a predictive imaging biomarker for cytotoxic and anti-angiogenic treatments in malignant gliomas | J. Neuro-oncology | 102 | 95-103 | 2011 | http://www.springerlink.com/content/p276m0067p324202/ | 11 | Glioma | TTP and OS | 50 | 1.5 | GE | 1 | 2 | 0, 1000 | 3 | Monoexp. 2-pt | DW SS EPI & Propellor | Head | No | Implicit | 5 | 240 | 128x128 |  | 102 | 8000 | no | Serial change in ADC > 0.4 | Rate of fDM change, implies progression. Voxel-by-voxel differences
Paldino, MJ et al | Repeatability of quantitative parameters derived from diffusion tensor imaging in patients with glioblastoma multiforme | J. Mag. Reson. Imag | 29 | 1199-1205 | 2009 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21732/full | 6 | GBM | Reapeatability of ADC & FA | 16 | 1.5 | Siemens | 1 | 2 | 0, 1000 | 12 | Mean eigen & FA from tensor | DTI SS EPI | Head | No | Implicit | 5 | 220 | 128x128 | 4 | 100 | 6000 | 2 | Test Retest CI's : 7.4% and 13.9% | Highly relevant paper to DWI Profile
Sasaki, M et al | Variability in Absolute Apparent Diffusion Coefficient Values across Different Platforms May Be Substantial: A Multivendor, Multi-institutional Comparison Study | Radiology | 249 | 624-630 | 2008 | http://radiology.rsna.org/content/249/2/624.short | 30 | Normal | Intra- Inter-vendor variability | 12 | 1.5(8) & 3(2) | 4 vendors/10systems | 1 | 2 | 0, 1000 | 3 | Monoexp. 2-pt | DW SS EPI | Head | No | Implicit | 6 | 220 | 128x80-128 | 2 | 70-100 | 4000-8000 | no | Intra- Inter- system CoV
Chenevert, TL et al | Diffusion magnetic resonance imaging: an early surrogate marker of therapeutic efficacy in brain tumors | J. Nat. Canc. Instit. | 92 | 2029-2036 | 2000 | http://jnci.oxfordjournals.org/content/92/24/2029.short | 381 | Tumor | Change in ADC vs time post Tx | 2 | 1.5 | GE | 1 | 2 | 1, 1000 | 3 | Monoexp. 2-pt | DW SS EPI | Head | No | Implicit | 6 | UNK (200-220cm) | UNK (>128) | 1 | 80 | 10000 | No | no
Burdette, JH et al | Calculation of apparent diffusion coefficients (ADCs) in brain using two-point and six-point methods | J. Comput. Assit. Tomogr. | 22 | 792-794 | 1998 | http://journals.lww.com/jcat/Abstract/1998/09000/Calculation_of_Apparent_Diffusion_Coefficients.23.aspx | 55 | Disease / Normal | correlation of ADC methods | 10 | 1.5 | GE | 6 | 1 | 1, 100,, 200, 500, 800, 1000 |  | Monoexp. 2-pt fit; and nonlin regression | DW SS EPI | Head | No | Implicit | 5 | UNK (200-220cm) | 128x128 | 1 | 98 | 10000 | No | no, high correlation between methods | Two point ADC is good enough
Galloway, NR et al | Diffusion-Weighted Imaging Improves Outcome Prediction in Pediatric Traumatic Brain Injury | J. of Neurotrauma | 25 | 1153-1162 | 2008 | http://online.liebertpub.com/doi/abs/10.1089/neu.2007.0494?2 | 26 | TBI / Normal | ADC vs GCS score | 37 / 10controls (2-18yo) | 1.5 | Siemens | 0 | 2 | 500, 1000 | 3 | UNK | DW SS EPI | Head CP | No | Implicit | 5 | UNK (200-220cm) | UNK | UNK | 110 | 4000 | No | no | Dichotomized good vs poor groups, ADC ROI group differences; lower ADC with severity
Mills, SJ et al | Candidate Biomarkers of Extravascular Extracellular Space: A Direct Comparison of Apparent Diffusion Coefficient and Dynamic Contrast-Enhanced MR Imaging--Derived Measurement of the Volume of the Extravascular Extracellular Space in Glioblastoma Multiforme | Am. J. Neurorad. | 31 | 549-553 | 2010 | http://www.ajnr.org/content/31/3/549.short | 7 | GBM | ADC vs Ve (DCE) | 19 | 3 | Philips | 1 | 2 | 0, 1000 | 6 | Mean eigen value | DTI SS EPI | Head | No | Implicit | 4 | 230 | 128x128 | UNK | 68 | 2319 | UNK | no | No correlation found between ADC and Ve, either in voxel-by-voxel, or VOI)
Wilde, EA et al | Diffusion tensor imaging of acute mild traumatic brain injury in adolescents | Neurology | 70 | 948-955 | 2008 | http://www.neurology.org/content/70/12/948.short | 141 | Mild acute TBI | Concussion sympt & distress | 10 (14-19yo) & 10 contols | 3 | Philips | 1 | 2 | 0, 1000 | 30 | Mean eigen value | DTI SS EPI | Head 8ch | No | Implicit | 2 | 224 | 112x112 | 2 | 51 | 6318 | Yes = 2 | ICC&CC>0.98. Meas in CorpCallo. TBI vs Control grp differences | TBI incr FA, decr ADC than controls
Moffat, BA et al | Functional diffusion map: A noninvasive MRI biomarker for early stratification of clinical brain tumor response | Proc. Nat. Acad. Sci | 102 | 5524-5529 | 2005 | http://www.pnas.org/content/102/15/5524.short | 308 | PrimBrainTumor / Normal | Tx Response pre vs 3wk | 20 | 1.5 | GE | 1 | 2 | 0, 1000 | 3 | Monoexp. 2-pt fit | DW SS EPI | Head | No | Implicit | 6 | UNK (200-220cm) | UNK (>128) | 1 | 80 | 10000 | No | 95% CI of voxels, contralat brain | Voxel-by-voxel difference (fDM introduced)
Jansen, J et al | Reproducibility of Quantitative Cerebral T2 Relaxometry, Diffusion Tensor Imaging, and 1H Magnetic Resonance Spectroscopy at 3.0 Tesla | Invest Radiol | 42 | 327–337 | 2007 | http://www.ncbi.nlm.nih.gov/pubmed/17507802 | 20 | Normal (20-28 yo) | CV | 10 | 3.0T | Philips Achieva | Assume 1 | 1 | 800 | 15 | Trace (on scanner) | Single Shot EPI-SE | 8-channel |  |  | 2 | 256x256 | 128x128 |  | 62 | 6600 | Acceleration 2.5 | CV 1.4-1.6%, Repeatability Coefficient 32-36, ICC 0.73-0.74 | Scan times 10 min, CV FA≤5.3%, Median ADC 786-827
Pfefferbaum A, et al | Replicability of diffusion tensor imaging measurements of fractional anisotropy and trace in brain | J Magn Reson Imaging | 18 | 427– 433 | 2003 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.10377/abstract;jsessionid=E8CAE14C07CA756D88388051723B33EF.d03t03 | 58 | Normal (21-33 yo) | CV | 10 | 1.5T | GE scanners | 6 | 1 | 860 | 6 | Trace | Single Shot EPI-SE |  |  |  | 4 |  | 128x128 | 6 | 103 | 10000 |  | CV 0.93 (within) 1.15-1.64% (across scanners) whole brain, 0.84(within) 3.5-3.73%(across scanners) white matter,Corpus callosum CV=2.6% within scanners and CV=8.2-8.6% across scanners | Corpus callosum 380-410,420-430 whole brain
Bonekamp, D et al | Diffusion tensor imaging in children and adolescents: reproducibility, hemispheric, and age-related differences | Neuroimage | 34 | 733-42 | 2007 | http://www.sciencedirect.com/science/article/pii/S1053811906009621 | 83 | Normal (5-19 yo) | CV | 40 | 1.5T | GE | 2 | 1 | 1000 | 15 | Trace (DTI Studio) | Single Shot EPI-SE | CP Birdcage Head Coil |  |  | 5 | 240 | 96x96 | 1 | 93.7 |  |  | Intra-rater CV ≤2.7%, inter-rater CV ≤4.6%, between scan ≤3.4% | White matter 660-881
Walker, L et al | A Framework for the Analysis of Phantom Data in Multicenter Diffusion Tensor Imaging Studies | HBM | Epub |  | 2012 | http://onlinelibrary.wiley.com/doi/10.1002/hbm.22081/abstract |  | Normal Male (55 yo) | ICC | 1 - scanned 20 times | 1.5T | 3 Siemens, 2 GE | 1 | 1 | 1000 | 6 | Trace | Single Shot EPI-SE |  |  |  | 3-4 mm | Varied to get 3x3 mm in-plane |  | 4 | minTE | minTR=3000 |  | Inter-rater ICC 0.6-0.84, Intrarater ICC 0.16-0.40. After EPI Correction: Inter ICC=0.64-0.72, Intra ICC 0.28-0.36
Lee, C et al | Normal regional fractional anisotropy and apparent diffusion coefficient of the brain measured on a 3 T MR scanner | Neuroradiology | 51 | 3–9 | 2009 | http://www.springerlink.com/content/u1542574q66k0264/ | 16 | Normal (19-62) | Means of ADC in regions | 31 | 3.0T | Philips Intera | Assume 1 | 1 | 1000 | 32 | Trace (PRIDE) | Single Shot EPI-SE | 6-channel SENSE |  |  | 2.5 mm | 240 | 96x96 | 3 | 86 | 6000 |  | Only reported means of ADC in 20 regions
Fox, RJ et al | A Validation Study of Multicenter Diffusion Tensor Imaging: Reliability of Fractional Anisotropy and Diffusivity Values | AJNR Am J Neuroradiol. | 33 | 695-700 | 2012 | http://www.ajnr.org/content/33/4/695.long | 0 | Normal (35 yo) | CV, Concordance | 2 (scanned 5 times) | 3.0 T | 3 Siemens Trio, 2 GE Signas | 9 | 1 | 1000 | 33 | Trace | Twice-refocused EPI-SE |  |  |  | 2.5 mm | 256 | 104x104 | 4 | 100/94 (Siemens/GE) | 2500/4200 (Siemens/GE) |  | 4.8% (whole brain) and 4.1% (white matter) across 5 scanners, 2.2% same scanner, Concordance 0.88
Rana, AK et al | Apparent diffusion coefficient (ADC) measurements may be more reliable and reproducible than lesion volume on diffusion-weighted images from patients with acute ischaemic strokeÐimplications for study design | Magn Reson Imaging. | 21 | 617-24 | 2003 | http://www.mrijournal.com/article/S0730-725X%2803%2900087-0/abstract | 11 | Stroke | CV | 15 (scanned 38 times) | 2T | GE | 1 | 1 | 700 | 3 | Trace | Single Shot EPI-SE |  |  |  | 6 mm (skip 20%) | [6 empty] | Absolute ADC: Inter-rater CoV 20.5%, Intra-rater CoV 14.8%; Relative ADC: Inter-rater 73.2%, Intra-rater 18.3%
Brander, A et al | Diffusion tensor imaging of the brain in a healthy adult population: Normative values and measurement reproducibility at 3 T and 1.5 T | Acta Radiol. | 51 | 800-7. | 2010 | http://ar.rsmjournals.com/content/51/7/800.long | 5 | Normals (19-61 yo) |  | 40 | 3.0T (30 subjects)/1.5T (10 subjects) | Siemens Trio/Siemens Avanto | Assume 1 | 1 | 1000 | 20/12 |  | Single Shot EPI-SE | 12-channel Head Matrix |  |  | 3 mm (skip 0.9mm)/ 5 mm (skip 1.5 mm) | 230 | 128x128 | 3 | 92/96 | 5144/3600 |  | CV varied from 3.7% to 37% | ADC varied from 643 to 858

## Breast
Author | Title | Journal | Reproducibility assessed? | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Reproducibility Assessed:
Tagliafico, A et al | Diffusion tensor magnetic resonance imaging of the normal breast: reproducibility of DTI-derived fractional anisotropy and apparent diffusion coefficient at 3.0 T | Radiol Med | y | Epub ahead of print | ?? | 2012 | http://www.springerlink.com.offcampus.lib.washington.edu/content/25026586u403124w/fulltext.pdf |  | Normal (contralateral breast in women evaluated for suspicous findings) | Inter/Intraobserver reproducibility (4 observers), Test-retest (coffee break rescan) | 60 (n=11 for test/retest) | 3 | GE | 1 | 2 | 0, 1000 | 32 | No Details Given | DW - EPI | 8-Ch Breast Coil (GE) | No | not specified | 2 | 260 | 96x128 | not given | min | 16675 | Not specified | Inter/Intraobserver CoV=15% (n=60), Test/Retest (n=11 coffee break rescan) CoV=15% | ADC range 1.85 - 2.01x10-3 mm2/s (mean, 1.92x10-3 mm2/s) for normal breast tissue
Partridge, SC et al | Diffusion tensor magnetic resonance imaging of the normal breast | Mag. Reson. Imag. | y | 28 | 320-328 | 2010 | http://www.sciencedirect.com.offcampus.lib.washington.edu/science/article/pii/S0730725X09002835# |  | Normal | Test/retest | n=9 for test/retest | 1.5 | GE | 1 | 2 | 0,600 | 6 | 2 b-values, Monoexponential | SE-EPI | 8-Ch Breast Coil (GE) | No | SPECIAL | 4,5 | 360 | 192x192 | 3 | 71.5 | 7000 | Yes, R=2 | Test/Retest (coffee break) CoV=4.5% | ADC range 1.53 - 2.34x10-3 mm2/s (mean, 1.95x10-3 mm2/s) for normal breast tissue
O'Flynn, EA et al | Diffusion weighted imaging of the normal breast: reproducibility of apparentdiffusion coefficient measurements and variation with menstrual cycle and menopausal status. | Eur. Radiol. | y | 22 | 1512-1518 | 2012 | http://www.springerlink.com.offcampus.lib.washington.edu/content/p233p74326051552/fulltext.pdf |  | Normal | Inter/intraobserver variability | 31 | 3 | Philips | 1 | 3-7 | 0, 100, 150, 200, 350, 700, 1200 |  | all b-vals (ADC total)/ b0 excluded (ADC high) | SS-EPI | 7-Ch Breast Coil | No | SPAIR/SSGR | 2.5,3 | 180 | not specified |  | 67/67 | 1220/3771 | Not specified | Intraobserver ICC = 0.93, interobserver kappa = 0.83 | ADC range 1.78 - 1.89x10-3 mm2/s (mean, 1.84x10-3 mm2/s)for normal breast tissue (all b-vals)
Petralia G. et al | Intraobserver and interobserver variability in the calculation of apparent diffusion coef? cient (ADC) from diffusion-weighted magnetic resonance imaging (DW-MRI) of breast tumours | La Radiologia Medica. | y | 116 | 466-476 | 2011 | http://www.springerlink.com.offcampus.lib.washington.edu/content/w37072166j876016/fulltext.pdf |  | Breast Carcinoma | Inter/intraobserver variability | 28 | 1.5 | Siemens | 1 | 4 | 0, 250, 500, 1000 |  | all b-vals? (assumed) | not specified | 8-Ch Breast Coil (GE) |  | SPAIR | 5 | not specified | 150x90 |  | 71 | 4800 | Yes, R=2 | Intraobserver CoV = 1.1%, interobserver CoV = 2% | ADC mean = 1.07 +/- 0.138x10-3 mm2/s for breast carcinomas; DWI threshold used to select ROI
Scaranelo, AM. et al | Accuracy of Unenhanced MR Imaging in the Detection of Axillary Lymph Node Metastasis: Study of Reproducibility and Reliability | Radiology | y | 262 | 425-434 | 2012 | http://radiology.rsna.org.offcampus.lib.washington.edu/content/262/2/425 |  | Metastatic axillary lymph nodes | Inter/intraobserver variability | 61 | 1.5 | Siemens | 0 | 4 | 50, 300, 700 ,1000 | 3 | all b-vals? (assumed), nonzero bmin | not specified | 8-Ch Breast Coil (GE) |  | not specified | 4 | 408 | Not specified |  | not specified | not specified | yes | Intraobserver Lin CCC = 0.933, interobserver Lin CCC = 0.959 | ADC mean = 0.663-0.698x10-3 mm2/s in malignant lymph nodes, 0.700-0.739x10-3 mm2/s in benign lymph nodes.
Rubesova, E et al | Quantitative diffusion imaging in breast cancer: A clinical prospective study | J. Mag. Reson. Imag. | y | 24 | 319-324 | 2006 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.20643/full | 118 | Lesions (path not specified) | Inter/intraobserver reproducibility (2 observers) | 37 | 1.5 | Siemens | 1 | 5 | 0, 200, 400, 600, 1000 | 3 | all b-vals | DW - EPI | Phased array breast coil | No | not specified | 4 | 160 | 128X128 |  | 110 | 2600 | Not specified | CCC, CoV not calculated. Bland altman plot shows good inter/intraobserver agreement | Malignancies: ADC=0.95x10–3 mm2/s for reader 1, 0.99x10–3 mm2/s for reader 2.Benign lesions: ADC=1.51x10–3 mm2/s for reader 1, 1.47x10–3 mm2/s for reader 2.
Lima M, et al | Apparent diffusion coefficient as an MR imaging biomarker of low risk ductal carcinoma in situ: a pilot study | Radiology | y | 260 | 364-372 | 2011 | http://radiology.rsna.org/content/260/2/364.full.pdf+html |  | DCIS lesions | Interobserver variability | 25 | 1.5 | Philips | 1 | 2 | 0, 1000 | 3 | 2 b-values | DW - EPI | 4-Ch Breast Coil | No | not specified | 5 | 400 | 128x104 | not given | 96 | 8000 | Yes, R=2 | Interobserver ICC=0.72,Pearson's r=0.91 | DCIS: Mean ADC=1.42x10-3 mm2/s for low grade, 1.23x10-3 mm2/s intermediate, 1.19x10-3 mm2/s for high grade, 2.18x10-3 mm2/s for normal tissue
Parsian, S et al (Partridge SC) | Nonmalignant breast lesions: apparent diffusion coefficients of benign and high risk subtypes assessed as false-positives on dynamic contrast-enhanced MR imaging | Radiology | y | In Press | ?? | 2012 |  |  | Benign Lesions | Inter/intraobserver variability | 30 | 1.5 | GE | 1 | 2 | 0, 600 | 6 | 2 b-values, Monoexponential | SE-EPI | 8-Ch Breast Coil (GE, Sentinelle) | No | SPECIAL | 5 | 360 | 192x192 | 3 | 71.5 | 7000 | Yes, R=2 | Intraobserver CCC=0.91, interobserver CCC=0.90 | Benign Lesions (n=175): ADC mean = 1.77 +/-0.45x10-3mm2/s
Reproducibility not Assessed:
Sharma, U et al | Longitudinal study of the assessment by MRI and diffusion‐weighted imaging of tumor response in patients with locally advanced breast cancer undergoing … | NMR in Biomedicine | n | 22 | 104-113 | 2009 | http://onlinelibrary.wiley.com/doi/10.1002/nbm.1245/abstract | 84 |  | Advanced cancer vs. benign tumors | 81 | 1.5 | Siemens |  | 3 | 0, 500, 1000
Partridge, SC et al | Diffusion-weighted MRI: influence of intravoxel fat signal and breast density on breast tumor conspicuity and apparent diffusion coefficient measurements | Mag. Reson. Imag. | n | 29 | 1215-1221 | 2011 | http://www.sciencedirect.com/science/article/pii/S0730725X11002864
Woodhams, R et al | Diffusion-weighted Imaging of the Breast: Principles and Clinical Applications | RadioGraphics | n | 31 | 1059-1084 | 2011 | http://radiographics.rsna.com/content/31/4/1059.short | 5
Woodhams, R et al | Diffusion-Weighted Imaging of Malignant Breast Tumors: The Usefulness of Apparent Diffusion Coefficient (ADC) Value and ADC Map for the Detection of Malignant Breast Tumors and Evaluation of Cancer Extension | J. Comp. Assist. Tomogr. | n | 29 | 644-649 | 2005 | http://journals.lww.com/jcat/Abstract/2005/09000/Diffusion_Weighted_Imaging_of_Malignant_Breast.15.aspx | 140
Yankeelov, TE et al | Integration of quantitativeDCE-MRI and ADCmapping to monitortreatmentresponse in humanbreast cancer: initialresults | Mag. Reson. Imag. | n | 25 | 1-13 | 2007 | http://www.sciencedirect.com/science/article/pii/S0730725X0600289X | 121
Lo, GG et al | Diffusion-Weighted Magnetic Resonance Imaging of Breast Lesions: First Experiences at 3T | J. Comp. Assist. Tomogr. | n | 33 | 63-69 | 2009 | http://journals.lww.com/jcat/Abstract/2009/01000/Diffusion_Weighted_Magnetic_Resonance_Imaging_of.11.aspx | 31
Kuroki, Y and Nasu, K | Advances in breast MRI: diffusion-weighted imaging of the breast | Breast Cancer | n | 15 | 212-217 | 2008 | http://www.springerlink.com/content/ch7r82278x77kq58/ | 26
Tsushima, Y et al | Magnetic resonance (MR) differential diagnosis of breast tumors using apparent diffusion coefficient (ADC) on 1.5‐T | J. Mag. Reson. Imag. | n | 30 | 249-255 | 2009 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21854/full | 30
Park, MJ et al | The Role of Diffusion-Weighted Imaging and the Apparent Diffusion Coefficient (ADC) Values for Breast Tumors | Korean J. Radiol. | n | 8 | 390-396 | 2007 | http://synapse.koreamed.org/DOIx.php?id=10.3348/kjr.2007.8.5.390 | 60
Bogner, W et al | Diffusion-weighted MR for Differentiation of Breast Lesions at 3.0 T: How Does Selection of Diffusion Protocols Affect Diagnosis?1 | Radiology | n | 253 | 341-351 | 2009 | http://radiology.rsna.org/content/253/2/341.short | 46
Partridge, SC et al | Quantitative Diffusion-Weighted Imaging as an Adjunct to Conventional Breast MRI for Improved Positive Predictive Value | Am. J. Roentgenology | n | 193 | 1716-1722 | 2009 | http://www.ajronline.org/content/193/6/1716.short | 29 | [22 empty] | Reproducibility not assessed
Theilmann, RJ et al | Changes in Water Mobility Measured by Diffusion MRI Predict Response of Metastatic Breast Cancer to Chemotherapy | Neoplasia |  | 6 | 831-837 | 2004 | http://www.neoplasia.com/abstract.php?msid=263 | 142 | [22 empty] | Liver mets (not breast measures)
Pickles, MD et al | Diffusion changes precede size reduction in neoadjuvant treatment of breast cancer | Mag. Reson. Imag. | n | 24 | 843-847 | 2006 | http://www.sciencedirect.com/science/article/pii/S0730725X06000907 | 156
Guo, Y et al | Differentiation of clinically benign and malignant breast lesions using diffusion-weighted imaging | J. Mag. Reson. Imag. | n | 16 | 172-178 | 2002 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.10140/full | 307
Yabuuchi, H et al | Enhanced mass on contrast-enhanced breast MR imaging: Lesion characterization using combination of dynamic contrast-enhanced and diffusion-weighted MR images | J. Mag. Reson. Imag. | n | 28 | 1157-1165 | 2008 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21570/full | 52
Marini, C et al | Quantitative diffusion-weighted MR imaging in the differential diagnosis of breast lesion | Eur. Radiol. | n | 17 | 2646-2655 | 2007 | http://www.springerlink.com/content/a43v1910380456r4/ | 79
Kim, SH et al | Diffusion-weighted imaging of breast cancer: Correlation of the apparent diffusion coefficient value with prognostic factors | J. Mag. Reson. Imag. | n | 30 | 615-620 | 2009 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21884/full | 22
Peter, NHGM et al | Quantitative diffusion weighted imaging for differentiation of benign and malignant breast lesions: The influence of the choice of b-values | J. Mag. Reson. Imag. | n | 31 | 1100-1105 | 2010 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22152/full | 9
Khouli, RHE et al | Diffusion-weighted Imaging Improves the Diagnostic Accuracy of Conventional 3.0-T Breast MR Imaging | Radiology | n | 256 | 64-73 | 2010 | http://radiology.rsna.org/content/256/1/64.short | 27
Kuhl, CK et al | Contrast-enhanced MR Imaging of the Breast at 3.0 and 1.5 T in the Same Patients: Initial Experience | Radiology | n | 239 | 666-676 | 2006 | http://radiology.rsna.org/content/239/3/666.short | 85
Yoshikawa, MI et al | Relation between cancer cellularity and apparent diffusion coefficient values using diffusion-weighted magnetic resonance imaging in breast cancer | Radiat. Med | n | 26 | 222-226 | 2008 | http://www.springerlink.com/content/gp6x381515m1825p/ | 33

## General
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Padhani, AR and Ollivier, L | The RECIST criteria: implications for diagnostic radiologists | Brit. J. Radiol. | 74 | 983-986 | 2001 | http://bjr.birjournals.org/content/74/887/983.short | 168
Therasse, P et al | RECIST revisited: A review of validation studies on tumour assessment | Eur. J. Cancer | 42 | 1031-1039 | 2006 | http://www.sciencedirect.com/science/article/pii/S0959804906001511 | 233
Sun, YS et al | Early Evaluation of Cancer Response by a New Functional Biomarker: Apparent Diffusion Coefficient | Am. J. Roentgenology | 197 | W23-W29 | 2011 | http://www.ajronline.org/content/197/1/W23.short | 4
Thoeny, HC and Keyzer, F | Extracranial applications of diffusion-weighted magnetic resonance imaging | Eur. Radiol. | 17 | 1385-1393 | 2007 | http://www.springerlink.com/index/F8641X3Q21073P43.pdf | 120
Marcus, CD et al | Imaging techniques to evaluate the response to treatment in oncology: Currentstandards and perspectives | Crit. Reviews in Onc./Hematology | 72 | 217-238 | 2009 | http://www.sciencedirect.com/science/article/pii/S1040842808001534 | 44
Yanagihara, TK et al | Statistical Power in Quantitative Diffusion MRI of Tumor Response | Acad. Radiology | 19 | 326-330 | 2011 | http://www.sciencedirect.com/science/article/pii/S1076633211005113
Padhani, AR and Khan, AA | Diffusion-weighted (DW) and dynamic contrast-enhanced (DCE) magnetic resonance imaging (MRI) for monitoring anticancer therapy | Targeted Oncol. | 5 | 39-52 | 2010 | http://www.springerlink.com/content/n0656253h4831107/ | 16
Padhani, AR et al | Diffusion-weighted magnetic resonance imaging as a cancer biomarker: consensus and recommendations |  | 11 | 102-125 | 2009 | http://www.neoplasia.com/abstract.php?msid=2119 | 294
Barnhart, HX and Barboriak DP | Applications of the Repeatability of QuantitatiMagnetic Resonance Assessment of Response to Therapy: Tumor Change Measurement, Truth Data and Error Sourcesve Imaging Biomarkers: A Review of Statistical Analysis of Repeat Data Sets | Transl. Oncol. | 2 | 231-235 | 2009 | http://www.ncbi.nlm.nih.gov/pmc/articles/PMC2781067/ | 6
Jackson, EF et al | Magnetic Resonance Assessment of Response to Therapy: Tumor Change Measurement, Truth Data and Error Sources | Transl. Oncol. | 2 | 211-215 | 2009 | http://www.ncbi.nlm.nih.gov/pmc/articles/PMC2781079/ | 9
Hamstra, DA et al | Diffusion magnetic resonance imaging: a biomarker for treatment response in oncology | J. Clin. Oncol. | 25 | 4104-4109 | 2007 | http://jco.ascopubs.org/content/25/26/4104.short | 112
Moffat, BA et al | The Functional Diffusion Map: An Imaging Biomarker for the Early Prediction of Cancer Treatment Outcome | Neoplasia | 8 | 259-267 | 2006 | http://www.ncbi.nlm.nih.gov/pmc/articles/PMC1600674/ | 94
Patterson, DM et al | Technology insight: water diffusion MRI—a potential new biomarker of response to cancer therapy | Nat. Rev. Clin. Oncol. | 5 | 220-233 | 2008 | http://www.nature.com/nrclinonc/journal/v5/n4/full/ncponc1073.html | 93
Koh, DM et al | Reproducibility and changes in the apparent diffusion coefficients of solid tumours treated with combretastatin A4 phosphate and bevacizumab in a two-centre phase I clinical trial | Eur. Radiol. | 19 | 2728-2738 | 2009 | http://www.springerlink.com/content/dw62331861832151/ | 42
Koh, DM and Thoeny, HC | Diffusion-Weighted MR Imaging | Book |  |  | 2010 | http://rd.springer.com/book/10.1007/978-3-540-78576-7/page/1
Sorensen, AG | Magnetic Resonance As a Cancer Imaging Biomarker | J. Clin. Oncol. | 24 | 3274-3281 | 2006 | http://jco.ascopubs.org/content/24/20/3274.short | 45
Thoeny, HC and Ross, BD | Predicting and monitoring cancer treatment response with diffusion‐weighted MRI | J. Mag. Reson. Imag | 32 | 2-16 | 2010 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22167/full | 36
Padhani, AR and Koh, DM | Diffusion MR imaging for monitoring of treatment response | Magn. Reson. Imag. Clin. N. Amer. | 19 | 181-209 | 2011 | http://www.ncbi.nlm.nih.gov/pubmed/21129641 | 19
Meyer, CR et al | Quantitative Imaging to Assess Tumor Response to Therapy: Common Themes of Measurement, Truth Data, and Error Sources | Transl. Oncol. | 2 | 198-210 | 2009 | http://www.ncbi.nlm.nih.gov/pmc/articles/PMC2781075/ | 23
Luna, A et al | Diffusion MRI Outside the Brain: A Case-Based Review and Clinical Applications | Book |  |  |  | http://www.springerlink.com/content/978-3-642-21051-8/#section=990797&page=1

## Gyno
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Harry, VN et al | Diffusion-weighted magnetic resonance imaging in the early detection of response to chemoradiation in cervical cancer | Gynec. Oncol. | 111 | 213-220 | 2008 | http://www.sciencedirect.com/science/article/pii/S0090825808005647 | 68
McVeigh, PZ et al | Diffusion-weighted MRI in cervical cancer | Eur. Radiol. | 18 | 1058-1064 | 2008 | http://www.springerlink.com/content/169062252l350q41/ | 77
Naganawa, S et al | Apparent diffusion coefficient in cervical cancer of the uterus: comparison with the normal uterine cervix | Eur. Radiol. | 15 | 71-78 | 2005 | http://www.springerlink.com/content/xwewmjf6m8gw7kgd/ | 94
Tamai, K et al | Diffusion-weighted MR imaging of uterine endometrial cancer | J. Mag. Reson. Imag. | 26 | 682-687 | 2007 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.20997/full | 152
Chen, J et al | The utility of diffusion-weighted MR imaging in cervical cancer | Eur. J. Radiol. | 74 | e101-e106 | 2010 | http://www.sciencedirect.com/science/article/pii/S0720048X0900223X | 26
Levy, A et al | Interest of diffusion‐weighted echo‐planar MR imaging and apparent diffusion coefficient mapping in gynecological malignancies: A review | J. Mag. Reson. Imag. | 33 | 1020-1027 | 2011 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22546/full | 7
Lin, G et al | Detection of lymph node metastasis in cervical and uterine cancers by diffusion‐weighted magnetic resonance imaging at 3T | J. Mag. Reson. Imag. | 28 | 128-135 | 2008 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21412/full | 78
Kim, JK et al | Feasibility of diffusion‐weighted imaging in the differentiation of metastatic from nonmetastatic lymph nodes: Early experience | J. Mag. Reson. Imag. | 28 | 714-719 | 2008 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21480/full | 68

## Head and Neck
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Galbán, CJ et al | A Feasibility Study of Parametric Response Map Analysis of Diffusion-Weighted Magnetic Resonance Imaging Scans of Head and Neck Cancer Patients for Providing Early Detection of Therapeutic Efficacy | Transl. Oncol. | 18 | 184-190 | 2009 | http://www.ncbi.nlm.nih.gov/pmc/articles/PMC2730136/ | 38
Sumi, M et al | Discrimination of Metastatic Cervical Lymph Nodes with Diffusion-Weighted MR Imaging in Patients with Head and Neck Cancer | AJNR | 24 | 1627-1634 | 2003 | http://www.ajnr.org/content/24/8/1627.short | 157
Razek, AAKA et al | Role of Diffusion-Weighted Magnetic Resonance Imaging in Differentiation between the Viable and Necrotic Parts of Head and Neck Tumors | Acta. Radiol. | 49 | 364-370 | 2008 | http://informahealthcare.com/doi/abs/10.1080/02841850701777390 | 27
Razek, AAKA et al | Role of diffusion-weighted MR imaging in cervical lymphadenopathy | Eur. Radiol. | 16 | 1468-1477 | 2006 | http://www.springerlink.com/content/wp549978280g211h/ | 157
Vandecaveye, V et al | Detection of head and necksquamouscellcarcinoma with diffusionweightedMRIafter (chemo)radiotherapy: Correlation between radiologic and histopathologicfindings | Int. J. Rad. Oncol.Biol.Phys. | 67 | 960-971 | 2007 | http://www.sciencedirect.com/science/article/pii/S0360301606029890 | 92
Kato, H et al | Head and neck squamous cell carcinoma: usefulness of diffusion-weighted MR imaging in the prediction of a neoadjuvant therapeutic effect | Eur. Radiol. | 19 | 103-109 | 2009 | http://www.springerlink.com/content/p132570682515218/ | 32
Hermans, R and Vandecaveye, V | Diffusion-weighted MRI in head and neck cancer | Cancer Imag. | 7 | 126-127 | 2007 | http://www.ncbi.nlm.nih.gov/pmc/articles/PMC1974861/ | 17
King, AD et al | Malignant Cervical Lymphadenopathy: Diagnostic Accuracy of Diffusion-weighted MR Imaging | Radiology | 245 | 806-813 | 2007 | http://radiology.rsna.org/content/245/3/806.short | 89
Vandecaveye, V et al | Predictive value of diffusion-weighted magnetic resonance imaging during chemoradiotherapy for head and neck squamous cell carcinoma | Eur. Radiol. | 20 | 1703-1714 | 2010 | http://www.springerlink.com/content/m866712012732814/ | 32
Srinivasan, A et al | Differentiation of benign and malignant pathology in the head and neck using 3T apparent diffusion coefficient values: early experience | AJNR | 29 | 40-44 | 2008 | http://www.ajnr.org/content/29/1/40.short | 48
Kwee, TC et al | ADC measurements of lymph nodes: Inter- and intra-observerreproducibilitystudy and an overview of the literature | Eur. J. Radiol. | 75 | 215-220 | 2010 | http://www.sciencedirect.com/science/article/pii/S0720048X09001442 | 33
Kim, S et al | Diffusion-Weighted Magnetic Resonance Imaging for Predicting and Detecting Early Response to Chemoradiation Therapy of Squamous Cell Carcinomas of the Head and Neck | Clin. Cancer Res. | 15 | 986-994 | 2009 | http://clincancerres.aacrjournals.org/content/15/3/986.short | 76
Razek, AAKA et al | Characterization of pediatric head and neck masses with diffusion-weighted MR imaging | Eur. Radiol. | 19 | 201-208 | 2009 | http://www.springerlink.com/content/c86u3n05810057l4/ | 20
Chen X, et al. | Role of periodically rotated overlapping parallel lines with enhanced reconstruction diffusion-weighted imaging in correcting distortion and evaluating head and neck masses using 3 T MRI. | Clin Radiol | 69 | 403-9 | 2014 | http://www.sciencedirect.com.libproxy.usc.edu/science/article/pii/S0009926013005412 | 0 | [21 empty] | No but compared PROPELLER vs EP DWI
Zheng D, et al. | Early assessment of induction chemotherapy response of nasopharyngeal carcinoma by pretreatment diffusion-weighted magnetic resonance imaging. | J Comput Assist Tomogr | 37 | 673-80 | 2013 | http://journals.lww.com/jcat/pages/articleviewer.aspx?year=2013&issue=09000&article=00005&type=abstract | 0 | [21 empty] | No
Hong J, etal. | Value of magnetic resonance diffusion-weighted imaging for the prediction of radiosensitivity in nasopharyngeal carcinoma. | Otolaryngol Head Neck Surg | 149 | 707-13 | 2013 | http://oto.sagepub.com.libproxy.usc.edu/content/149/5/707.long | 0 | [21 empty] | No
Shang DS, et al. | Differentiating laryngeal carcinomas from precursor lesions by diffusion-weighted magnetic resonance imaging at 3.0 T: a preliminary study. | PLoS One | 8 | e68622 | 2013 | http://www.plosone.org/article/info%3Adoi%2F10.1371%2Fjournal.pone.0068622 | 1 | [21 empty] | No
Shi, et al. | Utility of diffusion-weighted imaging in differentiating malignant from benign thyroid nodules with magnetic resonance imaging and pathologic correlation. | J Comput Assist Tomogr. | 37 | 505-10 | 2013 | http://journals.lww.com/jcat/pages/articleviewer.aspx?year=2013&issue=07000&article=00005&type=abstract | 0 | [21 empty] | No
Hwang, et al. | Differentiation of recurrent tumor and posttreatment changes in head and neck squamous cell carcinoma: application of high b-value diffusion-weighted imaging. | AJNR | 34 | 2343-8 | 2013 | http://www.ajnr.org.libproxy.usc.edu/content/34/12/2343.long | 0 | [21 empty] | No
Verhappen MH, et al. | Diffusion-weighted MR imaging in head and neck cancer: comparison between half-fourier acquired single-shot turbo spin-echo and EPI techniques. | AJNR | 33 | 1239-46 | 2012 | http://www.ajnr.org.libproxy.usc.edu/content/33/7/1239.full.pdf+html | 4 | [21 empty] | No
Yamashita K, et al. | High-resolution three-dimensional diffusion-weighted imaging of middle ear cholesteatoma at 3.0 T MRI: usefulness of 3D turbo field-echo with diffusion-sensitized driven-equilibrium preparation (TFE-DSDE) compared to single-shot echo-planar imaging. | Eur J Radiol | 82 | e471-5 | 2013 | http://zb5lh7ed7a.search.serialssolutions.com.libproxy.usc.edu/uschsl?sid=Entrez:PubMed&id=pmid:23701953 | 0 | [21 empty] | No
Aoyagi T, et al. | Apparent diffusion coefficient correlation with oesophageal tumour stroma and angiogenesis. | Eur Radiol | 22 | 1171-7 | 2012 | http://download.springer.com.libproxy.usc.edu/static/pdf/171/art%253A10.1007%252Fs00330-011-2359-0.pdf?auth66=1411105019_439b86035b0b7bb44fe9207d681aa0d2&ext=.pdf | 4 | [21 empty] | No
Mutlu H, et al. | Role of apparent diffusion coefficient values and diffusion-weighted magnetic resonance imaging in differentiation between benign and malignant thyroid nodules. | Clin Imaging | 36 | 2014-01-07 | 2012 | http://ac.els-cdn.com.libproxy.usc.edu/S0899707111000805/1-s2.0-S0899707111000805-main.pdf?_tid=3e9fbd04-3e2d-11e4-9c73-00000aacb35f&acdnat=14109326 | 4 | [21 empty] | No
Ichikawa Y, et al. | Efficacy of diffusion-weighted imaging for the differentiation between lymphomas and carcinomas of the nasopharynx and oropharynx: correlations of apparent diffusion coefficients and histologic features | AJNR | 33 | 761-6 | 2012 | http://www.ajnr.org.libproxy.usc.edu/content/33/4/761.long | 1 | [21 empty] | No
Wu LM, et al. | Value of magnetic resonance imaging for nodal staging in patients with head and neck squamous cell carcinoma: a meta-analysis | Acad Radiol | 19 | 331-40 | 2012 | http://www-ncbi-nlm-nih-gov.libproxy.usc.edu/pubmedhealth/PMH0041220/ | 2 | [21 empty] | No
Park JK, et al. | High-resolution diffusion-weighted imaging of neck lymph nodes using 2D-single-shot interleaved multiple inner volume imaging diffusion-weighted echo-planar imaging at 3T. | AJNR | 32 | 1173-7 | 2011 | http://www.ajnr.org.libproxy.usc.edu/content/32/7/1173.long | 1 | [21 empty] | No
Bhatia KS, et al. | Can diffusion-weighted imaging distinguish between normal and squamous cell carcinoma of the palatine tonsil? | Br J Radiol | 83 | 753-8 | 2010 | http://www.birpublications.org.libproxy.usc.edu/doi/abs/10.1259/bjr/58331222?url_ver=Z39.88-2003&rfr_id=ori:rid:crossref.org&rfr_dat=cr_pub%3dpubmed | 0 | [21 empty] | No
Vandecaveye V, et al. | Applications of diffusion-weighted magnetic resonance imaging in head and neck squamous cell carcinoma. | Neuroradiology | 52 | 773-84 | 2010 | http://www-ncbi-nlm-nih-gov.libproxy.usc.edu/pubmed/20631998 | 5 | [21 empty] | No
Fong D, et al. | Diagnostic accuracy of diffusion-weighted MR imaging for nasopharyngeal carcinoma, head and neck lymphoma and squamous cell carcinoma at the primary site. | Oral Oncol | 46 | 603-6 | 2010 | http://www.sciencedirect.com.libproxy.usc.edu/science/article/pii/S1368837510001661 | 1 | [21 empty] | No
King, AD et al | Squamous cell carcinoma of the head and neck: diffusion-weighted MR imaging for prediction and monitoring of treatment response. | Eur Radiol | 20 | 2213-20 | 2010 | http://link.springer.com.libproxy.usc.edu/article/10.1007%2Fs00330-010-1769-8 | 11 | [21 empty] | No
Perrone A, et al. | Diffusion-weighted MRI in cervical lymph nodes: differentiation between benign and malignant lesions. | Eur J Radiol | 77 | 281-6 | 2011 | http://www.sciencedirect.com.libproxy.usc.edu/science/article/pii/S0720048X0900477X | 5 | [21 empty] | No
Aschenbach R, et al. | Diffusion-weighted imaging and ADC mapping of head-and-neck paragangliomas: initial experience. | Klin Neuroradiol | 19 | 215-9 | 2009 |  | 1 | [21 empty] | No
Dirix P, et al. | Dose painting in radiotherapy for head and neck squamous cell carcinoma: value of repeated functional imaging with (18)F-FDG PET, (18)F-fluoromisonidazole PET, diffusion-weighted MRI, and dynamic contrast-enhanced MRI. | J Nucl Med | 50 | 1020-7 | 2009 | http://jnm.snmjournals.org.libproxy.usc.edu/content/50/7/1020.long | 11 | [21 empty] | No
Akter M, et al. | Diffusion tensor tractography in the head-and-neck region using a clinical 3-T MR scanner. | Acad Radiol | 16 | 858-65 | 2009 | http://www.sciencedirect.com.libproxy.usc.edu/science/article/pii/S1076633209000531 | 2 | [21 empty] | No
Vandecaveye, V et al | Head and neck squamous cell carcinoma: value of diffusion-weighted MR imaging for nodal staging. | Radiology | 251 | 134-46 | 2009 | http://pubs.rsna.org.libproxy.usc.edu/doi/abs/10.1148/radiol.2511080128?url_ver=Z39.88-2003&rfr_id=ori:rid:crossref.org&rfr_dat=cr_pub%3dpubmed | 18 | [21 empty] | No
Friedrich KM, et al. | Diffusion-weighted magnetic resonance imaging of head and neck squamous cell carcinomas | Eur J Radiol | 68 | 493-8 | 2008 | http://www-ncbi-nlm-nih-gov.libproxy.usc.edu/pubmed/19189427 | 2 | [21 empty] | No
Sakurada A, et al. | Diagnostic performance of diffusion-weighted magnetic resonance imaging in esophageal cancer. | Eur Radiol | 19 | 1461-9 | 2009 | http://link.springer.com.libproxy.usc.edu/article/10.1007%2Fs00330-008-1291-4 | 6 | [21 empty] | No
Bozgeyik Z, et al. | Diffusion-weighted MR imaging of thyroid nodules. | Neuroradiology | 51 | 193-8 | 2009 | http://link.springer.com.libproxy.usc.edu/article/10.1007%2Fs00234-008-0494-3 | 6 | [21 empty] | No
de Bondt RB, et al. | Diagnostic accuracy and additional value of diffusion-weighted imaging for discrimination of malignant cervical lymph nodes in head and neck squamous cell carcinoma. | Neuroradiology | 51 | 183-92 | 2009 | http://link.springer.com.libproxy.usc.edu/article/10.1007%2Fs00234-008-0487-2 | 6 | [21 empty] | No
Holzapfel K, et al. | Value of diffusion-weighted MR imaging in the differentiation between benign and malignant cervical lymph nodes | Eur J Radiol | 72 | 381-7 | 2009 | http://www.sciencedirect.com.libproxy.usc.edu/science/article/pii/S0720048X08005147 | 10 | [21 empty] | No
Schueller-Weidekamm C, et al | Can quantitative diffusion-weighted MR imaging differentiate benign and malignant cold thyroid nodules? Initial results in 25 patients. | AJNR | 30 | 417-22 | 2009 | http://www.ajnr.org.libproxy.usc.edu/content/30/2/417.long | 2 | [21 empty] | No
Sakamoto J, et al. | Tissue characterization of head and neck lesions using diffusion-weighted MR imaging with SPLICE. | Eur J Radiol | 69 | 260-8 | 2009 | http://www-ncbi-nlm-nih-gov.libproxy.usc.edu/pubmed/18023549 | 4 | [21 empty] | No
Abdel Razek AA, et al. | Role of diffusion-weighted echo-planar MR imaging in differentiation of residual or recurrent head and neck tumors and posttreatment changes. | AJNR | 28 | 1146-52 | 2007 | http://www-ncbi-nlm-nih-gov.libproxy.usc.edu/pubmed/17569975 | 15 | [21 empty] | No
Habermann CR, et al. | Diffusion-weighted echo-planar MRI: a valuable tool for differentiating primary parotid gland tumors? | Rofo | 177 | 940-5 | 2005 | http://www-ncbi-nlm-nih-gov.libproxy.usc.edu/pubmed/15973595 | 7 | [21 empty] | No
Hoang JK, et al. | Diffusion-weighted imaging for head and neck squamous cell carcinoma: quantifying repeatability to understand early treatment-induced change. | AJR | 203 | 1104-8 | 2014 |  | 0 | [21 empty] | Yes | Can't get access to full text for some reason. Seems like single center, but can't tell how many scanners used fro abstract. Done in head and neck squamous cell carcinoma patients. Patients imaged with two baseline DWI examinations 1 week apart and a third DWI examination during the 2nd week of curative-intent chemoradiation therapy. Sixteen patients underwent double baseline imaging, with 14 also undergoing intratreatment imaging. Baseline nodal disease ADC could be measured in 16 patients, but ADC in primary tumors could only be measured in five patients. The nodal mean (SD) baseline %∆ADC was 8% (± 7%), which was significantly different compared with intratreatment changes of 32% (± 31%) (p = 0.01). Baseline ICC was 0.86 for nodal disease and 0.99 for primary tumor (excellent correlation). The calculated repeatability coefficient for baseline nodal ADC was 15%. No patients had decreases in intratreatment ADC of more than 15%. CONCLUSION: Baseline ADC variability for HNSCC is less than intratreatment ADC change for nodal disease. Assessment of response should consider intrinsic baseline variability.
Kolff-Gart AS, et al. | Diffusion-Weighted Imaging of the Head and Neck in Healthy Subjects: Reproducibility of ADC Values in Different MRI Systems and Repeat Sessions. | AJNR | 36 | 384-90 | 2015 |  | 0 | [21 empty] | Yes | Done in 7 healthy subjects on 5 MR scanners (4 1.5T, 1 3T) in 2 centers. ADC maps of EPI (with 2 b-values and 6 b-values) and TSE sequences were compared. 3 time points obtained (Two examinations were performed on the same day (between examinations, the subject was removed from the MR imaging system), and 1 examination, at least 1 month later.) Mean ADC values for different tissues (submandibular gland, sternocleidomastoid muscle, spinal cord, subdigastric lymph node, and tonsil) were used to evaluate intra and intersubject, intersystem, and intersequence variability by using a linear mixed model. The standard error of ADCmeasurement was the smallest for the tonsil and spinal cord (standard error of measurement 151.2 10 6 mm/s2 and 190.1 10 6 mm/s2, respectively). The intersystem difference for mean ADC values and the influence of the MR imaging system on ADC values among the subjects were statistically significant (P .001). The mean difference among examinations was negligible (ie, 10 10 6 mm/s2). CONCLUSIONS: In this study, the spinal cord was the most appropriate reference tissue and EPI-DWI with 6 b-values was the most reproducible sequence. ADC values were more precise if subjects were measured on the same MR imaging system and with the same sequence. ADC values differed significantly between MR imaging systems and sequences.
Matoba M, et al. | Fractional change in apparent diffusion coefficient as an imaging biomarker for predicting treatment response in head and neck cancer treated with chemoradiotherapy. | AJNR | 35 | 379-85 | 2013 |  | 1 | [21 empty] | No
Sumi, M et al. | Head and neck tumours: combined MRI assessment based on IVIM and TIC analyses for the differentiation of tumors of different histological types | Eur Radiol | 24 | 223-31 | 2014 |  | 1 | [21 empty] | No
Varoquaux A, et al. | Functional imaging of head and neck squamous cell carcinoma with diffusion-weighted MRI and FDG PET/CT: quantitative analysis of ADC and SUV. | Eur J Nucl Med Mol Imaging | 40 | 842-52 | 2013 |  | 7 | [21 empty] | Yes | Single center-multiscanner, retrospective analysis of H and N squamous cell cancer patients. 23 patients on 1.5T Siemens Espree and 11 on 3T Trio Siemens. DWI protocol = single-shot echo planar technique with fat suppression on both machines. 2 experienced radiologists measured tumor ADC independently. Intra- and inter-observer agreement for ADC measurements was assessed. Intraclass correlation coefficient (ICC) analysis showed almost perfect reproducibility (>0.90) for ADCmean, ADCmin values for intra-observer and inter-observer agreement. Mean ADCmean and ADCmin in HNSCC were 1.05±0.34 × 10−3mm2/s and 0.65±0.29 × 10−3mm2/s, respectively. Conclusion Our data suggest that ADC and SUV values are reproducible and independent biomarkers in HNSCC (I omitted SUV data here).
Hauser T, et al. | Characterization and therapy monitoring of head and neck carcinomas using diffusion-imaging-based intravoxel incoherent motion parameters-preliminary results. | Neuroradiology | 55 | 527-36 | 2013 |  | 2 | [21 empty] | No
Li PM, et al. | Evaluating the utility of non-echo-planar diffusion-weighted imaging in the preoperative evaluation of cholesteatoma: a meta-analysis. | Laryngoscope | 123 | 1247-50 | 2013 |  | 0 | [21 empty] | No
Kato H, et al. | Necrotic cervical nodes: usefulness of diffusion-weighted MR imaging in the differentiation of suppurative lymphadenitis from malignancy. | Eur J Radiol | 82 | e28-35 | 2012 |  | 2 | [21 empty] | No
Lu Y, et al. | Extension of the intravoxel incoherent motion model to non-gaussian diffusion in head and neck cancer. | J Magn Reson Imaging | 36 | 1088-96 | 2012 |  | 3 | [21 empty] | No
Ahn SJ, et al. | Histogram analysis of apparent diffusion coefficient map of standard and high B-value diffusion MR imaging in head and neck squamous cell carcinoma: a correlation study with histological grade. | Acad Radiol | 19 | 1233-40 | 2012 |  | 1 | [21 empty] | No
Hatakenaka M, et al. | Apparent diffusion coefficient calculated with relatively high b-values correlates with local failure of head and neck squamous cell carcinoma treated with radiotherapy. | AJNR | 32 | 1904-10 | 2011 |  | 0 | [21 empty] | No
Huins CT, et al. | Detecting cholesteatoma with non-echo planar (HASTE) diffusion-weighted magnetic resonance imaging. | Otolaryngol Head Neck Surg | 132 | 141-6 | 2010 |  | 0 | [21 empty] | No
Sumi M, et al. | Diagnostic importance of focal defects in the apparent diffusion coefficient-based differentiation between lymphoma and squamous cell carcinoma nodes in the neck. | Eur Radiol | 19 | 975-81 | 2009 |  | 2 | [21 empty] | No
Abdel Razek AA, et al. | Characterization of pediatric head and neck masses with diffusion-weighted MR imaging. | Eur Radiol | 19 | 201-8 | 2009 |  | 11 | [21 empty] | No
Razek, AA et al | Role of diffusion-weighted magnetic resonance imaging in differentiation between the viable and necrotic parts of head and neck tumors. | Acta. Radiol. | 49 | 364-70 | 2008 |  | 11 | [21 empty] | No
Suzuki C, et al. | Apparent diffusion coefficient of subcutaneous epidermal cysts in the head and neck comparison with intracranial epidermoid cysts. | Acad Radiol | 14 | 1020-8 | 2007 |  | 0 | [21 empty] | No

## Kidney
Author | Title | Journal | Volume | Pages | Year | URL | Citations | # of Subj | Vendor and Model | Field strength | Imaging Sequence | TR/TE(ms) | Gating | Orientation | Sl Thk (mm) | FOV (mmxmm) | Matrix | Lipid Suppresion | Parallel Imaging? | Directions | b-values (s/mm^2) | CV(Cortex) | CV(Medulla) | Test/Retest or Confidence Interval | Notes
Notohamiprodjo, M et al | Diffusion Tensor Imaging (DTI) of the Kidney at 3 Tesla–Feasibility, Protocol Evaluation and Comparison to 1.5 Tesla | Inv. Radiol. | 45 | 245-254 | 2010 | http://journals.lww.com/investigativeradiology/Abstract/2010/05000/Diffusion_Tensor_Imaging__DTI__of_the_Kidney_at_3.3.aspx | 17
Cutajar, M et al | Test–retest reliability and repeatability of renal diffusion tensor MRI in healthy subjects | Eur. J. Radiol. | 80 | e263-e268 | 2011 | http://www.sciencedirect.com/science/article/pii/S0720048X10006121 | 3 | 7 | Siemens TIM Avanto | 1.5t | MS-EPI | 730/74 |  | Obl Cor | 5 |  |  | Yes | 2 |  | 0,400 | 3.44% | 3.56%
Kılıçkesmez, O et al | Diffusion-Weighted Imaging of the Renal and Adrenal Lesions | J. Comp. Assis. Tomog. | 33 | 828-833 | 2009 | http://journals.lww.com/jcat/Abstract/2009/11000/Diffusion_Weighted_Imaging_of_the_Renal_and.3.aspx | 20
Notohamiprodjo, M et al | Diffusion Tensor Imaging of the Kidney With Parallel Imaging: Initial Clinical Experience | Inv. Radiol. | 43 | 677-685 | 2008 | http://journals.lww.com/investigativeradiology/Abstract/2008/10000/Diffusion_Tensor_Imaging_of_the_Kidney_With.1.aspx | 49
Ries, M et al | Diffusion tensor MRI of the human kidney | J. Mag. Reson. Imag. | 14 | 42-49 | 2001 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.1149/full | 139
Gürses, B et al | Diffusion tensor imaging of the kidney at 3 Tesla MRI: normative values and repeatability of measurements in healthy volunteers | Diagn. Interv. Radiol. | 17 | 317-322 | 2011 | http://www.ncbi.nlm.nih.gov/pubmed/21108183 | 4 | 44 | Philips Achieva | 3 | ss-EPI | 10000/60 | No | COR | 3 | 218x167 ? | 132X100 | No | 2 | Nex= 2, 16 dir DTI | 0,700 |  |  | Intra-observer p=0.46; inter-observer p=0.66
Cova, M et al | Diffusion-weighted MRI in the evaluation of renal lesions: preliminary results | Brit. J. Radiol. | 77 | 851-857 | 2004 | http://bjr.birjournals.org/content/77/922/851.short | 138
Xu, X et al | Diffusion-weighted MR imaging of kidneys in patients with chronic kidney disease: initial study | Eur. Radiol. | 20 | 978-983 | 2010 | http://www.springerlink.com/content/nk6q8261238p1152/ | 13
Yoshikawa, T et al | ADC Measurement of Abdominal Organs and Lesions Using Parallel Imaging Technique | Am. J. Roent. | 187 | 1521-1530 | 2006 | http://www.ajronline.org/content/187/6/1521.short | 145
Thoeny, HC et al | Diffusion-weighted MR Imaging of Native and Transplanted Kidneys | Radiology | 259 | 25-38 | 2011 |  |  | N/A | Non-specific | Non-specific | ss-EPI | >2500/71 | Resp | Cor | 5 | 400x400 | 128x128 | No | 2 | NEX=3-6 | 0-900 |  |  |  | A reviw article on Kidney DWI
Thoeny, HC et al | Diffusion-weighted MR Imaging of Kidneys in Healthy Volunteers and Patients with Parenchymal Diseases: Initial Experience | Radiology | 235 | 911-917 | 2005 | http://radiology.rsna.org/content/235/3/911.short | 182
Taouli, B et al | Renal Lesions: Characterization with Diffusion-weighted Imaging versus Contrast-enhanced MR Imaging | Radiology | 251 | 398-407 | 2009 | http://radiology.rsna.org/content/251/2/398.short | 71
Zhang, JL et al | Variability of Renal Apparent Diffusion Coefficients: Limitations of the Monoexponential Model for Diffusion Quantification | Radiology | 254 | 783-792 | 2010 | http://radiology.rsna.org/content/254/3/783.short | 31 | 3 | Siemens TIM Trio | 3T | SS-EPI |  | No | Cor | 6 | 410x345 | 192x162 |  | 2 |  | 0 - 1300 @ 50 inc
Giannarini, G et al | Potential and Limitations of Diffusion-Weighted Magnetic Resonance Imaging in Kidney, Prostate, and Bladder Cancer Including Pelvic Lymph Node Staging: A Critical Analysis of the Literature | Eur. Urology | 61 | 326-340 | 2012 | http://www.sciencedirect.com/science/article/pii/S0302283811010268 | 7
Damasio, MB et al | Diffusion-weighted MRI sequences (DW-MRI) of the kidney: normal findings, influence of hydration state and repeatability of results | La Radiologia Medica | 113 | 212-224 | 2008 | http://www.springerlink.com/content/7x1805nj58267704/ | 11
Kido, A et al | Diffusion Tensor MRI of the kidney at 3.0 and 1.5 Tesla | ACTA RADIOLOGICA | 9 | 1060-1063 | 2010 |  |  | 16 | Siemens Trio & Symphony | 3T and 1.5T | ss-EPI | 2000/74 | Resp | Cor | 3 | 320x? | 128x? | No | 2 | NEX=3, 6 dir | 0, 400 |  |  |  | a small but signifi cant difference in FA and ADC values small but signifi cant difference in FA and ADC values between 1.5 and 3.0 T
Wittsack, H et al | Statistical Evaluation of Diffusion-Weighted Imaging of the Human Kidney | MRM | 64 | 616-622 | 2010 |  |  | 5 | Siemens Trio | 3T | ss-EPI | 3000/75 | No | Cor |  | 400x400 | 192x192 | No | 2 | NEX=3, 3 dir | 0-750 |  |  |  | most voxels in the kidney show biexponential Most voxels in the kidney show biexponential behavior.
Zhang, JL et al | Optimization of b-Value Sampling for Diffusion-Weighted Imaging of the Kidney | MRM | 67 | 89-97 | 2012 |  |  | 4 | Siemens Trio | 3T | ss-EPI | 2000/80 |  | Cor | 6 | 410x345 | 192x162 | No | 2 | NEX=9, 3 dir | 0-750 @ 50 inc |  |  |  | Optimal values for 4 b-values: 0,52,268,750
MIQUEL, ME et al | In vitro and in vivo repeatability of abdominal diffusion-weighted Mri | Brit. J. Radiol. | Online before print |  | 2012 |  |  | 10 | Philips | 1.5T | ss-EPI | 5800/62 | No | Ax | 6 | 400x300 | 256x112 | No | ? | NEX=3 and 3 dir | 0-1000 mono fit | coefficient of reproducibility: 7.7% (VOI) 20% (ROI)
Bilgili | Reproductibility of apparent diffusion coefficients measurements in diffusion-weighted MRI of the abdomen with different b values | Eur. J. Radiol. | 81 | 2066-2068 | 2012 |  |  | 11 | Philips Infinion | 1.5T | ss-EPI | 2500/86 | BH | Ax | 8 | 300x? | 28x128 | SPIR | ? | NEX=2 | 0,100,500 | 10.4% (using b=500), 14.7% (using b=100)

## Liver
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of subjects | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Respiratory gating vs BH vs Free breathing | Lipid Suppresion | Slice Thickness/Gap (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Cui, Y et al | Apparent Diffusion Coefficient: Potential Imaging Biomarker for Prediction and Early Detection of Response to Chemotherapy in Hepatic Metastases | Radiology | 248 | 894-900 | 2008 | http://radiology.rsna.org/content/248/3/894.short | 213 | D | Response liver mets to therapy | 23 | 1.5 | GE |  | 2 | 0-800 | 1 | GE Functool | SSEPI | PA | BH | Yes | 6/1.0 | 36-40 | 128x90 | 1 | ? | 2750 | Yes | No
Bruegel, M et al | Characterization of focal liver lesions by ADC measurements using a respiratory triggered diffusion-weighted single-shot echo-planar MR imaging technique | Eur. Radiol. | 18 | 477-485 | 2008 | http://www.springerlink.com/content/80857p41747458n0/ | 143 | D | Lesion characterization | 102 | 1.5 | Siemens |  | 3 | 50-300-600 | 3 | Siemens Syngo | SSEPI | PA | Navigator triggered | Yes | 5/0.5 | 26-35 | 144x192 | 3 | 69 | 1 respiratory cycle | Yes | No
Lu, TLC et al | Interobserver and intraobserver variability of the apparent diffusion coefficient in treated malignant hepatic lesions on a 3.0T machine: Measurements in the whole lesion versus in the area with the most restricted diffusion | J. Mag. Reson. Imag. | 32 | 647-653 | 2010 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22301/full | 3 | D | Intra and inter-observer variability | 25 | 3 | Siemens |  | 3 | 0-300-600 | 3 | Siemens Syngo | SSEPI | PA | FB | Yes | 6/0 |  | voxel 3.4x2.7x6.0 | 5 | 72 | 3100 | Yes | Yes | Variability -0.17 to +0.17
Bilgili, Mirace YK | Reproductibility of apparent diffusion coefficients measurements in diffusion-weighted MRI of the abdomen with different b values | Eur. J. of Radiol. |  |  | 2011 | http://www.sciencedirect.com/science/article/pii/S0720048X11005869 |  | N | reproducibility ADC liver spleen kidney | 11 | 1.5 | Philips |  | 3 | 0-100-500 |  | Philips | SSEPI | PA | BH | Yes | 8/1.0 | 300 | 128x128 | 2 | 86 | 2500 | ? | Yes | CV values range 7.3% and 14.7%.
Braithwaite, AC et al | Short- and midterm reproducibility of apparent diffusion coefficient measurements at 3.0-T diffusion-weighted imaging of the abdomen | Radiology | 250 | 459-465 | 2009 | http://radiology.rsna.org/content/250/2/459.long | 41 | N | reproducibility ADC liver spleen pancreas | 12 | 3 | Siemens |  | 2 | 0-400 | 3 | Siemens Syngo | SSEPI | PA | FB | Yes | 6/? |  | 108X128 | 4 | 76 | 3200 | Yes | Yes | mean CV for ADC 14%
Rosenkrantz, M et al | Diffusion‐weighted imaging of the abdomen at 3.0 Tesla: Image quality and apparent diffusion coefficient reproducibility compared with 1.5 Tesla | J. Mag. Reson. Imag. | 33 | 128-135 | 2011 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22395/full | 13 | N | image quality, ADC values and ADC reproucibility in liver, renal cortices, pancreas, and spleen | 8 | 1.5 and 3 | Siemens | 6 | 11 | 0-50-100-150-200-300-400-500-600-700-800 | 3 | Siemens Syngo | SSEPI |  | BH(5) and FB(1) | Yes | 7/1.4 | 262x350 | 144x192 | 1 | 76 | 1900 | Yes | Yes
Kwee, TC et al | Comparison and reproducibility of ADC measurements in breathhold, respiratory triggered, and free‐breathing diffusion‐weighted MR imaging of the liver | J. Mag. Reson. Imag. | 28 | 1141-1148 | 2008 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21569/full | 70 | N | ADC reproducibility in liver | 11 | 1.5 | Philips | 1 | 2 | 0-500 | 3 | Philips | SSEPI | PA | RG, BH and | Yes | 7/1.0 (BH, RT and FB) anf 5/0 (FB) | 360x292 | 208x208 | 5 | 70 | 2247 | Yes | Yes
Mo, YH et al | Hepatic ADC value correlates with cirrhotic severity of patients with biliary atresia | Eur. J. of Radiol. | 80 | e253-e257 | 2011 | http://www.sciencedirect.com/science/article/pii/S0720048X10005474 | 1
Mannelli, l et al | Assessment of Tumor Necrosis of Hepatocellular Carcinoma After Chemoembolization: Diffusion-Weighted and Contrast-Enhanced MRI With Histopathologic Correlation of the Explanted Liver | Am. J. of Roent. | 193 | 1044-1052 | 2009 | http://www.ajronline.org/content/193/4/1044.short | 35 | D | assessment of necrosis of HCC after TACE | 21 | 1.5 | Siemens |  | 3 | 0-50-500 | 3 | Siemens Syngo | SSEPI | PA | BH and RT | Yes | 7/1.4 | 300-400mm2 | 144x192 | 2(BH) and 4(RT) | 67-82 | 1300(BH), 2600(RT) | Yes | No
Miquel, ME et al | In vitro and in vivo repeatability of abdominal diffusion-weighted MRI | Brit. J. of Radiol. |  |  | 2012 | http://bjr.birjournals.org/content/early/2012/06/06/bjr.32269440.abstract |  | N | repeatability of ADC in liver, spleen, kidney, gallbladder | 10 | 1.5 | Philips |  | 5 | 100-200-500-750-1000 | 3 | Philips | EPI |  | FB |  | 6/1.0 | 400-450 x 400-450 | 112x256 | 3 | 62 | 5300-5800 | No | No
Eccles, CL et al | Change in diffusion weighted MRI during liver cancer radiotherapy: Preliminary observations | Acta Oncologica | 48 | 1034-1043 | 2009 | http://informahealthcare.com/doi/abs/10.1080/02841860903099972 | 22 | D | evaluation of changes in ADC with conformal radiotherapy | 11 | 1.5 | GE |  | 2 | 0-600 |  | GE | EPI and T2-weighted fast spin-echo | PA | BH | No and Yes | 8/0.0 and 6/0.0 | 37.9-43.8 | 128x128 and 256x160 | 2 and 1 | 67.8-96.8 and 90 | 3000-3775 and 3000 | No | No
Colagrande, S et al | MR‐diffusion weighted imaging of healthy liver parenchyma: Repeatability and reproducibility of apparent diffusion coefficient measurement | J. Mag. Reson. Imag. | 31 | 912-920 | 2010 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22117/full | 8 | N | repeatability and reproducibility of ADC in liver | 30 | 1.5 | Philips |  | 2 | 0-1000 |  |  | SSEPI | PA | BH | Yes | 9/1.0 | 300-420 | 128x64 | 2 | 66 | 2000 | No | Yes
Kim, SY et al | Reproducibility of measurement of apparent diffusion coefficients of malignant hepatic tumors: Effect of DWI techniques and calculation methods | J. Mag. Reson. Imag. |  |  | 2012 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.23744/full |  | D | ADC calculation methods, selection of b-values and reproducibility on malignant hepatic tumors | 19 | 1.5 | Siemens |  | 3 and 5 | 0-50-500 and 0-50-300-500-1000 | 3 | Matlab (4D Analyzer) | SSEPI | 6-channel torso array | BH and RT | Yes | 6/1.8 and 6/1.2 | 350x240-280 | 122x150 | 2 and 4 | 73 | 1500 and 2900 | Yes | Yes
Kim, SY et al | Malignant Hepatic Tumors: Short-term Reproducibility of Apparent Diffusion Coefficients with Breath-hold and Respiratory-triggered Diffusion-weighted MR Imaging | Radiology | 255 | 815-823 | 2010 | http://radiology.rsna.org/content/255/3/815.short | 14 | D | reproducibility of ADC in malignant hepatic tumors | 49 | 1.5 | Siemens |  | 2 | 50-500 | 3 | Matlab (4D Analyzer) | SSEPI | 6-channel torso array | BH and RT | Yes | 6/1.8 and 6/1.2 | 350x240-280 | 156x192 | 2 and 4 | 73 | 1500 and 3500 | Yes | Yes
Yoshikawa, T et al | ADC Measurement of Abdominal Organs and Lesions Using Parallel Imaging Technique | Am. J. Roent. | 187 | 1521-1530 | 2006 | http://www.ajronline.org/content/187/6/1521.short | 145 | D | reliability of parallel imaging for ADC in liver, spleen, pancreas, gallbladder, kidney, back muscle | 200 | 1.5 | Philips |  | 2 | 0-600 | 3 | Philips | SSEPI | PA | BH | Yes | 6/2.0 | 400x450 | 96x96 | 2 | 66 | 1500 | Yes | No
Parikh, T et al | Focal Liver Lesion Detection and Characterization with Diffusion-weighted MR Imaging: Comparison with Standard Breath-hold T2-weighted Imaging | Radiology | 246 | 812-822 | 2008 | http://radiology.rsna.org/content/246/3/812.long |  | D | Lesion detection and characterization | 53 | 1.5 | Siemens |  | 3 | 0-50-500 | 3 | Siemens Syngo | SSEPI | PA | BH and RT | Yes | 7/1.4 | 300x400 | 144x192 |  | 67-82 | 1300-1400 | Yes | No
Taouli, B et al | Evaluation of liver diffusion isotropy and characterization of focal hepatic lesions with two single-shot echo-planar MR imaging sequences: prospective study in 66 patients | Radiology | 226 | 71-78 | 2003 | http://radiology.rsna.org/content/226/1/71.long |  | D | Lesion characterization | 66 | 1.5 | Philips |  | 5 | 0-500 and 0-134-267-400 | 3 and 1 | Philips | SSEPI | PA | BH | Yes | 6/2.0 | 320 | 82x128 | 1 | 104 | 2400-3106 | No | No
excluded | review papers
Taouli, B et al | Diffusion-weighted MR Imaging of the Liver | Radiology | 254 | 47-66 | 2010 | http://radiology.rsna.org/content/254/1/47.short | 106

## Lung
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Imaging Sequence | Gating | Slice Thickness (mm) | FOV (mm) | Matrix | Lipid Suppresion | Parallel Imaging? | Other imaging notes | b-values (s/mm^2) | Test/Retest or Confidence Interval | Notes
Erasmus, JJ et al | Interobserver and Intraobserver Variability in Measurement of Non–Small-Cell Carcinoma Lung Lesions: Implications for Assessment of Tumor Response | J. Clin. Oncology | 21 | 2574-2582 | 2003 | http://jco.ascopubs.org/content/21/13/2574.short | 211
Regier, M et al | Diffusion weighted MRI and 18F-FDG PET/CT in non-small cell lung cancer (NSCLC): Does the apparent diffusion coefficient (ADC) correlate with tracer uptake (SUV)? | Eur J Radiol | 81 | 2913-2918 | 2012 | http://www.ejradiology.com/article/S0720-048X%2811%2900828-X/abstract |  | single-shot spin-echo (SE) echo-planar imaging (EPI) | respiration, no ECG | 5 | 425 | 256 | STIR impulse | T1w 3D-GRE |  | 0, 500 | interobserver agreement between 2 observers: mean difference 0.010, SD 0.087
Komori et al | 2-[Fluorine-18]-fluoro-2-deoxy-D-Glucose positron emission tomography / computed tomography versus whole-body diffusion-weighted MRI for detection of malignan lesions: initial experience | Ann Nucl Med | 21 | 209-215 | 2007 | http://www.ncbi.nlm.nih.gov/pubmed/17581719 | 83 | Disease | Sensibility of DWIBS as compared to 18F-FDG-PET. Differences between ADC of malignant lesions and reference organs (organs around lesions). | 16 | 1.5 | Philips | N/A | 1 (besides b0?) | 1000 | 3 | monoexponential model | SS-EPI | whole body | no (free breathing) | STIR | 6 (gap = 1) | 440 | 176 x 256 | 10 | 70 | 5009 | N/A | No
Mori et al. | Diffusion-weighted magnetic resonance imaging for diagnosing malignant pulmonary nodules/masses: comparison with positron emission tomography | J Thorac Oncol | 3 | 358-64 | 2008 | http://www.ncbi.nlm.nih.gov/pubmed/18379353 | 76 | Disease | Sensibility / specificity of ADC for discrimination between benign and malignant lesions of the lung. Comparison with FDG-PET | 140 | 1.5 | Philips | N/A | 1 (besides b0?) | 1000 | 3 | monoexponential model | SS-SE-EPI | N/A | no (shallow-breathing) | STIR | 6 (gap = 1) | 400 | 112x100 | 6 | 60 | 5900 | N/A | No | 1. The authors have used the minimum value of ADC inside a ROI because "the value whithin the entire tumor might not characterize the tumor because of its heterogeneity". Decision supported by Reference [10] of the article. This is important because some papers on DW-MRI report no significant differences in ADC values between benign and malignant lesions, which could be explained by the use of mean ADC ... 2. Quite interesting paper because of the high number of patients involved in the study (140!) and the availability of histophatological analysis
Ohno et al. | Diffusion-weighted MRI versus 18F-FDG PET/CT: performance as predictors of tumor treatment response and patient survival in patients with non-small cell lung cancer receiving chemoradiotherapy. | AJR Am J Roentgenol | 198 | 75-82 | 2012 | http://www.ncbi.nlm.nih.gov/pubmed/22194481 | 2 | Disease | Sensibility / specificity of ADC and SUVmax for response to treatment prediction | 64 | 1.5 | Philips | N/A | 2 | 0,1000 | 3 | monoexponential model | SS-SE-EPI | body coil (whole body coil?) | No (free breathing) | STIR | N/A | 300-350 x 300-350 | 96x96, 256x256 | 4 | 70 | 5000 | N/A | Inter- and intra-observer = 0.0 +/- 0.2 x 10^-3. This value is small when compared to the mean ADC value (2-3 x 10^-3)
Okuma et al. | Assessment of early treatment response after CT-guided radiofrequency ablation of unresectable lung tumours by diffusion-weighted MRI: a pilot study. | Br J Radiol | 82 | 989-94 | 2009 | http://www.ncbi.nlm.nih.gov/pubmed/19470575 | 22 | Disease | Differences in ADC before/after ablation between responders and non-responders | 17 | 1.5 | Siemens | N/A | 2 | 0,1000 | N/A | N/A | N/A | N/A | No (quiet breathing) | STIR | 6 | N/A | 128x128 | N/A | 82 | 5700 | N/A | No
[14 empty] | 3 | 0,500,1000 | [8 empty] | 128x128 |  | 81 | 5700
Sun et al. | Early evaluation of cancer response by a new functional biomarker: apparent diffusion coefficient. | AJR Am J Roentgenol | 197 | W23-9 | 2011 | http://www.ncbi.nlm.nih.gov/pubmed/21700991 | 4 | Disease | Differences in ADC between responders and non-responders | 102 (21 with lung cancer) | 1.5 | GE | N/A | 1 (besides b0?) | 0,500 | 3 | N/A | SS-EPI | eight-channel phased-array body coil | N/A | N/A | 8 | 420 | 384x256 | 4 | 90 | 2500 | N/A | No
Tondo et al. | Role of diffusion-weighted imaging in the differential diagnosis of benign and malignant lesions of the chest-mediastinum | Radiol Med | 116 | 720-33 | 2011 | http://www.ncbi.nlm.nih.gov/pubmed/21293944 | 8 | Disease | Sensibility / specificity of ADC for discrimination between benign and malignant lesions of the chest/mediastinum. | 34 | 1.5 | Philips | N/A | 3 | 0, 500, 1000 | 3 | monoexponential model | SS-EPI | phased-array sensitivity encoding (SENSE) body coils | N/A | STIR | 4 (gap = 4) | 520×303×329 | 132x74 | 2 | 66 | 3360 | N/A | No
Usuda et al. | Diffusion-weighted imaging is superior to positron emission tomography in the detection and nodal assessment of lung cancers | Ann Thorac Surg | 91 | 1689-95 | 2011 | http://www.ncbi.nlm.nih.gov/pubmed/21619964 | 8 | Disease | Sensibility / specificity of ADC for detection of lung lesions. Comparison with FDG-PET | 63 | 1.5 | Siemens | N/A | 2 | 0, 800 | 3 | N/A | SS-EPI | 2 anterior six-channel body phased-array coils and 2 posterior spinal clusters (six channels each) | N/A | Fat Sat | 6 | 350 | 128x128 | 5 | 65 | 4500 | N/A | No
Uto et al. | Higher sensitivity and specificity for diffusion-weighted imaging of malignant lung lesions without apparent diffusion coefficient quantification | Radiology | 252 | 247-54 | 2009 | http://www.ncbi.nlm.nih.gov/pubmed/19420317 | 55 | Disease | Sensibility / specificity of ADC and LSR for discrimination between benign and malignant lesions of the lung | 28 | 1.5 | GE | N/A | 2 | 0, 1000 | 3 | monoexponential model | SE-EPI | phased-array body coil (Eight-Channel Body Array Coil; GE Healthcare) | No (quiet breathing) | N/A | 5 | 400-480 | N/A | N/A | N/A | N/A | N/A | No

## Lymphoma
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | TI (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes | summary
Kwee, TC et al | ADC measurements of lymph nodes: Inter- and intra-observer reproducibility study and an overview of the literature | Eur. J. Radiol. | 75 | 215-220 | 2010 | http://www.sciencedirect.com/science/article/pii/S0720048X09001442 | 33 | normal | inter/intra observer variability | 20 | 1.5 | Phillips | 1 | 2 | 0, 1000 | 3 | ? Monoexponential | stejskal tanner single shot EPI | 4 channel phased array | ? | STIR and CHESS | 4 | 450x360 |  | 3 | 78 | 8612 | 180 | SENSE 2
Doskaliyev, A et al | Lymphomas and glioblastomas: differences in the apparent diffusion coefficient evaluated with high b-value diffusion-weighted magnetic resonance imaging at 3T | Eur. J. Radiol. | 81 | 339-344 | 2012 | http://www.sciencedirect.com/science/article/pii/S0720048X10005504 |  | 10 cerebral lymphoma/14 GBM | histopathological correlation | 24 | 3 | GE | 1 | 3 | 0,1000; 0,4000 | 3 | monoexponential | SSEPI | 8 channel | no | ? | 6 | 22 x 22 | 128 x 128 | 2 | 73 (b1000); 100 b(4000) | 5000 |  |  | none | All ADC values were statistically associated with tumor cellularity. ADCmin at b (4000) was associated with tumor cellulartiy more significantly than ADC min at b 1000. All ADC values were lower for lymphome than GBM and the statistical difference was larger at b 4000 than b 1000.
Kwee, TC et al | Whole-body magnetic resonance imaging, including diffusion-weighted imaging, for diagnosing bone marrow involvement in malignant lymphoma | Brit. J. Hameatology | 149 | 628-630 | 2010 | http://onlinelibrary.wiley.com/doi/10.1111/j.1365-2141.2010.08093.x/full | 10
Lin, C et al | Whole-body diffusion-weighted magnetic resonance imaging with apparent diffusion coefficient mapping for staging patients with diffuse large B-cell lymphoma | Eur. Radiol. | 20 | 2027-2038 | 2010 | http://www.springerlink.com/content/302766g58150756r/ | 29
Kwee, TC et al | Whole-Body MRI, Including Diffusion-Weighted Imaging, for the Initial Staging of Malignant Lymphoma | Invest. Radiol. | 44 | 683-690 | 2009 | http://journals.lww.com/investigativeradiology/Abstract/2009/10000/Whole_Body_MRI,_Including_Diffusion_Weighted.8.aspx | 39 | Staging malignant lymphoma vs. CT | Staging | 33 | 1.5 | Phillips |  | 2 | 0, 1000 | 3 | not calculated or used. Qualitative | SSEPI | 4 channel phased array | no | STIR | 4 | 450x360 | 128x81 | 3 | 78 | 8612 | 180 | SENSE 2 | not performed | CT vs. MRI (+/-) DWI; Seamless coronal whole-body T1W and STIR images were created by merging separately acquired stations, using software implemented in the standard operating console. Axial diffusionweighted images were first coronally reformatted with a slice thickness/ gap of both 3.5/0 and 7/0 mm, and then merged to create seamless coronal whole-body diffusion-weighted images. Wholebody diffusion-weighted images were displayed using greyscale inversion. | Good agreement between final staging and MRI + DWI. No ADC.
King, AD et al | Malignant Cervical Lymphadenopathy: Diagnostic Accuracy of Diffusion-weighted MR Imaging | Radiology | 245 | 806-813 | 207 | http://radiology.rsna.org/content/245/3/806.short | 89 | Prospective trial neck adenopathy origin (lymphoma, SCC, nasoph. CA) | histopathological correlation | 43 | 1.5 | Phillips |  | 6 | 0,100,200,300,400,500 | 3 | monoexponential | SSEPI | receive only | none | none | 4 | 23 x 23 | 112 x 112 | 4 | 75 | 2000 |  | none | none | Good ability to differentiate lymphoma, SCC and NPC by using threshold.
Sumi, M et al | Discrimination of Metastatic Cervical Lymph Nodes with Diffusion-Weighted MR Imaging in Patients with Head and Neck Cancer | AJNR | 24 | 1627-1634 | 2003 | http://www.ajnr.org/content/24/8/1627.short | 157 | squamous cell head and neck/lymphoma/reactive lymph nodes | 50 lymph nodes/histopathological correlation | 31 | 1.5 | GE | 0 | 2 | 5001000 | 3 | monoexponential | SSEPI | GE neurovascular array | none |  | 5mm | 24 x 24 | 128x128 | 4 | 64 (b500), 74 (b1000) | 10000 |  | none | none | The ADC was significantly greater in metastatic lymph nodes than in benign lymphadenopathy. Nodal lymphomas showed even lower levels of ADC.
Kwee, TC et al | Whole-body diffusion-weighted imaging for staging malignant lymphoma in children | Ped. Radiol. | 40 | 1592-1602 | 2010 | http://www.springerlink.com/content/586j50543w502820/ | 14

## Prostate
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Shimofusa, R et al | Diffusion-Weighted Imaging of Prostate Cancer | J. Comp. Assis. Tomo. | 29 | 149-153 | 2005 | http://journals.lww.com/jcat/Abstract/2005/03000/Diffusion_Weighted_Imaging_of_Prostate_Cancer.1.aspx | 135
Giles, SL et al | Apparent Diffusion Coefficient as a Predictive Biomarker of Prostate Cancer Progression: Value of Fast and Slow Diffusion Components | Am. J. Roent. | 196 | 586-591 | 2011 | http://www.ajronline.org/content/196/3/586.short | 12
Rosen, MA et al | American College of Radiology Imaging Network and Quantiative Imaging Biomarkers Alliance | American College of Radiology Imaging Network and Quantiative Imaging Biomarkers Alliance, ACRIN 6701 |  |  | 2012 | http://qibawiki.rsna.org/images/0/03/ACRIN-6701_master_031512.pdf
Yoshizako, T et al | Apparent diffusion coefficient of line scan diffusion image in normal prostate and prostate cancer--comparison with single-shot echo planner image | Mag. Reson. Imag. | 29 | 106-110 | 2011 | http://www.sciencedirect.com/science/article/pii/S0730725X10002171
Rosenkrantz, AB et al | Diffusion‐weighted imaging of the abdomen at 3.0 Tesla: Image quality and apparent diffusion coefficient reproducibility compared with 1.5 Tesla | J. Mag. Res. Imag | 33 | 128-135 | 2011 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22395/full | 14
deSouza, NM et al | Diffusion-weighted magnetic resonance imaging: a potential non-invasive marker of tumour aggressiveness in localized prostate cancer | Clin. Radiol. | 63 | 774-782 | 2008 | http://www.sciencedirect.com/science/article/pii/S0009926008000822 | 69
Afaq, A et al | Clinical utility of diffusion-weighted magnetic resonance imaging in prostate cancer | BJU Internat. | 108 | 1716-1722 | 2011 | http://onlinelibrary.wiley.com/doi/10.1111/j.1464-410X.2011.10256.x/full | 4
Lichy, MP et al | Tumor Detection by Diffusion-Weighted MRI and ADC-Mapping- Initial Clinical Experiences in Comparison to PET-CT | Invest. Radiol. | 42 | 605-613 | 2007 | http://journals.lww.com/investigativeradiology/Abstract/2007/09000/Tumor_Detection_by_Diffusion_Weighted_MRI_and.1.aspx | 78
Lim, HK et al | Prostate Cancer: Apparent Diffusion Coefficient Map with T2-weighted Images for Detection--A Multireader Study | Radiology | 250 | 145-151 | 2009 | http://radiology.rsna.org/content/250/1/145.short | 72
Katahira, K et al | Ultra-high-b-value diffusion-weighted MR imaging for the detection of prostate cancer: evaluation in 201 cases with histopathological correlation | Eur. Radiol. | 21 | 188-196 | 2011 | http://www.springerlink.com/content/t0167666575l12m5/ | 20
Vargas, HA, et al | Diffusion-weighted Endorectal MR Imaging at 3 T for Prostate Cancer: Tumor Detection and Assessment of Aggressiveness | Radiology | 259 | 775-784 | 2011 | http://radiology.rsna.org/content/259/3/775.short | 23
Tanimoto, A, et al | Prostate cancer screening: The clinical value of diffusion-weighted imaging and dynamic MR imaging in combination with T2-weighted imaging | J. Mag. Res. Imag | 25 | 146-152 | 2007 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.20793/full | 141
Kim, JH et al | Apparent diffusion coefficient: prostate cancer versus noncancerous tissue according to anatomical region | J. Mag. Res. Imag | 28 | 1173-1179 | 2008 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21513/full | 40
Tamada, T et al | Age-related and zonal anatomical changes of apparent diffusion coefficient values in normal human prostatic tissues | J. Mag. Res. Imag | 27 | 552-556 | 2008 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21117/full | 30
Tamada, T et al | Apparent diffusion coefficient values in peripheral and transition zones of the prostate: Comparison between normal and malignant prostatic tissues and correlation with histologic grade | J. Mag. Res. Imag | 28 | 720-726 | 2008 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21503/full | 75
Xu, J et al | Magnetic resonance diffusion characteristics of histologically defined prostate cancer in humans | Mag. Reson. Med. | 61 | 842-850 | 2009 | http://onlinelibrary.wiley.com/doi/10.1002/mrm.21896/full | 42
Mazaheri, Y et al | Prostate cancer: identification with combined diffusion-weighted MR imaging and 3D 1H MR spectroscopic imaging--correlation with pathologic findings | Radiology | 246 | 480-488 | 2009 | http://radiology.rsna.org/content/246/2/480.short | 114
Kitajima, K et al | High b-value diffusion-weighted imaging in normal and malignant peripheral zone tissue of the prostate: effect of signal-to-noise ratio | Mag. Reson. Med. Sci. | 7 | 93-99 | 2008 | http://www.ncbi.nlm.nih.gov/pubmed/18603841 | 42
Gibbs, P et al | Diffusion Imaging of the Prostate at 3.0 Tesla | Invest. Radiol. | 41 | 185-188 | 2006 | http://journals.lww.com/investigativeradiology/Abstract/2006/02000/Diffusion_Imaging_of_the_Prostate_at_3_0_Tesla.18.aspx | 104
Hosseinzadeh, K and Schwarz SD | Endorectal diffusion-weighted imaging in prostate cancer to differentiate malignant and benign peripheral zone tissue | J. Mag. Res. Imag | 20 | 654-661 | 2004 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.20159/full | 196
Kumar, V et al | Correlation between metabolite ratios and ADC values of prostate in men with increased PSA level | Mag. Reson. Imag | 24 | 541-548 | 2006 | http://www.sciencedirect.com/science/article/pii/S0730725X06000257 | 47
Kim, CK et al | Diffusion-Weighted MRI at 3 T for the Evaluation of Prostate Cancer | Am. J. Roent. | 194 | 1461-1469 | 2010 | http://www.ajronline.org/content/194/6/1461.short | 22
Zelhof, B et al | Correlation of diffusion-weighted magnetic resonance data with cellularity in prostate cancer | BJU Internat. | 103 | 883-888 | 2009 | http://onlinelibrary.wiley.com/doi/10.1111/j.1464-410X.2008.08130.x/full | 55
Messiou, C et al | Assessing response in bone metastases in prostate cancer with diffusion weighted MRI | Eur. Radiol. | 21 | 2169-2177 | 2011 | http://www.springerlink.com/content/t5v6247243v512v7/ | 3
Jacobs, MA et al | Diffusion-weighted imaging with apparent diffusion coefficient mapping and spectroscopy in prostate cancer | Top. Magn. Reson. Imag. | 19 | 261-272 | 2008 | http://www.ncbi.nlm.nih.gov/pubmed/19512848 | 18
Lawrence, EM et al | The emerging role of diffusion-weighted MRI in prostate cancer management | Nat. Rev. Urol. | 9 | 94-101 | 2012 | http://www.nature.com/nrurol/journal/v9/n2/full/nrurol.2011.222.html
Gibbs, P et al | Repeatability of echo-planar-based diffusion measurements of the human prostate at 3 T | Mag. Reson. Imag. | 25 | 1423-1429 | 2007 | http://www.sciencedirect.com/science/article/pii/S0730725X07002354 | 21 | normal | ADC, FA, Dxx, Dyy, Dzz | 8 | 3T | GE |  | 2 | 0,700 | 6 | NA | SS-SE-EPI | 8ch ph array | none | not specifieid (yes by fig inspection) | 2.7 | 350 | 128 x 128 | 6 | 65 | 6200 | 2x | repeatability ( = 2.77 * "common SD") | BW = +/- 250 kHz , coffee break and 1 month interval
Gibbs, P et al | Repeatability of echo-planar-based diffusion measurements of the human prostate at 3 T | Mag. Reson. Imag. | 25 | 1423-1429 | 2007 | http://www.sciencedirect.com/science/article/pii/S0730725X07002354 | 21 | normal | ADC-y | 8 | 3T | GE |  | 2 | 0,500 | 1 | NA | SS-SE-EPI | 8ch ph array | none | not specifieid (yes by fig inspection) | 5 | 260 | 224 x 224 | 16 | 66 | 4000 | 2x | repeatability ( = 2.77 * "common SD") | BW = +/- 250 kHz , coffee break and 1 month interval

## Sarcoma
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Baunin, C et al | Value of diffusion-weighted images in differentiating mid-course responders to chemotherapy for osteosarcoma compared to the histological response: preliminary results | Skeletal Radiology |  |  | 2012 | http://www.springerlink.com/content/v6868589k98148x0/
van Rijswijk, CSP et al | Diffusion-weighted MRI in the characterization of soft-tissue tumors | J. Mag. Reson. Imag. | 15 | 302-307 | 2002 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.10061/full | 76
Schnapauff, D et al | Diffusion-weighted echo-planar magnetic resonance imaging for the assessment of tumor cellularity in patients with soft-tissue sarcomas | J. Mag. Reson. Imag. | 29 | 1355-1359 | 2009 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.21755/full | 22
Einarsdóttir, H et al | Diffusion-weighted MRI of soft tissue tumours | Eur. Radiol. | 14 | 959-963 | 2004 | http://www.springerlink.com/content/a0aw0uy8kkkxhf8v/ | 71
Humphries, PD et al | Tumors in Pediatric Patients at Diffusion-weighted MR Imaging: Apparent Diffusion Coefficient and Tumor Cellularity | Radiology | 245 | 848-854 | 2007 | http://radiology.rsna.org/content/245/3/848.short | 74
Baur, A and Reiser, MF | Diffusion-weighted imaging of the musculoskeletal system in humans | Skeletal Radiology | 29 | 555-562 | 2000 | http://www.springerlink.com/content/lqnj5f8ayqde1mkk/ | 100
Uhl, M et al | Osteosarcoma: Preliminary Results of In Vivo Assessment of Tumor Necrosis After Chemotherapy With Diffusion- and Perfusion-Weighted Magnetic Resonance Imaging | Invest. Radiol. | 41 | 618-623 | 2006 | http://journals.lww.com/investigativeradiology/Abstract/2006/08000/Osteosarcoma__Preliminary_Results_of_In_Vivo.3.aspx | 62
Uhl, M et al | Evaluation of tumour necrosis during chemotherapy with diffusion-weighted MR imaging: preliminary results in osteosarcomas | Ped. Radiol. | 36 | 1306-1311 | 2006 | http://www.springerlink.com/content/d0r62n935g8051j4/ | 42
Dietrich, O et al | A comparative evaluation of a RARE-based single-shot pulse sequence for diffusion-weighted MRI of musculoskeletal soft-tissue tumors | Eur. Radiol. | 15 | 772-783 | 2005 | http://www.springerlink.com/content/ghmgenkvw37y20lq/ | 27

## Whole Body
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Koh, DM et al | Whole-Body Diffusion-Weighted MRI: Tips, Tricks, and Pitfalls | Am. J. Roent. | 199 | 252-262 | 2012 | http://www.ajronline.org/content/199/2/252.short
Padhani, AR et al | Whole-Body Diffusion-weighted MR Imaging in Cancer: Current Status and Research Directions | Radiology | 261 | 700-718 | 2011 | http://radiology.rsna.org/content/261/3/700.short | 1
Koh, DM and Collins DJ | Diffusion-Weighted MRI in the Body: Applications and Challenges in Oncology | Am. J. Roent. | 188 | 1622-1635 | 2007 | http://www.ajronline.org/content/188/6/1622.full | 499
Thoeny, HC and Keyzer, F | Extracranial applications of diffusion-weighted magnetic resonance imaging | Eur. Radiol. | 17 | 1385-1393 | 2007 | http://www.springerlink.com/index/F8641X3Q21073P43.pdf | 120
Lin, C et al | Whole-body diffusion-weighted magnetic resonance imaging with apparent diffusion coefficient mapping for staging patients with diffuse large B-cell lymphoma | Eur. Radiol. | 20 | 2027-2038 | 2010 | http://www.springerlink.com/content/302766g58150756r/ | 29 | lymphoma | correlation to FDG PET/CT | 15 | 1.5 | Siemens | 0 | 3 | 50400800 | 3 |  | SSEPI | phased array | respiratory | spectrally selective | 5mm | 330-400 | 130x108 |  | 73 | 3000 | GRAPPA 2 | none | A total of 296 lymph node regions in the 15 patients were analyzed. DWI findings matched PET/CT findings in 277 regions (94%) (kappa=0.85, p<0.0001), yielding sensitivity and specificity for DWI nodal involvement detection of 90% and 94%. Combiniing visual ADC analysis with size measurement increased DWI specificity to 100% with 81% sensitivity. For organ involvement, the two techniques agreed in all 20 recorded organs (100%). All involved organ lesions showed restricted diffusion.
Rosenkrantz, M et al | Diffusion‐weighted imaging of the abdomen at 3.0 Tesla: Image quality and apparent diffusion coefficient reproducibility compared with 1.5 Tesla | J. Mag. Reson. Imag. | 33 | 128-135 | 2011 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22395/full | 13 | normal | reproducibility | 8 | 1.5/3 | Siemens | 1 | multiples of 3 | (0-50-100, 0-150-200, 0-300-400, 0-500-600, and 0-700-800, (0, 50, 100, 150, 200, 300, 400, 500, 600, 700, and 800 s/mm2) s/mm2 | 3 | monoexponential | EPI | 6+6 spine/torso |  | SPAIR | 7mm | 262x350 | 144x192 interpolated to 388x384 | 1 | 76 | 1900 | GRAPPA 2 |  | Subjective image quality was significantly lower at 3.0T for all comparisons (P  0.0078– 0.0156). ADC values were similar at 1.5T and 3.0T for all assessed organs, except for lower liver ADC at 3.0T using b0-500- 600 and breathhold technique. ADC reproducibility was moderate at both 1.5T and 3.0T, with no significant dif- ference in coefficient of variation of ADC between field strengths
Kwee, TC et al | Whole-body diffusion-weighted imaging for staging malignant lymphoma in children | Ped. Radiol. | 40 | 1592-1602 | 2010 | http://www.springerlink.com/content/586j50543w502820/ | 14 | lymphoma in children | suggestions only |  | 1.5 |  | 1 | 2 | 0,1000 | 3 | monoexponential | SSEPI |  |  | STIR/Spectral | 4 | 450x360 | 128x81 | 3 | 78 | 6962 | 2 | none
Kwee, TC et al | Diffusion-weighted whole-body imaging with background body signal suppression (DWIBS): features and potential applications in oncology | Eur. Radiol. | 18 | 1937-1952 | 2008 | http://www.springerlink.com/content/87x516136g167053/ | 156
Nakanishi, K et al | Whole-body MRI for Detecting Metastatic Bone Tumor: Diagnostic Value of Diffusion-weighted Images | Mag. Reson. Med. Sci. | 6 | 147-155 | 2007 | https://www.jstage.jst.go.jp/article/mrms/6/3/6_3_147/_article | 91
Kwee, TC et al | Whole-body diffusion-weighted magnetic resonance imaging | Eur. J. Radiol. | 70 | 409-417 | 2009 | http://www.sciencedirect.com/science/article/pii/S0720048X09001818 | 46
Giannarini G, et al | Potential and Limitations of Diffusion-Weighted Magnetic Resonance Imaging in Kidney, Prostate, and Bladder Cancer Including Pelvic Lymph Node Staging: A Critical Analysis of the Literature | Eur. Urology | 61 | 326-340 | 2012 | http://www.sciencedirect.com/science/article/pii/S0302283811010268 | 7
Messiou, C et al | Assessing response in bone metastases in prostate cancer with diffusion weighted MRI | Eur. Radiol. | 21 | 2169-2177 | 2011 | http://www.springerlink.com/content/t5v6247243v512v7/ | 3 | prostate CA + HRT | change in ADC following Tx | 26 | 1.5 | Siemens | 1 | 5 (0,100,250,500,750) | 5 (0,100,250,500,750) | 3 | monoexponential | SSEPI | phased array | none | SPAIR | 5 | 350 |  | 4 | 80 | 3000 |  | none | mean adc/adc/slow in lesions from responders and progressors wshowed a significant increase. Although the marjority of lesions demonstrated an ADC/ADCslow rise, some lesions in both responsders and progressors demonstrated a fall in ADC beyond the limits of reproducibiilty. mean adc is not an appropriate measure of response in bone mets. The heterogeneity of changes in ADC is likely to be related to the composition of bone marrow with changes that have opposing effects on ADC.

## Whole Body (2)
Author | Title | Journal | Volume | Pages | Year | URL | Citations | Disease/Normal | Outcome Measured | # of patients | Field (T) | Vendor | # b0 images | # b-values | b-values (s/mm^2) | Diffusion directions | ADC Calc. Method | Imaging Sequence | Coil Type | Gating | Lipid Suppresion | Slice Thickness (mm) | FOV (mm) | Matrix | NEX | TE (ms) | TR (ms) | Parallel Imaging | Test/Retest, CI, CoV | Notes
Koh, DM et al | Whole-Body Diffusion-Weighted MRI: Tips, Tricks, and Pitfalls | Am. J. Roent. | 199 | 252-262 | 2012 | http://www.ajronline.org/content/199/2/252.short |  | Both | Review of basic principles, current developement, and challenges. |  | 1.5 | Siemens | NA (1?) | 2 | 0-100, 600-1000 | NA | NA | Spin echo | NA | NA | STIR | 5 | 380x380 | 150x256 | 4 | 72 | 14000 | yes (factor 2) | No | Review article. Parameters correspond to typical values for 3 manufacturers.
[12 empty] | Philips |  | 2 | 0, 1000 | [6 empty] | 5 | 400x280 | 128x96 | 4(b=0), 12 (b=1000) | 70 | 8322 | yes (factor 2)
[12 empty] | GE |  | 1 | 600 | [6 empty] | 8 | 440x440 | 128x88 | 3 | 64.6 | 6625 | NA
Padhani, AR et al | Whole-Body Diffusion-weighted MR Imaging in Cancer: Current Status and Research Directions | Radiology | 261 | 700-718 | 2011 | http://radiology.rsna.org/content/261/3/700.short | 1 | Both | Review of basic principles, current developement, and challenges. | NA | 1.5 (best fat supression) | Siemens | NA (1?) | 2 | 50, 900 | NA | Mono exponential model for 2 b-values (Stejskal-Tanner) | SS SE-EPI | Phased-array surface coils (four stations) | Free-breathing | STIR | 5 | 450x450 | 128x128 | 6 | 68 | 9000 | yes (factor 2) | Yes (In references, please see notes) | Review article. Some parameter values are recommended in an appendix (http://radiology.rsna.org/content/261/3/700/suppl/DC1) . References for for ADC variability are are given: [99] for reproducibility of ADC, [20][119] for alternative methods to measure variability , and [96] for multicenter studies. The paper sais that there are not established thresholds for extracranial applications yet.
Koh, DM and Collins DJ | Diffusion-Weighted MRI in the Body: Applications and Challenges in Oncology | Am. J. Roent. | 188 | 1622-1635 | 2007 | http://www.ajronline.org/content/188/6/1622.full | 499 | Both | Review of basic principles, current developement, and challenges. | NA | 1.5 | Philips | NA (1?) | 2 | 0, 1000 | 3 | Mono exponential model | SS SE-EPI | NA | Free-breathing | STIR | 4 (gap = 1mm) | 380x400 | 160x256 | 6 | 72 | 3500 | yes (factor 2) | No | Review article. Parameters correspond to recommended sequences for whole-body acquisitions. Two more parameter sets are provided for liver, abdomen and pelvis in Table I
Thoeny, HC and Keyzer, F | Extracranial applications of diffusion-weighted magnetic resonance imaging | Eur. Radiol. | 17 | 1385-1393 | 2007 | http://www.springerlink.com/index/F8641X3Q21073P43.pdf | 120 |  | Review article | [19 empty] | No (but one reference provided on inter subject variability could help [70]) | Review article. No specific parameter values nor test/retest are provided. Reference [70} provides values of intersubject variability.
Lin, C et al | Whole-body diffusion-weighted magnetic resonance imaging with apparent diffusion coefficient mapping for staging patients with diffuse large B-cell lymphoma | Eur. Radiol. | 20 | 2027-2038 | 2010 | http://www.springerlink.com/content/302766g58150756r/ | 29 | Disease | Sensibility specificity of lesion detection. Comparison with PET/CT. | 15 | 1.5 | Siemens | NA (1?) | 3 | 50, 400, 800 | 3 | NA | SS SE-EPI | 5 sets of phased-array coils covering head, neck, chest, abdomen,pelvis. Spine coil embeded in patient table. | Respiration-gated | FAT SAT | 5 | 330-400 (patient-dependent) | 130x108 | 4 | 73 | 3000 | yes (factor 2) | No | FAT SAT was preferred to STIR since this last one is twice longer and the SNR is lower. Lowest b was fixed to 50 to eliminate signals from vessels and to reduce perfusion effects. This value also provides images with anatomical information. B=800 to enhance constrast lesion/tissue without image distortion and signal lost. Nice paper to define protocol!
Rosenkrantz, M et al | Diffusion‐weighted imaging of the abdomen at 3.0 Tesla: Image quality and apparent diffusion coefficient reproducibility compared with 1.5 Tesla | J. Mag. Reson. Imag. | 33 | 128-135 | 2011 | http://onlinelibrary.wiley.com/doi/10.1002/jmri.22395/full | 13 |  | Review article | NA | [19 empty] | Impossible to download it from the web site (?), but this article is included for review in abdomen and prostate, so other reviewers should have read it.
Kwee, TC et al | Whole-body diffusion-weighted imaging for staging malignant lymphoma in children | Ped. Radiol. | 40 | 1592-1602 | 2010 | http://www.springerlink.com/content/586j50543w502820/ | 14 | Both | Review article | NA | 1.5 | NA | NA (1?) | 2 | 0, 1000 | 3 | NA | SS SE-EPI | Specific coild for each section | Free-breathing | FAT SAT | 4 (no gap) | 450 x 360 | 128 x 81 | 3 | 78 | 6982 | yes (factor 2) | No | Review article. Parameters correspond to two suggested sequences for FAT SAT and STIR. The only different parameter is TR
8612
Kwee, TC et al | Diffusion-weighted whole-body imaging with background body signal suppression (DWIBS): features and potential applications in oncology | Radiology | 18 | 1937-1952 | 2008 | http://www.springerlink.com/content/87x516136g167053/ | 156 | Both | Review of basic principles, current developement, and challenges. | NA | 1.5 | NA | NA (1?) | 1 (besides b=0?) | 1000 | 1 | Monoexponential model | DWIBS | built-in body coil / surface body coils | Free-breathing | STIR | 4 | 400 (rectangular FOV of 70%) | 160 | 10 | 70 | 10205 | yes | No | b=500 for abdomen and b=1500/2000 for prostate
[8 empty] | Both |  | NA | 1.5 | NA | NA (1?) | 1 (besides b=0?) | 1000 | 1 | Monoexponential model | DWIBS | built-in body coil / surface body coils | Free-breathing | CHESS | 4 | 400 (rectangular FOV of 70%) | 160 | 6 | 70 | 7800 | yes | No
Nakanishi, K et al | Whole-body MRI for Detecting Metastatic Bone Tumor: Diagnostic Value of Diffusion-weighted Images | Mag. Reson. Med. Sci. | 6 | 147-155 | 2007 | https://www.jstage.jst.go.jp/article/mrms/6/3/6_3_147/_article | 91 | Disease | Sensibility specificity of metastases in bone tumor. | 30 | 1.5 | Philips | NA (1?) | 2 | 0, 600 |  |  | SS STIR-EPI | quadrature body coil (whole body) | Free breathing | STIR | 5 | 450 | 112 x 112 |  | 59 | 6243 |  | No | ADC was no computed. Only visual image inspection of DWI combined with other weightings to study changes in sensibility/specificity in the lesion detection
Kwee, TC et al | Whole-body diffusion-weighted magnetic resonance imaging | Eur. J. Radiol. | 70 | 409-417 | 2009 | http://www.sciencedirect.com/science/article/pii/S0720048X09001818 | 46 | Both | Review of basic principles and current developement. | NA | 1.5 | NA | NA (1?) | 0 | 0, 1000 | 3 | Monoexponential model | DWIBS | proposed the use of a single surface coil manually displaced between acquisitions. The individual images are then merged by using the scanner's software | Free breathing | STIR or CHESS | 4 (no gap) | 430 (rectangular FOV of 80%) | 160 | 10 | 70 | 10205 | yes (factor 2) | No | Review article. Parameters correspond to recommended sequences for whole-body acquisitions.
Giannarini G, et al | Potential and Limitations of Diffusion-Weighted Magnetic Resonance Imaging in Kidney, Prostate, and Bladder Cancer Including Pelvic Lymph Node Staging: A Critical Analysis of the Literature | Eur. Urology | 61 | 326-340 | 2012 | http://www.sciencedirect.com/science/article/pii/S0302283811010268 | 7
Messiou, C et al | Assessing response in bone metastases in prostate cancer with diffusion weighted MRI | Eur. Radiol. | 21 | 2169-2177 | 2011 | http://www.springerlink.com/content/t5v6247243v512v7/ | 3 | Disease | Correlation between ADC and response to treatment. | 26 | 1.5 | Siemens | NA (1?) | 6 | 0, 50, 100, 250, 500, 750 | 3 | Monoexponential model | Twice refocused spin echo planar | NA | NA | FAT SAT | 5 | 350 | NA | 3 | 80 | 3000 |  | No