QIBA Profile.  Computed Tomography: Lung Densitometry

Version 0.3

01-July-2011

Table of Contents

Open Issues:	2

Closed Issues:	3

I. Executive Summary	4

II. Clinical Context and Claims	4

Utilities and Endpoints for Clinical Trials	4

Claim 1: Longitudinal Stability of Lung-equivalent phantom Densitometry	5

Claim 2: Longitudinal Stability of Lung Densitometry	5

III. Profile Details	6

1. Subject Handling	7

2. Image Data Acquisition	8

3. Image Data Reconstruction	10

4. Image Analysis	14

IV. Compliance	14

Acquisition Device	14

Reconstruction Software	15

Software Analysis Tool	15

Image Acquisition Site	15

References	15

Appendices	16

Acknowledgements and Attributions	16

Background Information	16

Conventions and Definitions	20

Breathing Instructions	21

Model-specific Instructions and Parameters	24

# Open Issues:

The following issues have not been resolved to the satisfaction of the technical committee.  An open issue may be a short question prompting a proposed resolution or discussion. The issues and answers below may represent some of the directions the Committee is currently leaning. Feedback on these issues is encouraged, particularly during the Public Comment period for the profile.

Q. Is the claim appropriate/supported by the profile details and groundwork? A.

Q. What kind of additional study (if any) would best prove the profile claim? A.

Q. How do we balance specifying what to accomplish vs how to accomplish it? A. E.g. if the requirement is that the scan be performed the same way, do we need to specify that the system or the Technologist record how each scan is performed? If we don’t, how will the requirement to “do it the same” be met?

Q. Should there be a “patient appropriateness” or “subject selection” section? A. The protocol template includes such a section to describe characteristics of appropriate (and/or inappropriate) subjects. E.g. a requirement that the patient be able to hold their breath for 15 seconds. We could also discuss what constitutes an “assessable lesion” (the claim introduces this term)

Q. Are there sufficient DICOM fields for all of what we need to record in the image header, and what are they specifically? A. For those that exist, we need to name them explicitly. For those that may not currently exist, we need to work with the appropriate committees to have them added.

Q. Have we worked out the details for how we establish compliance to these specifications? A. We are continuing to work on how this is to be accomplished but felt that it was helpful to start the review process for the specifications in parallel with working on the compliance process.

Q. Should we specify all three levels (Acceptable, Target, Ideal) for each parameter? A. The vCT group has said No. As of this version for densitometry, we have left them in since group discussion still is considering this.

Q. What is the basis for our claim, and is it only aspirational? A. This is still a topic of discussion for the group.

Q. What about dose? A. This is still a topic of discussion for the group.

# Closed Issues:

The following issues have been considered closed by the technical committee.  They are provided here to forestall discussion of issues that have already been raised and resolved, and to provide a record of the rationale behind the resolution.

# I. Executive Summary

This document establishes the methods to estimate lung densities and lung volumes from CT images obtained during longitudinal studies.  Lung densities and lung volumes will serve as surrogate endpoints for severity of emphysema, degree of air trapping, and degree of hyperinflation in COPD and asthma.

QIBA has constructed a systematic approach for standardizing and qualifying densitometry as a biomarker. This QIBA Profile is expected to provide specifications that may be adopted by users as well as equipment developers to meet targeted levels of clinical performance in identified settings.

The intended audiences include:

- Technical staffs of software developers and device manufacturers who create products for this purpose

- Clinical trial scientists

- Practicing clinicians at healthcare institutions considering appropriate specifications for procuring new equipment

- Experts involved in quantitative medical image analysis

- Anyone interested in the technical and clinical aspects of medical imaging

Note that specifications stated as “requirements” here are only requirements to achieve the claim, not “requirements on standard of care.”   Specifically, meeting the goals of the profile are secondary to properly caring for the patient.

# II. Clinical Context and Claims

## Utilities and Endpoints for Clinical Trials

This imaging protocol will be able to establish three endpoints to serve a clinical trial:

###### 1. It will establish the severity of emphysema.

Define specific metrics below. Emphysema metrics are global measurements based on CT attenuation histogram. Methods used for calculation of extent of emphysema include the density mask technique, where the percent emphysema is calculated based on the percentage of lung voxels with CT attenuation below a given threshold. The threshold used for extent of emphysema has varied from -900 to -970 HU. Correlation with histologic and morphometric measures of emphysema suggests that the optimal threshold at total lung capacity, using thin-section densitometry CT, may be -960 or -970. Percentile values based on the CT attenuation histogram may be ….. Perc 15. (caveats- sliding scale, sensitivity vs specificity, scanner variability, slice thickness).

2. It will establish the degree of air trapping in the lungs.

Define specific metrics below. -856 HU at FRC or RV- pros and cons- limited data

3. It will establish inspiratory and expiratory lung volumes.

Define specific metrics below. TLC. Correlation with physiologic TLC. FRC/RV

## Claim 1: Longitudinal Stability of Lung-equivalent phantom Densitometry

Phil Judy: I recommend expressing the compliance levels for lung density metrics that are consistent with the CT number precisions specification (Section 7).  Have a description and use of various emphysema metrics and relationships with each other.  Then estimate expected precision of those emphysema measurements.  Instrumental claim for air trapping same as emphysema because based on same phantom measurements.

We describe procedures to establish reference values (Ref and Ref’)  in Section 8. Image Post-processing Metrics or in Section 9 Image Analysis. Group should decide which section

Using this profile, we will be able to achieve adequate calibration and longitudinal stability of CT scanners in order to permit measurement of longitudinal change in attenuation of air on CT phantom of ±1 HU. CT phantom measurements should be made at least once a month, and after every recalibration, hardware/software change, or CT service, to ensure long-term stability and precision of CT measurements. CT phantom should be optimized for measurement of air and lung equivalent material (COPDGene II phantom or equivalent phantom as approved by QIBA technical committee). Phantom should be scanned using the same CT parameters as used in the CT study, and at each available scan acquisition data collection diameter. While there are some data on short-term precision, there is little information on long-term precision of CT measurements of emphysema.

## Claim 2: Longitudinal Stability of Lung Densitometry

Claim regarding precision of in vivo patient measurement. Using this protocol, we will be able to measure severity of lung emphysema (measured using density mask technique at -950 HU) with repeatability of X%, or 5 HU shift in 15th percentile of lung attenuation (do we have data for this?) with repeatability of X%.

Compliance Levels for Measurement CT phantom measurements of air or lung equivalent material)

| Measurement Result | Performance Levels Achieved under Bull's Eye Conditions |
| --- | --- |
| CT attenuation (air inside phantom) | If Activities are Performed at Target Level -995 to -1005 HU |
| Longitudinal stability: CT attenuation (lung equivalent material) Deviation from baseline measurement | If Activities are Performed at Target Level ± 2 HU |

# III. Profile Details

A technical description of tests for the biomarker, identifying measurement activities and read-outs, is provided:

Figure 1: Replace with similar diagram for densitometry

Formally defined “Actors” who are required to meet these claims include the following:

- Hardware and software devices (acquisition, reconstruction, and analysis)

- Technologists

- Image Analysts

- Image Acquisition Sites

The following sections provide details for what the various components required for compliance:

Section 1, Subject Handling, is practiced by an Image Acquisition Site.

Section 2, Imaging Data Acquisition, is practiced by a Technologist at an Image Acquisition Site using an Acquisition Device.

Section 3, Imaging Data Reconstruction, is practiced by an Technologist at an Image Acquisition Site using Reconstruction Software.

Section 4, Image Analysis, is practiced by an Image Analyst using one or more Software Analysis Tools.

The requirements included herein are intended to establish a baseline level of capabilities. Providing higher performance or advanced capabilities is both allowed and encouraged.  The profile is not intended to be limiting in any way with respect to how these requirements are met by equipment suppliers.

## 1. Subject Handling

1.1 Timing Relative to Index Intervention Activity

The pre-treatment CT scan shall take place prior to any intervention to treat the disease. This scan is referred to as the “baseline” scan.  It should be acquired as soon as possible before the initiation of treatment, and in no case more than the number of days before treatment specified in the protocol.

1.2 Timing Relative to Confounding Activities

This document does not presume any timing relative to other activities.

1.3 Contrast Preparation and Administration

Contrast agents are never used in CT assessment of COPD or Asthma.

1.4 Subject Positioning

Discussion

Consistent positioning avoids unnecessary variance in attenuation, changes in gravity induced shape and fluid distribution, or changes in anatomical shape due to posture, contortion, etc. Significant details of subject positioning include the position of their upper extremities, the anterior-to-posterior curvature of their spines as determined by pillows under their backs or knees, the lateral straightness of their spines, and, if prone, the direction the head is turned.  Positioning the subject Supine/Arms Up/Feet first has the advantage of promoting consistency, and reducing cases where intravenous lines go through the gantry, which could introduce artifacts.

Specification

| Parameter | Specification |
| --- | --- |
| Subject Positioning | The Technologist shall position the subject the same as for prior scans. If the previous positioning is unknown, the Technologist shall position the subject Supine/Arms Up/Feet first if possible. |
| Table Height | The Technologist shall adjust the table height to place the mid-axillary line at isocenter. |
| Image Header | The Acquisition Device shall record the Table Height in the image header. |

1.5 Instructions to Subject During Acquisition

Discussion

Breath holding reduces motion that might degrade the image. Full inspiration inflates the lungs, which separates structures and makes lesions more conspicuous.

Although performing the acquisition in several segments (each of which has an appropriate breath hold state) is possible, performing the acquisition in a single breath hold is likely to be more easily repeatable and does not depend on the Technologist knowing where the lesions are located.

Specification

| Parameter | Specification |
| --- | --- |
| Breath hold | The Technologist shall ensure that image acquisition occurs at least near the high end inspiration. The Technologist shall ensure that for each lesion the breath hold state is the same as for prior scans. |
| Image Header | The Technologist shall record factors that adversely influence patient positioning or limit their ability to cooperate (e.g., breath hold, remaining motionless, agitation in patients with decreased levels of consciousness, patients with chronic pain syndromes, etc.). These shall be accommodated with data entry fields provided by the Acquisition Device. |

## 2. Image Data Acquisition

Discussion

CT scans for tumor densitometry analysis will be performed on equipment that complies with the specifications set out in this profile.  At this stage of development, we continue to recommend that all CT scans for an individual participant be performed on the same platform throughout the trial. In the rare instance of equipment malfunction, follow-up scans on an individual participant can be performed on the same type of platform. All efforts should be made to have the follow-up scans performed with identical parameters as the first.  This is inclusive of as many of the scanning parameters as possible, including the same field of view (FOV).

A set of scout images should be initially obtained. Pitch is chosen so as to allow completion of the scan in a single breath hold. In some cases two or more breaths may be necessary. In those cases, it is important that the target lesion be fully included within one of the sequences.

Faster scans shorten the scan time and reduce the breath hold requirements, thus reducing the likelihood of motion artifacts.  Scan Plane (transaxial is preferred) may differ for some subjects due to the need to position for physical deformities or external hardware.

Total Collimation Width (defined as the total nominal beam width) is often not directly visible in the scanner interface. Wider collimation widths can increase coverage and shorten acquisition, but can introduce cone beam artifacts which may degrade image quality.

Slice Width directly affects voxel size along the subject z-axis. Smaller voxels are preferable to reduce partial volume effects and provide higher accuracy due to higher spatial resolution.

X-ray CT uses ionizing radiation.  Exposure to radiation can pose risks.  It is recognized that there are tradeoffs between radiation dose and image quality.  As the radiation dose is reduced, image quality can be degraded.  It is expected that health care professionals will balance the need for good image quality with the risks of radiation exposure on a case-by-case basis.  It is not within the scope of this document to describe how these trade-offs should be resolved.

Specification

| Parameter | Compliance Levels whole lung |
| --- | --- |
| Whole lung scan duration | Target 10 sec |
| Parameter | Compliance Levels whole lung |
| Anatomic Coverage | Target Entire lung fields, bilaterally (lung apices through bases) |
| Parameter | Compliance Levels whole lung |
| Data co (if available) | Target Rib to rib |
| Parameter | Compliance Levels |
| IEC Pitch | Target 0.9-1.1 |

CT scans are acquired using a tube potential of 120 kV with no tube current modulation.

Radiation dose values specific to protocol, shall be tailored to patient size, and shall be kept constant for longitudinal studies.

Rotation speed should be ≤ 500 msec

Pitch specification is specific to protocol.  This is a sampling artifact – dose tradeoff.  Phil Judy would claim if artifact “specification” met then pitch of 2.0 is ideal.

Acquisition collimation (cone angle)?

Detector configuration?

Dose modulation- currently not recommended because of varying manufacturer specifications.

The following is from the corresponding section of vCT:

| Parameter | Specification |
| --- | --- |
| Scan Duration for Thorax | The Acquisition Device shall be capable of performing the required scans at an axial rate of at least 4cm per second. |
| Anatomic Coverage | The Technologist shall perform the scan such that the acquired anatomy is the same as for prior scans. |
| Scan Plane (Image Orientation) | The Technologist shall set the scan plane to be the same as for prior scans. |
| Total Collimation Width | The Acquisition Device shall be set up so as to achieve a total collimation width >=20mm. |
| IEC Pitch | The Acquisition Device shall be set up so as to achieve IEC pitch less than 1.5. |
| Tube Potential | The Acquisition Device shall be set up so as to achieve same kVp for all scans |
| Single Collimation Width | The Acquisition Device shall be set up so as to achieve single collimation width <= 1.5mm. |
| Image Header | The Acquisition Device shall record actual Anatomic Coverage, Field of View, Scan Duration, Scan Plane, Total Collimation Width, Single Collimation Width, Scan Pitch, Tube Potential, and Slice Width in the image header. |

## 3. Image Data Reconstruction

Discussion

Spatial Resolution quantifies the ability to resolve spatial details. Lower spatial resolution can make it difficult to accurately determine the borders of tumors, and as a consequence, decreases the precision of volume measurements.  Increased spatial resolution typically comes with an increase in noise. Therefore, the choice of factors that affect spatial resolution typically represent a balance between the need to accurately represent fine spatial details of objects (such as the boundaries of tumors) and the noise within the image. Spatial resolution is mostly determined by the scanner geometry (which is not usually under user control) and the reconstruction kernel (which is somewhat under user control as the user usually gets to choose from a limited set of choices of reconstruction kernels provided at the scanner).  It is stated in terms of “the number of line-pairs per cm that can be resolved in a scan of resolution phantom (such as the synthetic model provided by the American College of Radiology and other professional organizations).” –OR– “the full width at half of the line spread function”.

Noise Metrics quantify the magnitude of the random variation in reconstructed CT numbers.  Some properties of the noise can be characterized by the standard deviation of reconstructed CT numbers over a uniform region in phantom.  Noise (pixel standard deviation) can be reduced by using thicker slices for a given mAs. A constant value for the noise metric might be achieved by increasing mAs for thinner slices and reducing mAs for thicker slices.  The standard deviation is limited since it can vary by changing the reconstruction kernel, which will also impact the spatial resolution.  A more comprehensive metric would be the noise-power spectrum which measures the noise correlation at different spatial frequencies.

Reconstruction Field of View affects reconstructed pixel size because the fixed image matrix size of most CT scanners is 512 X 512.  If it is necessary to expand the field of view to encompass more anatomy, the resulting larger pixels may be insufficient to achieve the claim. A targeted reconstruction with a smaller field of view may be necessary, but a reconstruction with that field of view would need to be performed for every time point. Pixel Size directly affects voxel size along the subject x-axis and y-axis. Smaller voxels are preferable to reduce partial volume effects and provide higher measurement precision.  Pixel size in each dimension is not the same as resolution in each dimension; inherent resolution is different than how the data is reconstructed and is strongly affected by the reconstruction kernel.   When comparing data fields of different resolution, do not sacrifice higher resolution data to match the level of lower resolution data.

Reconstruction Interval (a.k.a. Slice spacing) that results in discontiguous data is unacceptable as they may “truncate” the spatial extent of the tumor, degrade the identification of tumor boundaries, confound the precision of measurement for total tumor volumes, etc.  Decisions about overlap (having an interval that is less than the nominal reconstructed slice thickness) need to consider the technical requirements of the clinical trial, including effects on measurement, throughput, image analysis time, and storage requirements. Reconstructing datasets with overlap will increase the number of images and may slow down throughput, increase reading time and increase storage requirements.  For multidetector row CT (MDCT) scanners, creating overlapping image data sets has NO effect on radiation exposure; this is true because multiple reconstructions having different kernel, slice thickness and intervals can be reconstructed from the same acquisition (raw projection data) and therefore no additional radiation exposure is needed.  <Note that the slice thickness is “nominal” since the thickness is not technically the same at the middle and the edges>

Reconstruction Kernel Characteristics need to be defined to optimize the analysis for each lesion while still meeting the requirements for noise and spatial resolution. A softer kernel can reduce noise at the expense of spatial resolution. An enhancing kernel can improve resolving power at the expense of increased noise.

The effects of iterative reconstructions on quantitative accuracy and reproducibility are currently not fully understood as of this writing of this profile version.

Specification

For quantification of whole tumor volumes, the reconstruction software produces images that meet the following specifications:

| Reconstruction field of view | Acceptable Complete thorax Target Rib to rib Ideal Rib to rib |
| --- | --- |
| Parameter | Compliance Levels whole lung |
| Matrix | Acceptable 512X512 Target 512X512 Ideal 1024X1024 |
| Slice thickness | Acceptable <= 2.0 mm Target <= 1.0 mm Ideal <= 0.5 mm |
| Slice Interval | Acceptable Shall be less 1.1 (contiguous) Target Should be less 0.75 (25% overlap) Ideal Should be less 0.5 (50% overlap) |
| Parameter | Compliance Levels |
| Reconstruction Kernel Characteristics: Emphysema | Target Relatively soft |

Insert COPDGene phantom measures

The full width at half of the line spread function shall be used as a measurement of resolution.  (Could use a cutoff of the MTF.  These compliance levels are driven by air way size measurement, density measurement may benefit from thicker sections.  Need to define a measurable slice thickness.

From previous document. “Spatial Resolution Metric quantifies the ability to resolve spatial details. It is stated in terms of the number of line-pairs per cm that can be resolved in a scan of resolution phantom (such as the synthetic model provided by the American College of Radiology and other professional organizations). Does not determine an objective quantitative measurement.  ACR subjective visual evaluation.  Lower spatial resolution can make it difficult to accurately determine the borders of tumors, and as a consequence, decreases the precision of volume measurements.  Spatial resolution is mostly determined by the scanner geometry (which is not usually under user control) and the reconstruction algorithm (which is under user control)”

| Axial FWHM | Acceptable <= 2.0 mm Target <= 1.0 mm Ideal <= 0.5 mm |
| --- | --- |
| Measured slice thickness | Acceptable <= 2.0 mm Target <= 1.0 mm Ideal <= 0.5 mm |
| Measured slice Interval | Acceptable Shall be less 1.1 (contiguous) Target Should be less 0.75 (25% overlap) Ideal Should be less 0.5 (50% overlap) |
| Alternative Spatial Resolution | Acceptable >= 6 lp/cm Target >= 7 lp/cm Ideal >= 8 lp/cm |

Noise Metrics quantify the magnitude of the random variation in reconstructed CT numbers.  Some properties of the noise can be characterized by the standard deviation of reconstructed CT numbers over a uniform region in phantom.  The standard phantom is a 20cm acrylic tube filled with water.   Noise (pixel standard deviation) can be reduced by using thicker slices for a given mAs. A constant value for the noise metric might be achieved by increasing mAs for thinner slices and reducing mAs for thicker slices.  The limitation of the standard deviation as complete measure of noise is the pixel standard deviation is modified by changes in spatial characteristics of reconstruction kernel in the absence of changes in radiation dose or slice thickness.  The reconstruction kernel and procedure for obtaining the noise metric for a given protocol should be described in that protocol.

| Delete this table Noise Metrics | Acceptable Std. dev. in 20cm water phantom < 40 HU Target Ideal |
| --- | --- |

“Pixel Size directly affects voxel size along the subject x-axis and y-axis. Smaller voxels are preferable to reduce partial volume effects and (likely) provide higher measurement precision.” Pixel size is determined by field of view, so don’t need to specify Compliance Levels. In any case should specify in Section 7.3 Image Reconstruction.

Scan Plane may differ for some subjects due to the need to position for physical deformities, but should be constant for each scan of a given subject. Slice plane specification is protocol specific.

How are artifact quantified? “Motion Artifacts may produce false targets and distort the size of existing targets. “Minimal” artifacts are such that motion does not degrade the ability of image analysts to detect the boundaries of target lesions.” This is statement is specific to volume measurements.  Artifact specification specific to protocol.

The following is from the corresponding section of vCT:

| Parameter | Specification |
| --- | --- |
| Spatial Resolution | The Reconstruction Software shall be set up so as to achieve spatial resolution >= 6 lp/cm –OR– Axial FWHM <= 0.8mm. |
| Voxel Noise | The Reconstruction Software shall be set up so as to achieve voxel noise standard deviation of < 5HU in 20cm water phantom. |
| Reconstruction Field of View | The Reconstruction Software shall be set up so as to achieve a reconstruction field of view spanning the entire lateral extent of the patient, but no greater than required to image the entire body; <same as previous scan> |
| Slice Thickness | The Reconstruction Software shall be set up so as to achieve slice thickness ≤2.5 mm. |
| Reconstruction Interval | The Reconstruction Software shall be set up so as to achieve reconstruction interval ≤2.5 mm. |
| Reconstruction Overlap | The Reconstruction Software shall be set up so as to achieve reconstruction overlap > 0 (i.e. no gap, and may have some overlap). |
| Reconstruction Kernel Characteristics | The Reconstruction Software shall be set up so as to utilize an equivalent kernel for all time points. |
| Image Header | The Reconstruction Software shall record actual Spatial Resolution, Noise, Pixel Spacing, Reconstruction Interval, Reconstruction Overlap, Reconstruction Kernel Characteristics, as well as the model-specific Reconstruction Software parameters utilized to achieve compliance with these metrics in the image header. |

## 4. Image Analysis

Discussion

Segmentation method- lung or lobe (protocol specific). Image processing shall use the same segmentation algorithm throughout each study.

(Which algorithm should be used? How do we ensure that image analysis algorithms are measuring the same thing?)

Histogram measurement

Fraction of lung less than “reference CT number”- usually -950 (but may vary depending on phantom measurements)

Percentile(“reference”)

Volume correction

Longitudinal analysis:

Shall be on same scanner, same acquisition parameters, similar breathhold.  If different scanner, adjust for inter-scanner differences, adjust for differences in breathhold.

9.3. Required Characteristics of Resulting Data

**Introduce the section.

Lung Volume

Emphysema metrics

Specification

| Parameter | Compliance Levels |
| --- | --- |
| Intra-rater reliability of fully automated read-outs | Acceptable Greater than 80% Target Greater than 90% Ideal Greater than 95% |

# IV. Compliance

## Acquisition Device

Compliance is certified according to specifications set out in the Image Acquisition section above.  Additionally, compliant Acquisition Devices shall provide means to record the information identified in the Subject Handling section as means to document compliance of the Image Acquisition Site to the specifications noted there.

## Reconstruction Software

Compliance to specifications as set out in the Image Reconstruction section above.  Additionally, compliant Reconstruction Software shall propagate the information collected at the prior Subject Handling and Imaging Acquisition stages and extend it with those items noted in the Reconstruction section.  See the compliance procedure notes associated with Acquisition Devices above for procedural assistance to identify Model Specific Parameters for Reconstruction Software.

## Software Analysis Tool

Compliance to specifications as set out in the Image Analysis section above.  Additionally, compliant Software Analysis Tools shall propagate the information collected at the prior Subject Handling, Imaging Acquisition, and Imaging Reconstruction stages and extend it with those items noted in the Analysis section

## Image Acquisition Site

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large patient population under consideration.  For imaging it is important to consider the availability of:

- appropriate imaging equipment and quality control processes,

- appropriate injector equipment and contrast media,

- experienced CT Technologists for the imaging procedure, and

- processes that assure imaging profile compliant image generation at the correct point in time.

A calibration and QA program shall be designed consistent with the goals of the clinical trial. This program shall include (a) elements to verify that sites are performing correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address issues relating to both radiation dose and image quality (which may include issues relating to water calibration, uniformity, noise, spatial resolution -in the axial plane-, reconstructed slice thickness z-axis resolution, contrast scale, CT number calibration and others). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.

# References

<fill in, mostly from those cited in the Background appendix>

# Appendices

## Acknowledgements and Attributions

This document is proffered by the Radiological Society of North America (RSNA) Quantitative Imaging Biomarker Alliance (QIBA) Densitometry Computed Tomography (v-CT) Technical Committee. The v-CT technical committee is composed of scientists representing the imaging device manufacturers, image analysis software developers, image analysis laboratories, biopharmaceutical industry, academia, government research organizations, professional societies, and regulatory agencies, among others. All work is classified as pre-competitive. A more detailed description of the v-CT group and its work can be found at the following web link: http://qibawiki.rsna.org/index.php?title=Densitometry_CT.

The Densitometry CT Technical Committee (in alphabetical order):

- …

The Densitometry CT Technical Committee is deeply grateful for the support and technical assistance provided by the staff of the Radiological Society of North America.

## Background Information

QIBA

The Quantitative Imaging Biomarker Alliance (QIBA) is an initiative to promote the use of standards to reduce variability and improve performance of quantitative imaging in medicine. QIBA provides a forum for volunteer committees of care providers, medical physicists, imaging innovators in the device and software industry, pharmaceutical companies, and other stakeholders in several clinical and operational domains to reach consensus on standards-based solutions to critical quantification issues. QIBA publishes the specifications they produce (called QIBA profiles), first to gather public comment and then for field test by vendors and users.

QIBA envisions providing a process for developers to test their implementations of QIBA profiles through a compliance mechanism. After a committee determines that a profile has undergone sufficient successful testing and deployment in real-world care settings, it is released for use.  Purchasers can specify conformance with appropriate QIBA profiles as a requirement in requests for proposal. Vendors who have successfully implemented QIBA profiles in their products can publish conformance statements (called QIBA Compliance Statements) represented as an appendix called “Model-specific Parameters.”  General information about QIBA, including its governance structure, sponsorship, member organizations and work process, is available at http://qibawiki.rsna.org/index.php?title=Main_Page.

CT Densitometry for Lung Disease

Inflammatory parenchymal lung diseases are common and are significant causes of disability and premature death. These diseases are the result of sub-acute/chronic or chronic inflammatory processes, and are linked to cigarette smoking, either as a cause or as a modifying agent. Chronic obstructive pulmonary disease (COPD) is currently the 12th leading cause of disability in the world and is predicted to be 5th by the year 2020 (201). In the United States alone, it has been estimated that the annual cost of morbidity and early mortality due to COPD is approximately 4.7 billion dollars (202). COPD is a complex condition in which environmental factors interact with genetic susceptibility to cause disease. Tobacco smoke is the most important environmental risk factor, and in susceptible individuals it causes an exaggerated inflammatory response that ultimately destroys the lung parenchyma (emphysema) and/or increases airway resistance by remodeling of the airway wall (203). It has long been known that the pathway varies between individuals; some patients have predominant emphysema while others can have similar degrees of airflow obstruction due to severe small airway disease with relatively preserved parenchyma, but the proportion and contribution of each to the pathogenesis of disease is still unknown. (175)

The pathologic events leading to emphysema are insidious and include structural and physiologic alterations that are characterized by inflammatory processes within the peripheral pulmonary parenchyma, thickening of arteriolar walls, and parenchymal destruction. A growing body of literature documents that these changes are likely to be associated with alterations in blood flow dynamics at a regional, microvascular level, and thus may serve as a beacon pointing toward the onset of early emphysema. Regional alterations in blood flow parameters may not only serve as an early marker for inflammatory processes but may also be a major etiologic component of the pathologic process, leading to emphysema in a subset of the smoking population (not all smokers have emphysema). (176)

Measures based on airflow or other measures of global lung function have reached their limits in their ability to provide new insights into the etiology of the disease, or even in leading us to an understanding of how lung volume reduction, in late stages of the disease, provides patient improvements. A number of articles have been written in which attempts are made to explain improvements of physiologic status post-LVRS (9, 16) on the basis of lung mechanics, and we find it difficult to understand how these relate to the observations from the NETT (15) showing that subjects with apical but not basal prevalence of disease receive the greatest benefit from surgery. However, if regional pulmonary perfusion is again brought into consideration, it makes sense that, if one removes apical lung that is not contributing well to gas exchange and blood is shunted to less diseased basal lung, gas exchange will be improved. Furthermore, by removing a diseased portion of the basal lung when the disease is predominantly basal, then it is likely that blood will be preferentially shunted to the contralateral basal lung. Using scintigraphy to assess regional V˙ /Q˙ , Moonen and colleagues (2) have recently concluded that an important mechanism for improvement in functional status post-LVRS relates to the reduction of regional shunt (i.e., blood flow may be directed toward regions of improved ventilation whereas regions receiving blood flow but that have poor ventilation are removed).

A recent international consensus statement on the diagnosis and therapy of COPD—the Global Strategy for the Diagnosis, Management, and Prevention of Chronic Obstructive Lung Disease (GOLD [Global Initiative for Chronic Obstructive Lung Disease])—has established diagnostic criteria that currently do not include CT findings (17). This is not surprising given that the consensus statement has been developed in part for the World Health Organization. It is notable that the summary makes the observation that different inflammatory events occur “in various parts of the lung,” a reference to the marked heterogeneity of the disease which cannot be defined without imaging. Of interest also are the future recommended research directions, which include identifying better defining characteristics of COPD, developing other measures to assess and monitor COPD, and recognizing the increasing need to identify earlier cases of the disease, all potential outcomes of improvements to quantitative lung imaging.

It has been well demonstrated that lung function declines with age (18–21) and perhaps also as a result of inflammation (22). This has confounded research related to the effects of smoking cessation on lung health. There are mixed results as to whether or not smoking cessation halts the progression of emphysematous lung disease (20, 23–27). Work by Bosse and colleagues (28) attempted to take into account the aging process and suggested that the disease process is slowed if one stops smoking. However, tests were not sensitive enough to conclude this definitively. There appear to be important sex differences in the effects of cigarette smoking and cessation (29). No reliable specific biochemical markers of disease presence or progression have been identified (30), in part perhaps because of the lack of a sensitive standard to diagnose and follow the diseases. More recently, CT parameters have been shown to be likely more sensitive to disease progression (31). Furthermore, the long time course of these diseases means that clinical trials using only whole lung function as primary outcome measures require huge numbers of subjects for extremely long periods of time.

Anatomic–Physiologic Correlates of Emphysema

The lack of a direct marker for emphysema has meant that epidemiologic studies have been limited to COPD, and these perhaps give a limited view as to the epidemiology of emphysema, a specific subset of COPD but not identified as such by spirometry. The direct effect of cigarette smoking on lung function has been widely studied, with differences in relative changes in FEV1 and the effects of smoking noted (28, 32–35). Although some studies show an increased rate of loss of FEV1 for current smokers, there is a less significant decrease in FEV1 for reformed Hoffman, Simon, and McLennan: CT-Based Lung Structure and Function 521 smokers (28, 32).However, more recent studies have found similar FEV1 declines with age in both smokers and never-smokers (34, 35). It must be emphasized that these changes are likely related to bronchial hyper-reactivity (36, 37) rather than to emphysema, highlighting again the need for objective measurement tools to assess emphysema. As indicated, chronic airflow limitation (COPD) is commonly seen in emphysema, but it is not essential. Measurements of lung physiology are not always able to distinguish the abnormalities that result from emphysema from those which result from the other causes of COPD, such as chronic bronchitis or asthma (38). The strongest positive association between an index of airflow limitation, FEV1 (% predicted) and a pathologically derived emphysema score comes from the National Institutes of Health Intermittent Positive-pressure Breathing Trial (39). There were only 48 subjects in this study, as autopsies were required for the pathologic assessment to be performed. Pulmonary function tests were performed every 3 months during the study, and were therefore available at some point before death. However, these subjects were highly selected; to enter the study, they were required to have very significant airflow obstruction, and could not be severely hypoxic; and to complete the study, they had to die during the observation period. In contrast, a study examining pathologic lung specimens taken during surgery, and appropriately fixed, showed no relationship between the pathologic emphysema rating and indices of airflow (40). Furthermore, an autopsy study enrolling 242 subjects over 6 years demonstrated that, although those subjects with greater pulmonary disability tended to have a greater degree of pathologic emphysema, 17 subjects with greater than 30% pathologic emphysema had no evidence for clinical COPD (41). Other pulmonary function tests—namely, diffusing capacity for carbon monoxide (DlCO) and the exponential description of the deflation pressure/volume curve (K)—have been used to identify, and to obtain a measure of, severity for pulmonary emphysema. A number of studies have found that measurement of DlCO has a very weak correlation with the pathologic assessment of emphysema (42–44). Measurements of elastic recoil pressure curves in life compared with pathologic assessment of emphysema at subsequent lung resection or postmortem have yielded conflicting results on the value of static compliance and K as a measure of emphysema (45–49). More recent studies show a weak but significant correlation between K and macroscopic emphysema (r _ 0.49) (47, 48), with K believed to be a measure of alveolar distensibility. This background highlights the continuing search for a marker for emphysema presence and severity.

Even though recent research has advanced our understanding of COPD pathogenesis, leading to the identification of potential targets and pathways for drug development, there are still major difficulties in conducting clinical trials designed to evaluate the benefits of new drug treatment for several reasons. These reasons include:

(1) the lack of validated short- and intermediate-term endpoints (or surrogates) that are predictive of future hard clinical outcomes, and

(2) the lack of a method to provide precise phenotypes suitable for large-scale studies.

It is for these reasons that computed tomography (CT) has become such an important tool in COPD research. CT provides a noninvasive method to obtain images of the lung that look similar to anatomic assessment, and CT images themselves are densitometry maps of the lung. Therefore any change in the structure of the lung will change the densitometry of the lung and, therefore, the image. Virtually every clinical center in all regions of the world has access to a CT scanner, so it is thought that CT images should be quite easy to obtain and it should be easy to conduct large, meaningful clinical studies.

<use a table like the example below>

Table Summarizing Precision/reproducibility of densitometry measurements from clinical studies reported in the literature

| Scan | Reader | # of Readers | # of Patients | # of Nodules | Lesion Size, Mean (range) | Organ System | Densitometry, 95% CI of Measurement Difference | Densitometry, Measurement Difference % | 1D Measurement, 95% CI of Measurement Difference | 1D, Mean Measurement Difference % | Slice Thickness /Recon Interval, mm | Author, Year |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| repeat scans | intra-reader | 1 | 20 | 218 | 9.85 mm | lung, mets | -21.2 to 23.8% | 1.30% |  |  | 1.0/0.7 | Gietama et al. 2007 [9] |
| repeat scans | intra-reader | 3 | 32 | 32 | 38 mm (11–93 mm) | lung, NSCLC | -12 to 13.4% | 0.70% | -7.3% to 6.2% | -0.60% | 1.25/1.25 | Zhao et al. 2009 [11] |
| same scan | intra-reader | 1 | 10 | 50 | 6.9 mm (2.2–20.5 mm) | lung, mets | -3.9 to 5.7% | 0.90% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [4] |
| same scan | inter-reader | 2 | 10 | 50 | 6.9 mm (2.2–20.5 mm) | lung, mets | -5.5 to 6.6% | 0.50% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [4] |
| repeat scans | not specified | not specified | 10 | 151 | 7.4 (2.2–20.5 mm) | lung, mets | -20.4 to 21.9% | 1.50% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [4] |
| repeat scans | not specified | not specified | 10 | 105 | <10 mm | lung, mets | -19.3 to 20.4% | 1.70% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [4] |
| same scan (5 sets, 1 set/phase) | intra-reader ? (consensus by 2 readers), 3 x reading | 2 | 30 | 73 | ~1–9 mm [25.3 (0.2–399 mm3)] | lung, noncalcified nodules | coefficient of variance as large as 34.5% (95% CI not reported) | not reported | not reported | not reported | 0.75/0.6 | Boll et al. 2004 [27] |
| same scan | inter-reader | 2 | 33 | 229 | 10.8 mm (2.8–43.6 mm), median 8.2 mm | lung, primary or mets | -9.4 to 8.0% | 0.70% | -31.0 to 27% | -2.00% | 1.0/0.8 | Hein et al. 2009 [12] |
| same scan | inter-reader, inter-algorithms (6 readers x 3 algorithms) | 6 | 16 | 23 | not reported | lung, nodules | 55% (upper limit) | not reported | not reported | not reported | 1.25/0.625 | Meyer et al. 2006 [28] |
| same scan | intra-reader | 2 | 50 | 202 | 3.16–5195 mm3, median 182.22 mm3 | lung, mets | % not reported | 0.15 to 0.22% | % not reported | 2.34–3.73% (p<0.05 1D vs 3D) | 0.75/0.70 | Marten et al. 2006 [7] |
| same scan | inter-reader | 2 | 50 | 202 | 3.16–5195 mm3, median 182.22 mm3 | lung, mets | % not reported | 0.22 to 0.29% | % not reported | 3.53–3.76% (p<0.05 1D vs 3D) | 0.75/0.70 | Marten et al. 2006 [7] |
| same scan | inter-reader | 2 | 2239 | 4225 | 15–500 mm3 (effective diameter 3.1–9.8 mm) | lung, nodules | -13.4 to 14.5% | 0.50% | not reported | not reported | 1.0/0.7 | Wang et al. 2008 [10] |
| same scan | intra-reader | 2 | 24 | 52 | 8.5 mm (<5 to 18 mm) | lung, noncalcified nodules | 8.9 % (upper limit) | not reported | not reported | not reported | 1.25 or 2.5/not specified | Revel et al.[6] |
| same scan | inter-reader (3 readers x 3 measurements) | 3 | 24 | 52 | 8.5 mm (< 18 mm) | lung, noncalcified nodules | 6.38 % (upper limit) | not reported | not reported | not reported | 1.25 or 2.5/not specified | Revel et al. [6] |

Abbreviations: 1D = unidimensional; mets = metastasis; CI = confidence interval

## Conventions and Definitions

Acquisition vs. Analysis vs. Interpretation: This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be computational steps that transform the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge. (The transformation of knowledge into wisdom is beyond the scope of this document.)

Other Definitions:

Image Analysis, Image Review, and/or Read: Procedures and processes that culminate in the generation of imaging outcome measures, such tumor response criteria. Reviews can be performed for eligibility, safety or efficacy. The review paradigm may be context specific and dependent on the specific aims of a trial, the imaging technologies in play, and the stage of drug development, among other parameters.

Image Header: The Image Header is that part of the file or dataset containing the image other than the pixel data itself

Imaging Phantoms: Devices used for periodic testing and standardization of image acquisition. This testing must be site specific and equipment specific and conducted prior to the beginning of a trial (baseline), periodically during the trial and at the end of the trial.

Intra-Rater Variability is the variability in the interpretation of a set of images by the same reader after an adequate period of time inserted to reduce recall bias.

Inter-Rater Variability is the variability in the interpretation of a set of images by the different readers.

A Time Point is a discrete period during the course of a clinical trial when groups of imaging exams or clinical exams are scheduled.

## Breathing Instructions

Rigorous adherence to instructing the patient how to breathe before the scan is obtained in order to attain TLC or RV volume before each chest CT is performed. These must be done verbally as the breathing should mimic breathing instructions typically provided in Pulmonary Function Laboratories.

Breath holding reduces motion which might degrade the image. Full inspiration inflates the lungs which is necessary to separate structures and make lesions more conspicuous.  Full expiration allows for measurement of air-trapping for determination of amount of emphysema present.

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| Breath hold | Acceptable At least near the high end inspiration & expiration Target Subjects should be instructed to hold a single breath at full inspiration and full expiration Ideal Subjects holding at max inspiration and max expiration using the instructions outlined. |

These are instructions given by the Technologist which are listed in sequence of scan acquisition.  Breathing instructions are for two standard scout scans, PA and Lateral, and two lung volumes: inspiration or total lung capacity (TLC) and expiration or residual volume (RV) volumetric scans.

* Side notes for technologist in italics

It is important for the Technologist to ensure that the subject is able to follow the commands appropriately and also that the subject is appropriately coached PRIOR to performing the actual imaging procedures. Practice breathing maneuvers are recommended and built into the breathing instructions to reduce subject anxiety and limit mistakes during actual scan acquisition.

(a) Practice Breathing (TLC)

For the first part of this scan, I am going to ask you to take a couple of deep breaths in and out before we have you hold your breath all the way in.

First let’s practice

Take a deep breath in (watch chest to ensure a deep breath as far in as possible then start scan)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Now breathe all the way IN…IN…IN… (watch chest to ensure a deep breath in as far as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure spine remains on the table, patients is not shaking- watch for these throughout the study!)

Breathe and relax

(b) Scout views: PA & Lateral (TLC)

OK let’s get started

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Now breathe all the way IN... IN... IN... and hold it (watch chest to ensure a deep breath as far in as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure the spine remains on the table, subject is not shaking and then start the scout scan)

Perform scout

At the end of the scout- Breathe and relax

Scout views: Lateral (TLC)

OK let’s get started

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Now breathe all the way IN... IN... IN... and hold it (watch chest to ensure a deep breath as far in as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure the spine remains on the table, subject is not shaking and then start the scout scan)

Perform scout

At the end of the scout- Breathe and relax

(c) Inspiratory CT (TLC)

Now we’re ready again so please

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Now breathe all of the way IN...IN...IN as far as possible and hold it in (watch chest to ensure a deep breath in as far as possible)

Keep holding your breath – DO NOT BREATHE!

At the end of scan- Breathe and relax

(d) Practice Breathing (RV)

For the second part of this scan, I am going to ask you to take a couple of deep breaths in and out before we have you hold your breath all the way out.

First let’s practice

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take another deep breath in (watch chest to ensure a deep breath in)

Now breathe all the way OUT…OUT…OUT…as far as possible and hold it out (watch chest to ensure a deep breath out as far as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure spine remains on the table, patients is not shaking- watch for these throughout the study!)

Breathe and relax

(e) Scout views: PA & Lateral (RV)

Scout views: PA (RV)

OK let’s get started

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take another deep breath in (watch chest to ensure a deep breath in)

Now breathe all the way OUT…OUT…OUT…as far as possible and hold it out (watch chest to ensure a deep breath out as far as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure spine remains on the table, patients is not shaking- watch for these throughout the study!)

Breathe and relax

Scout views: Lateral (RV)

OK let’s get started

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take another deep breath in (watch chest to ensure a deep breath in)

Now breathe all the way OUT…OUT…OUT…as far as possible and hold it out (watch chest to ensure a deep breath out as far as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure spine remains on the table, patients is not shaking- watch for these throughout the study!)

Breathe and relax

(f) Expiratory CT (RV)

Now we’re ready again so please

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take another deep breath in (watch chest to ensure a deep breath in)

Now breathe all of the way OUT...OUT...OUT… as far as possible and hold it out (watch chest to ensure a deep breath out as far as possible)

Keep holding your breath – DO NOT BREATHE!

At the end of scan- Breathe and relax

## Model-specific Instructions and Parameters

Compliance with a profile involves meeting a variety of requirements of which operating by these parameters is just one. To determine if a product (and a specific model/version of that product) is compliant, please refer to the Compliance section above.

Sites using models listed here are encouraged to consider using these parameters for both simplicity and consistency. Sites using models not listed here may be able to devise their own settings that result in data meeting the requirements but this is outside the formal scope of QIBA compliance. Use of a product as indicated here, and according to these settings, is expected to achieve compliance with the profile.

In some cases, parameter sets may be available as an electronic file for direct implementation on the imaging platform.

Table Model-specific Parameters for Specific Acquisition Devices Demonstrated to Achieve Compliance

IMPORTANT NOTE with respect to this example table: The presence of specific product models/versions in the following tables, as well as specific settings shown for them, shall not be taken to imply that those products are fully compliant with the QIBA Profile.  These settings were determined by the team in the 1C study as an example of how it could be done but more strict attention to all parameters identified in the Profile are necessary in order for a company to claim any particular model is compliant.  That said, we appreciate the good will and help that the vendors represented here have provided in this early phase of QIBA.

| Acquisition Device | Product Setting to Achieve Compliance Levels |
| --- | --- |
| GE Discovery HD750 sct3 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.625 Gantry Rotation Time in seconds 1 mA 120 Pitch 0.984 Scan FoV Large Body (500mm) |
| Philips Brilliance 16 IDT mx8000 | kVp 120 Number of Data Channels (N) 16 Width of Each Data Channel (T, in mm) 0.75 Gantry Rotation Time in seconds 0.75 Effective mAs 50 Pitch 1.0 Scan FoV 500 |
| Philips Brilliance 64 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.625 Gantry Rotation Time in seconds 0.5 Effective mAs 70 Pitch 0.798 Scan FoV 500 |
| Siemens Sensation 64 | kVp 120 Collimation (on Operator Console) 64 x 0.6 (Z-flying focal spot) Gantry Rotation Time in seconds 0.5 Effective mAs 100 Pitch 1.0 Scan FoV 500 |
| Toshiba Aquilion 64 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.5 Gantry Rotation Time in seconds 0.5 mA TBD Pitch .828 Scan FoV Medium and Large |

Table Model-specific Parameters for Specific Reconstruction Software Demonstrated to Achieve Compliance

IMPORTANT NOTE: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  These settings were determined by the team in the 1C study as an example of how it could be done but more strict attention to all parameters identified in the Profile are necessary in order for a company to claim any particular model is compliant.  That said, we appreciate the good will and help that the vendors represented here have provided in this early phase of QIBA.

| Reconstruction Software | Product Setting to Achieve Compliance Levels |
| --- | --- |
| GE Discovery HD750 sct3 | Reconstructed Slice Width, mm 1.25 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel STD |
| Philips Brilliance 16 IDT mx8000 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm (contiguous) Display FOV, mm 350 Recon kernel B |
| Philips Brilliance 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm (contiguous) Display FOV, mm 350 Recon kernel B |
| Siemens Sensation 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel B30 |
| Toshiba Aquilion 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm Display FOV, mm TBD Recon kernel FC12 |

Table Model-specific Parameters for Specific Image Analysis Software Demonstrated to Achieve Compliance

IMPORTANT NOTE: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  In particular, the following example table only has placeholders for these example products which need to be replaced with product model-specific settings in order to claim compliance.

| Image Analysis Software | Product Setting to Achieve Compliance Levels |
| --- | --- |
| Siemens LunCARE | a <settings to achieve…> b <settings to achieve…> c <settings to achieve…> d <settings to achieve…> |
| GE Lung VCAR | e <settings to achieve…> f <settings to achieve…> g <settings to achieve…> h <settings to achieve…> |
| R2 ImageChecker CT Lung System | i <settings to achieve…> j <settings to achieve…> k <settings to achieve…> l <settings to achieve…> |
| Definiens (name specific product) | m <settings to achieve…> n <settings to achieve…> o <settings to achieve…> p <settings to achieve…> |
| Median (name specific product) | q <settings to achieve…> r <settings to achieve…> s <settings to achieve…> t <settings to achieve…> |
| Intio (name specific product) | u <settings to achieve…> v <settings to achieve…> w <settings to achieve…> x <settings to achieve…> |

# Comments

Compare with DCEMRI claim

Insert data from Iowa, COPDGene, SARP, UCLA, ? Europe (Stoel/Wigstrom)

Corrected for lung volumes

995-1000

Needs discussion

? do we need this

? delete

Need to resolve attributions