## Sheet1
SPECT Biomarker Committee
Point-by-point response to public comments about Version 1.0 of Tc99m SPECT Profile
SPECT-CT: Estimating the Concentration of 99mTc Based Imaging Biomarkers in Large and Small Volumes of Interest
# | experts' names | Line # | Section # | Issue raised by public experts | suggestion raised by public & their editorial comments | BC response | (concrete) proposal
1 | Tateishi (Japan) | 95 | Q1 | Determination of minimal acceptable counts and optimal counts | Random Poisson sampling technique is considered to be suitable for evaluating minimal acceptable counts and optimal counts. At present, however, GE is the only manufacturer that provides system with this technology. Therefore, this technology cannot be used by users of other manufacturers. Manufacturers not equipped with this technology should be encouraged to develop and provide this technology from RSNA QIBA. | We appreciate the comment that Poisson resampling is the best technique for evaluating minimal acceptable counts and optimal counts. It is unfotunate that only one vendor currently offers this tool with their workstation software. We can recommend it, but not mandate it.
2 | Tateishi (Japan) | 95 | Q5 | Issue about the measurand. | I think we should basically use kBq/cc. KBq/cc should be an absolute indicator if the QC of SPECT system is normal. | We appreciate the comment. There were some on the committee who also preferred to use kBq/CC. In the end the committee decided to include both kBq/CC and target-to-background ratio.
3 | Tateishi (Japan) | 455 (now line 466) |  | Supply of NIST-traceable 99mTc | Since there is no system in Japan to purchase or get NIST-traceable 99mTc, it is considered difficult to conduct an accuracy test using NIST-traceable 99mTc. In Japan, the accuracy test and the consistency test are usually conducted by manufacturer of calibrator. | We appreciate the comment. Text has been revised to state NIST-traceable (or equivalent) 99mTc standard … It is our hope that this will make conformance possible everywhere in the world that quantitative SPECT is indicated. See line 466 (Table, Row 1) | Revised text.
1 | Laforest (Wash U) | 121-131 | 1. Executive Summary | This paragraph does not make sense. | re-write | We appreciate the commnent. The paragraph has been revised to clarify the content. It has again been vetted by our committee who hope it is now more understandable. | Revised text.
2 | Laforest (Wash U) | 168 | 1. Executive Summary | Specify what this profile 'Help' | Suggest to add the Profile helps clinicians and .. make informed decision on the accuracy and reproducibility of SPECT derived biomarkers. | We appreciate the commnent. The paragraph has been revised. Committee added to additional terms to clarify what 'help' means. | Revised text.
3 | Laforest (Wash U) | 216 | 2. Clinical Context and Claims | 3.37-4.63 | Did math wrong. Had wrong range listed. | We appreciate correcting our math error. Range has been fixed. | Math error fixed.
4 | Laforest (Wash U) | 224 | 2. Clinical Context and Claims | There should be dispositions to ensure quality control for absolute camera calibration. |  | We appreciate the comment. Methods to provide quantitative calibrations are now better described in Section 4.7. | Refer users to Section 4.7 for methods on how to apply quantitative corrections.
5 | Laforest (Wash U) | 264 (now Line 266) | 2. Clinical Context and Claims | Add same blood flow as the target. |  | Thank you. We agree. Text now states "Same biological properties as the target, including blood flow, or a known method of compensating for differenes, such as blood flow." | Revised text.
6 | Laforest (Wash U) | 285 | 2. Clinical Context and Claims | Should there be a need to add that the tracer uptake with the region of interest be constant over the time of data acquisition. |  | We appreciate the comment. Text now states "The tracer uptake within the region of interest can be considered constant over the time of data acquisition up to the point of saturation in situations where non-tracer doses of the pharmacophore are present with the tracer." | Revised text.
7 | Laforest (Wash U) | 340 | 2. Clinical Context and Claims | This line appears twice. |  | Thank you for finding the typo. It has been corrected. | Repeated line has been removed.
8 | Laforest (Wash U) | 469 | 3.5.3 Ancillary Equipment | Large sky scapers? |  | We appreciate the comment. Text has been revised. Text is provided to make sure that time of study and activity measurements are synchronized with an international standard. | Revised text.
9 | Laforest (Wash U) | 210 | 2. Clinical Context and Claims; and Section 3.9 | Do the 2 million refers to the 2 million per projection view (section 3.9) or per image plane? |  | We appreciate the question. The text on line 621 provides instructions on how to determine the time per view angle. First one needs to determine the number of events to collect per view (i.e., equation is 2 million counts/number of projection views) and then one can determine the time per view by dividing the number of counts to acquire per view by the count rate of the detector. | Clarification of text.
10 | Laforest (Wash U) |  | Document | Many other edits needed throughout text. |  | We appreciate the comment. The committee has gone through the whole document and tried to edit for content. | Committee went through whole document and revised for content.
1 | Robinson (NPL, UK) | 450-460 | 3.5.3 Ancillary Equipment | The guidance on using radionuclide calibrator is potentially misleading. Performing a monthly accuracy check with a traceable Tc-99m source is likely to be not possible. Acceptance of a 2.5% variation from the "standard" value without considering the statistical significance and an uncertainty budget for the calibrator does not define the uncertainty on the calibrator measurements. "shall confirm the measured and decay corrected values activity..." - this should just be the decay corrected activity, the ratio between measured and decay corrected values depends on the reference time and so they may not be both within 2.5% of the "standard value" | refer to https://www.npl.co.uk/special-pages/guides/establishing-maintaining-calibration-radionuclide.pdf or similar document when revising this section. Consultation with an appropriate national measurement institute may be beneficial. | We appreciate the comment. We consulted with two individuals from NIST and believe that the statement reflects the best that can be implemented globally at all the sites we are trying to harmonize at this time. | Consulted with two individuals from NIST.
2 | Robinson (NPL, UK) | 210 | 2. Clinical Context and Claims | No evidence is supplied to support the 8% COV claim for activity concentration measurement. Without this evidence it is hard to evaluate the applicability of the document. | Provide the data/references in the report. | We appreciate the comment. We agree it is difficult to evaluate COV. The 8% is based upon the experience from most clinical trialists in the group conducting multi-site clinical trials. | 8% COV was selected based upon experience of clinical trialists on the committee.
3 | Robinson (NPL, UK) | 524 | 3.6 Imaging Protocol Design | "no larger than 4mm voxel" | P51 "no larger than 4.8 mm voxels". P62 "voxel size <= 4.8 mm" | We appreciate the comment. Other reviewers also felt that 4.8 mm was too small. Therefore we have adjusted up to <= 5 mm and made this consistent through out the document. | The text has been revised and made consistent throughout the document.
1 | Knowland (Lucerno Dynamics) | 695 | 3.11.2 Specification | Injection site is to be imaged if count rate is low in order to "make sure there is little to no extravasation at the injection site." | Recent published evidence suggests that extravasation of the injection resolves over time due to biological clearance. This resolution can result in little to no extravasation visible at the imaging time, but dramatically alters the kinetics of biodistribution throughout the uptake time. Technologists should employ other means of ensuring proper delivery of the radiopharmaceutical and not rely on injection site imaging alone as kinetics and quantification can be affected in unpredictable ways. | We appreciate the comment. Suggestion accepted and text has been revised. | Text has been revised based upon comments from reviewer.
1 | AAPM | 342 | 2. Clinical Context and Claims | This sentence is redundant (i.e., a repeat of the sentence immediately above it). |  | We appreciated the comment. We have removed the redundant sentence. | Text has been revised.
2 | AAPM | 344 | 3. Profile Activities | In the Section column, there is an extra "." after 3.0 and 3.10. |  | We appreciated the comment. We have removed the the extra ".'s". | Text has been revised.
3 | AAPM | 387 | 3.2 Product Validation | Table 3.2.2: We question how does the 16% wCV limit for Target Activity Concentration or Target to Background Ratio for Image Analysis Tool relate to the 22% specified earlier? |  | Text has been corrected. The 16% wCV limit for Target Activity Concentration was an error. It has been corrected to state 22%. Explanation on how 22% was derived is included in Table 3.2.2. | Text has been revised.
4 | AAPM | 413 | 3.4 System Acceptance Tests | Add AAPM reference (See Report No. 177) |  | We appreciated the comment. The reference has been added. | The reference has been added.
5 | AAPM | 443 | 3.5 Periodic QA | Table 3.5.2: CT Attenuation map registration. We believe that daily is too onerous, and we would recommend that it be changed to quarterly. |  | We appreciate the comment. The text has been revised to state that daily is ideal and quarterly is the minimum acceptable standard. | The text has been revised.
6 | AAPM | 443 | 3.5 Periodic QA | Table 3.5.2: Overall System Performance. No frequency is specified. |  | Line 434 now reads, "measures . . . should be monitored on a regular basis. Manufacturers’ recommendations and institutional policy may require additional or more frequent tests." | The text has been revised.
7 | AAPM | 459 now line 464 | 3.5.3 Ancillary Equipment | Table: Accuracy. No NIST-traceable Tc-99m or simulated radionuclide calibrator standards are currently available. We would suggest allowing for the use of a Co-57 evial as a surrogate for this test. Also note that the k=2 (95% CI) and uncertainty of any such standard should be much less than 2.5%. |  | Now accounted for. Text now states: "NIST or equivalent (e.g., from a nuclear pharmacy with established traceablity)" | The text has been revised.
8 | AAPM | 459 | 3.5.3 Ancillary Equipment | Linearity is a quarterly test per regulation, whereas, Accuracy is annual. |  | Thank you. The text has now been revised as suggested to conform with the minimum regulatory standards. | The text has been revised.
9 | AAPM | 490 | 3.6 Imaging Protocol Design | AAPM recommends caution regarding this requirement. A less than 8 mm at 10 cm limit constrains the collimation to low-energy high (or ultra-high)-resolution, which constrains counting sensitivity (a big problem with SPECT vs. PET imaging). Iterative reconstruction techniques include resolution recovery, which can compensate for the use of higher efficiency but lower-resolution collimation 8mm resolution requirement may need discussion. |  | The text has now been expanded as suggested to (1) note the issue you raised, and (2) verify that LEAP collimators are unlikely to meet minimum specifications for conformance. | The text has been revised.
10 | AAPM | 514 | 3.6 Imaging Protocol Design | If what is meant by "total system cps" is that for each projection (as opposed to the total for two heads combined on a dual-head system), then total number of time frames to use in the formula should be independent of the number of heads. For example, if the total system cps is 1,000, and there are two heads and 120 total projections over 360 degrees, the time per view calculated by the current formula would be (2,000,000/1,000)/60 = 33.3 seconds; but the total acquired counts would be 1,000x33.3x120 = 4,000,000, which would be incorrect. |  | "Total system cps" refers to the total count rate of all detector heads summed together, i.e., the "total" system count per second. It is not the count rate for each individual detector head. Therefore, we belive that the formula is correct as written.
11 | AAPM | 523 | 3.6 Imaging Protocol Design | “4mm” should be “5 mm” to match that specified in Table 3.6.2. |  | The text has now been revised as suggested to read, "less than or equal to 5 mm" | The text has been revised.
12 | AAPM | 535 | 3.6 Imaging Protocol Design | Table. 3.6.2: Reconstruction Protocol. We believe iterative reconstruction with CT-base attenuation correction, scatter correction and resolution recovery should be a requirement. |  | We agree in principle. However, QIBA does not typically dictate how specifications are met. We now note that, just for example, some new machine learning and artificial intelligence methods may be better than iterative reconstruction. | The text has been revised.
13 | AAPM | 561 | 3.8 Subject Handling | We believe it should be "intravenous" and "arterial". |  | The text has now been expanded as suggested to handle almost any route of administration. | The text has been revised.
14 | AAPM | 594 | 3.9 Image Data Acqusition | What is meant by "laying the subject on a calibrator"? We recommend clarification. |  | The text has now been deleted as suggested, since it applies to planar images of neonates and body parts, such as a wrist, are typically taped directly on a collimator (covered by a blue chux), not calibrator (typo). | The text has been revised.
15 | AAPM | 605 now line 612 | 3.9 Image Data Acqusition | Typical LEGP (a.k.a. LEAP) collimators do not meet the 8 mm at 10 cm limit. |  | The text has now been revised as suggested to observe that LEAP collimators seem unlikely to meet the specifications, but several others might. The text now better emphasizes the conformance standard, not the means by which an innovator might achieve it. | The text has been revised.
16 | AAPM | 620 | 3.9 Image Data Acqusition | Table 3.9.2: SPECT acquisition mode. 4.8 mm should be 5 mm. (The 128x128 natural pixel size for the GE 870 CZT pixelated camera is 4.92 mm. The software allows interpolated 6 smaller pixel sizes, e.g., 4.42 mm, but sinc filtering of the projections is applied to achieve that). |  | Thank you. The text has now been revised as suggested to state "less than or equal to five mm." | The text has been revised.
17 | AAPM | 685 | 3.10 Image Data Reconstruction | SPECT matrix/voxel size. "4.8 mm" should be "5 mm” (see comment above). |  | Thank you. Same as above. AAPM comment #11. | The text has been revised.
18 | AAPM | 850 now line 866 | 4.1 Assessment Procedure: Intrinsic Uniformity | The "distance of five times the longest dimension of detector's field-of-view" would be the diagonal of a rectangular detector. |  | Thank you. We have now clarified the text to state, "distance of five times the diagonal length of the detector." | The text has been revised.
19 | AAPM | 863 now Line 880 | 4.1 Assessment Procedure: Intrinsic Uniformity | Placing a sheet source in contact with the face of the collimator is not possible on systems with auto-contouring light rails or touchpads that will not allow acquisition if contacted. |  | We agree that defeating these safety features is easier on some instruments that others. We have now revised the text to state "on, or as close as possible" and noted that imaging scientists may use alternative methods when necessary as long as the specifications are met. | The text has been revised.
Legend:
|  |  | 2. Priority Key: | L: Low. Typo or other minor correction that an editor can manage; requires no group discussion. | M: Medium issue or clarification. Requires discussion, but should not lead to long debate. | H: High. Important issue where there is a major issue to be resolved; requires discussion/debate.
|  |  | 5. Issue | Describe your issue; include enough to indicate what you see as a problem.
|  |  | 6. Proposal | Propose a resolution to your issue, e.g., suggested new wording or description of a way to address the issue; leave blank if no resolution can be provided.