Draft QIBA profile for fMRI brain mapping:

- Intended use

fMRI is used as a tool for diagnosis and pre-treatment planning in patients with brain disorders. The presenting symptoms and location of the affected brain tissue determine the particular region or regions of the brain to be investigated and the stimulus paradigm(s) selected (e.g. motor task, language task). The change in BOLD signal (relative to a rest condition) provides information about the brain region(s) involved in task performance and the proximity of this functioning cortex to brain tissue to be treated.  Pre-treatment planning endpoints involve risk assessment (impact of treatment course on functioning cortex, e.g. surgical pathway, tissue resection, or off-target irradiation) and predictive value assessment (will loss of a particular region of functioning cortex result in a deficit) that influence patient treatment. Our goal is to determine the quantitative parameters under which BOLD fMRI acts as reproducible and reliable predictor of cortical function, that is, as a valid imaging biomarker for medically meaningful change in brain activity.

2. Claims

A. fMRI can provide brain maps (images) of function-related changes in blood flow and oxygenation evoked by a specific behavioral task paradigm. These maps consist of a 3-dimensional spatial distribution of individual sites (voxels) each having a response to the behavioral paradigm that is statistically different from random variation (fMRI noise). Using currently available technology, the effective spatial resolution of these maps is typically 2-4?? mm.

B. Single sites (voxels) of fMRI activation can be determined to be statistically different from random variation (fMRI noise) at a specified confidence level (false positive probability). Acceptable false positive probabilities are typically in the range of ??-?? for single sites which yields a whole-brain false positive probability of ??-?? after accounting for multiple comparisons related to the number of brain voxels. Alternately, groups of activated voxels can be identified by clustering methods and a false positive rate for clusters of a given size or greater can be established. Current practice is to use minimum cluster volumes of ??-?? which have an acceptable false positive probability in the range of ??-??

C. The fMRI brain maps (A) can be co-registered with other types of MRI images (or other imaging technologies) to permit measurement and visualization of the spatial relationships between fMRI activation sites and brain pathology (e.g. tumors), white matter tracts (DTI) and treatment scenarios (eg. proposed resection plan, radiation dosage maps). Co-registration can typically be achieved to an accuracy of ??, but is dependent on the degree of distortion of the fMRI images relative to the other types of images.

D. Quantitative measures of “risk” to eloquent brain structures…

3. Protocol

A. Patient evaluation and fMRI paradigm selection

B. Paradigm design/selection

C. Patient training on behavioral task and quantitative assessment of performance

D. Equipment Q/A checks and calibrations

E. Patient preparation, instructions, behavioral check, adjustment of peripheral equipment

F. Pre-scan setup, shimming, selection of slices and other imaging parameters

G. Performance of fMRI scans including patient task instructions, performance monitoring

H. Post-scan evaluations of alertness, performance

I. Near-real time scan evaluation, repeat scans

J. Post processing, artifact detection etc.

K.  Report generation and content, including technical Q/A and evaluation

L. Clinical interpretation

M. Archiving and export to treatment systems

N. Follow-up

4. Compliance checks

fMRI signals reflect changes in local blood oxygenation and volume triggered by net changes in local neuronal activity. Such signals are therefore an indirect measure of brain activity. Moreover, the functional specificity of the fMRI activity depends on the behavioral task paradigm and the patient’s ability to perform the task. Thus, fMRI signals and the ability to identify functionally-specific eloquent cortex can be compromised by:

1. Neurovascular uncoupling (NVU) – impairment of the hemodynamic mechanism linking changes in neuronal activity to changes in blood flow and oxygenation and thus to changes T2* that are detectable by the MRI scanner.  NVU poses a problem for clinical interpretation, since it can cause a loss of fMRI signal even though the underlying neural activity is intact. This could lead to the incorrect clinical interpretation that a region of brain tissue is non-functional and can be resected or irradiated without loss of function.

2. Poor behavioral compliance – Since the fMRI signal depends on the patient’s ability to perform the required behavioral task and thereby activate neurons in task-related brain areas, the occurrence, magnitude and extent of fMRI activation can vary in proportion to the accuracy, reliability, speed, etc of the patient’s performance. Thus if undetected, poor performance can reduce or eliminate an fMRI signal leading to clinical mis-interpretation of the fMRI brain maps.

3. Lack of functional specificity – The areas of brain activation produced by a specific behavioral paradigm can reflect all aspects of the task, both intended and unintended. Consequently, the resulting fMRI activation may reflect both intended functions (eg. hand movement) and unintended functions (eg. postural adjustments). For the intended use in clinical treatment planning and guidance, this poses the problem that not all sites of fMRI activation may necessarily cause a deficit in the intended behavioral function if damaged.

QIBA Profiles consist of:

1. Intended use (clinical context), e.g. risk assessment or clinical trials

2. Claims (level of performance, both state-of-the-art and areas for improvement)

3. Protocol (structure provided by UPICT template)

4. Compliance check

It tells a user what can be accomplished by following the Profile. ("Profile Claims")

E.g. you will be able to detect volume changes of greater than <20%> in Stage IV Lung Nodules which are <5mm> in diameter or greater.

It tells a vendor what they must implement in their product to state compliance with the Profile. ("Profile Details")

E.g. to comply, the scanner must be able to:

- scan a <Mark-324> Chest Phantom, identify the smallest resolvable target, display the diameter of that target

- demonstrate resolving targets at least as small as <2mm> diameter on the Mark-324 phantom

- scan patients according to the ACRIN NLST acquisition protocol

E.g. to comply, the quantification application must be able to:

- segment a nodule (automatically or manually), derive the volume, store it in a DICOM object

- run a user through a set of test data with known volumes and at the end display an accuracy score

It may also tell the user staff what they must do for the Profile Claims to be realized. ("Profile Details")

E.g. to comply, the site CT techs must be able to:

- scan the patient within 10 minutes of contrast injection

E.g. to comply, the radiologist must be able to:

- achieve a score of <95%> or better using their segmentation application on the <Lung Invaders> test set.

From the UPICT Template

Acquisition vs. Analysis vs. Interpretation

This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge.

Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject.  Analysis is primarily considered to be computational steps that transform the data into information, extracting important values.  Interpretation is primarily considered to be judgment that transforms the information into knowledge.

(The transformation of knowledge into wisdom is beyond the scope of this document.)

Bulls-eye Compliance Levels

Acquisition parameter values and some other requirements in this protocol are specified using a “bullseye” approach.  Three rings are considered from widest to narrowest with the following semantics:

ACCEPTABLE: failing to meet this specification will result in data that is likely unacceptable for the intended use of this protocol.

TARGET: meeting this specification is considered to be achievable with reasonable effort and equipment and is expected to provide better results than meeting the ACCEPTABLE specification.

IDEAL: meeting this specification may require unusual effort or equipment, but is expected to provide better results than meeting the TARGET.

An ACCEPTABLE value will always be provided for each parameter.  When there is no reason to expect better results (e.g. in terms of higher image quality, greater consistency, lower dose, etc.), TARGET and IDEAL values are not provided.

Some protocols may need sites that perform at higher compliance levels do so consistently, so sites may be requested to declare their “level of compliance”.  If a site declares they will operate at the TARGET level, they must achieve the TARGET specification whenever it is provided and the ACCEPTABLE specification when a TARGET specification is not provided.  Similarly, if they declare IDEAL, they must achieve the IDEAL specification whenever it is provided, the TARGET specification where no IDEAL level is specified, and the ACCEPTABLE level for the rest.