## UMMS
QIBA 1C Performance Based Protocol Worksheet
Step 1 - scanner desiption
Scanner Manufacturer | Philips
Scanner Model | 64
Site id for scanner | UMMS - IR
Software version
Date of Testing | 2010-03-16
Steps 2-4: Scanning Parameters that Meet Performance Metrics
Step 2: kVp selected | 120
Step 3: Scan Field of View (SFOV) or Patient Size (in mm) | 250
Step 4: Reconstructed Slice with Selected (in mm) | 1
Step 5: Collimation Rotation Time and Pitch
Number of Data Channels (N) | 64
Width of Each Data Channel (T, in mm) | 0.625
Gantry Rotation Time in seconds | 0.5
Pitch (close to 1.0 is preferred) | 0.891
Time to cover a 35 cm long thorax (desired < 15 seconds) | 6.006
Step 6: Iterating on Recon Kernels to obtain 7 lp/cm
If using mA, then mA setting | 446
(if using mA, then effective mAs equals this value): | 250.2806
if using mAs/rotation (philips) or effective mAs (Siemens), the record setting here | Philips (effective mAs)
Recontruction Filter setting (name) - Iteration 1 | Sharp (C)
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Recontruction Filter setting (name) - Iteration 2
Size of line pairs seen with this filter (Level =1100, Win=100)
Recontruction Filter setting (name) - Iteration 3
Size of line pairs seen with this filter (Level =1100, Win=100)
Name of Filter that meets 7 lp/cm performance criteria | Sharp(C)
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter
Step 7: Iterating on mA (or mAs) settings to obtain std dev of 12 +/1-1 HU
mA or mAs or Eff. mAs setting - Iteration 1 | 250
Standard deviation measured on image of ACR phantom | 11.18
mA or mAs or Eff. mAs setting - Iteration 2
Standard deviation measured on image of ACR phantom
mA or mAs or Eff. mAs setting - Iteration 3
Standard deviation measured on image of ACR phantom
mA or eff. mAs setting that meets Standard Deviation Performance criteria | 250
Scanner reported CTDIvol (in mGy) using this mA or eff. mAs setting | 14.7
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter

## DUKE
QIBA 1C Performance Based Protocol Worksheet
Step 1 - scanner desiption
Scanner Manufacturer | GE
Scanner Model | 64
Site id for scanner
Software version
Date of Testing
Steps 2-4: Scanning Parameters that Meet Performance Metrics
Step 2: kVp selected | 120
Step 3: Scan Field of View (SFOV) or Patient Size
Step 4: Reconstructed Slice with Selected (in mm) | 1.25 mm
Step 5: Collimation Rotation Time and Pitch
Number of Data Channels (N) | 64
Width of Each Data Channel (T, in mm) | 0.625
Gantry Rotation Time in seconds | 1
Pitch (close to 1.0 is preferred) | 0.984
Time to cover a 35 cm long thorax (desired < 15 seconds) | 8.8923
Step 6: Iterating on Recon Kernels to obtain 7 lp/cm
If using mA, then mA setting
(if using mA, then effective mAs equals this value): | 0
if using mAs/rotation (philips) or effective mAs (Siemens), the record setting here | bone/detail/standard - detail
Recontruction Filter setting (name) - Iteration 1
Size of line pairs seen with this filter (Level =1100, Win=100)
Recontruction Filter setting (name) - Iteration 2
Size of line pairs seen with this filter (Level =1100, Win=100)
Recontruction Filter setting (name) - Iteration 3
Size of line pairs seen with this filter (Level =1100, Win=100)
Name of Filter that meets 7 lp/cm performance criteria | detail
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter
Step 7: Iterating on mA (or mAs) settings to obtain std dev of 12 +/1-1 HU
mA or mAs or Eff. mAs setting - Iteration 1 | 200
Standard deviation measured on image of ACR phantom | CTDI?? 14.88
mA or mAs or Eff. mAs setting - Iteration 2 | 250
Standard deviation measured on image of ACR phantom | 13 + detail kernel
mA or mAs or Eff. mAs setting - Iteration 3 | 325
Standard deviation measured on image of ACR phantom | 12.15 sd for detail kernel
mA or eff. mAs setting that meets Standard Deviation Performance criteria | 325
Scanner reported CTDIvol (in mGy) using this mA or eff. mAs setting | 24 mGy
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter