## Introduction
Appendix E: Checklists for Implementation of the SWS Profile
The checklist steps are arranged in approximate chronological order rather than by actor or task to make them more usable in the clinic.
The checklists are in excel to make data analysis and retrieval easier and to allow convenient conformance scoring
The final TAB contains a checklist to be used by a site doing self attestation of conformance or by an auditor to check conformance. It allows for scoring of conformance. A passing score has yet to be determined.
In the future, on-line checklist forms (such as in google docs) may be created for user convenience
To start move to the next TAB "Pre Acquisition", then to "Subjects and Data Acquisition" etc.

## Pre Acquisition
QIBA SHEAR WAVE SPEED PROFILE EXECUTION CHECKLIST V5.0 Scanner QA Pre & Post Installation
ITEM# | TASK/PARAMETER | CRITERION | RESPONSIBILITY/ACTOR | COMPLETED BY [1= MFR; 2= MFR Engineer; 3 = Site Op Manager; 4 = Site QA Manager; 5 = Operator | COMPLETED? [YES, NO, NA] | NOTES/EXPLANATIONS
3.1.2.0 | Pre-delivery
3.1.2.1 | Acoustic Output (SWS Mode) | Manufacturer Certifies acoustic output levels meets FDA recommendation | MFR
3.1.2.2 | Acoustic Transmit Focusing | Manufacturer specifies and certifies SWS measurement and Imaging | MFR
3.1.2.3 | SWS Measurement Consistency | Manufacturer confirms SWS Measurement consistency of the ultrasound scanner is within +/- 5% | MFR |  |  | US Imaging Phantom Specifications: Attenuation: 0.5 dB/cm/MHz Back Scatter: Approximately 10-4 – 10-3 cm-1Str-1 at 3 MHz or sufficient to create mean speckle brightness comparable to a human liver-mimicking phantom (± 3 dB). Speed of Sound: 1520-1540 m/sec. Stiffness: Normal Liver Equivalent & Fibrotic F3 Liver equivalent. ± 5% of the specified values. QIBA phantoms are checked and calibrated at a calibration center (Mayo) See 4.2 Assessmen Procedure: SWS Measurement Performance
3.1.2.4 | Ultrasound Imaging Performance | Meets MFR specifications as published in scanner documentation | Scanner Vendor
3.1.2.5 | SWS Imaging Performance | Identification and display meet manufacturer specifications listed in Appendix D | MFR Image processing software
3.1.2.7 | Software verification | Software version equals version specified in QIBA profile (Appendix D) | MFR
3.1.2.6 | Hardware & transducer MFR specified parameters | All Scanner, Transducer and Other Specified Devices Are Present and Operating correctly | MFR |  |  | Required Scanner components are found in the scanner specific instructions appendix D.
3.2.2.0 | Installation
3.2.2.1 | Hardware Damage | No physical damage to medical device | MFR Engineer / Clinical Staff
3.2.2.2 | Software verification | S/W version equals the version specified in QIBA Conformance Statement or in the products listed in Appendix D | MFR / QA Manager or Designee
3.2.2.3 | SWS Measurement Concordance | Confirm SWS measurements are within +/- 5% of Elastic SWS phantom specifications using standard instrument settings. | QA Manager or Designee |  |  | If the phantom specifications and independent test values are slightly different, the average of the two values will be used.
3.3.2.0 | Staff Qualification
3.3.2.1 | Operator Training | Technologist trained and approved for SWS acquisition | Technologist
3.3.2.2 | Operator Qualification | Meets performance requirements on phantoms and subjects | Technologist |  |  | Phantom testing-wCV ≤ .05 and/or case review IQR/median ≤ 0.30
3.4.2.0 | Site Quality Assurance
3.4.2.1 | US Imaging QA | U/S system QA checks and conform to quality criteria using AIUM / ACR QA guidelines | QA manager |  |  | On arrival then annually unless potential problem found during operations
3.4.2.2 | SWS Measurement Consistency & system QA Testing using SWS Phantom | Shall confirm that measurements of SWS on a QIBA elastic phantom using standard instrument settings and acquisition procedures annually, and after any software change are within ± 5% of the values of the Elastic SWS phantom specifications as determined by testing with a Verasonics system. If the system is already known to give results more than 5% different from the phantom values, the system should give values within ± 5% of the previously obtained results. (See Measurement Concordance Test Procedure Profile Section 3.2.1) US Scanner shall be capable of performing SWS measurements at reproducible instrument settings using manufacture specific standard procedures [appendix D] | QA manager / Ultrasound Scanner |  |  | Annual testing and testing after any software/hardware change.
3.4.2.3 | Operator Training & Qualification Testing | Operator trained on patient workflow & SWS acquistion and evaluated qualification criteria are met (the requirements are in 3.3. Staff Qualification) | Site manager or QA manager |  |  | Training on initial operator selection
3.4.2.4 | US Imaging / SWS Phantom Characterization & Stability Testing | Confirmation of SWS Phantom Acoustic and Mechanical Properties at Independent Test Site Using QIBA procedures after construction and if a weight change of >0.5% has occurred | Operator / QA Manager Manufacturer or Independent Phantom QA Site |  |  | Begin stability testing every 6 months. If no changes are noted after six months can change interval to every 12 months. Recalibrate as necessary. Use phantom stability QA form to record results
Note on C7: On screen dialog message appears as manufacturer indicates i.e. "Failure to estimate a shear wave velocity is indicated by an "x.xx m/s" [GE]
Note on A25: Garra, Brian: Consider moving all QA activities to a separate tab since the QA activities do not fit in termporal order with the other components.

## Subjects & Data Acquisition
QIBA SHEAR WAVE SPEED PROFILE EXECUTION CHECKLIST V5.0 Patient & Acquisition
ITEM# | TASK/PARAMETER | CRITERION | RESPONSIBILITY/ACTOR | COMPLETED BY [1= MFR; 2 = MFR Engineer; 3 = Site Op Manager; 4 = Site QA Manager; 5 = Operator | COMPLETED? [YES, NO, NA] | NOTES/EXPLANATIONS
3.5.2.0 | Subject Selection
3.5.2.1 | Clinical Indication | Assess liver stiffness for liver pathology that may lead to increased organ stiffness and increased shear wave speed (for example liver fibrosis). | Ordering Physician or Radiologist |  |  | A valid research protocol or a clinical concern supported by the literature is needed
3.5.2.2 | Approach | Confirm intercostal approach is feasible | Radiologist or Operator
3.5.2.3 | Body Wall Thickness | Confirm total body wall thickness < 4cm at intercostal location | Radiologist or Operator |  |  | allows SWS measurement to be made at a maximum depth of 6.5 cm from the skin surface and at least 2 cm below the liver capsule.
3.5.2.4 | Intercostal Space | Intercostal spacing wide enough for transducer placement | Radiologist or Operator |  |  | Evaluate intercostal space by manual palpation
3.5.2.5 | Breathing | Confirm patient can hold breath in suspended tidal for acquisition; | Radiologist or Operator |  |  | Refer to appendix D for manufacturers acquisition time
3.5.2.6 | Prior Surgery | Right lobe of liver present, no scarring or shadowing from surgery | Radiologist or Operator |  |  | Confirm subject has right lobe of liver if history of previous hepatic surgery
3.5.2.7 | Informed Consent | Informed consent obtained | Technologist or Radiologist |  |  | IC and HIPAA forms per institutional policy (sample consent form language in Appendix G)
3.6.2.0 | Subject Handling
3.6.2.1 | Patient Instructions | Fasting instructions provided | Technologist |  |  | Instruct patient on how to accomplish fasting and exceptions (e.g. oral medications, insulin); instruction sheet example Appendix G
3.6.2.2 | Fasting State | On query patient has successfully fasted 4 hours | Technologist and Radiologist |  |  | Offer to acquire data on a later date or later in the day if patient is not in a fasting state.
3.6.2.3 | Informed Consent | Informed consent confirmed | Technologist or Radiologist |  |  | IC and HIPAA forms per institutional policy. Sample consent form language in Appendix G.
3.6.2.4 | Patient Information | Provide specific information on SWE, acquisition process, number of breath holds, transducer pressure | Technologist or Radiologist |  |  | Refer to patient information-instruction sheet
3.7.2.0 | SWS Image Acquistion and Point SWS Measurement
3.7.2.1 | SWS Measurement Acquisition Device | Measurements are performed on the same equipment/manufacturer when successive measurements are made on the same patient | Ultrasound System |  |  | Obtain calibration values for successive measurements made on various machines for the same patient and forward along with test results for use during intepretation
3.7.2.2 | Patient Position | Patient positioned in supine or 30 degree left lateral decubitus | Technologist or Radiologist |  |  | Not full left lat decubitus which has been shown to produce a small but statistically significant increase in SWS
3.7.2.3 | Respiration | Patient has successfully practiced suspended tidal respiration and it is used during acquisition | Technologist |  |  | Patient to understand duration of required breath hold and ultrasound sensation
3.7.2.4 | Transducer Position | Transducer positioned at intercostal space from the upper right liver lobe | Technologist or Radiologist |  |  | Shall position the transducer at an intercostal space wide enough to accommodate the transducer and at the correct level to image/acquire from the upper right liver lobe (segments 5, 7, 8)). Shall position the transducer face long axis parallel to the intercostal space and check for correct positioning by inspection of the image for shadowing at the image edges. Shall position the transducer face in contact with the skin and parallel to the liver capsule so that the acoustic waves travel perpendicular to the capsule.
3.7.2.5 | Transducer Pressure | Maintain contact with skin surface; use lighter pressure during SWS acquisition | Technologist or Radiologist |  |  | May use more pressure to compress body wall when needed to enable ROI to be positioned in proper position in Liver
3.7.2.6 | Ultrasound Image Location Confirmation | Confirm absence of focal structures near image center, no acoustic shadowing from the ribs | Technologist or Radiologist |  |  | Keep the liver capsule parallel to the transducer face in both planes
3.7.2.7 | Measurement Region of Interest (ROI) Placement- Depth | ROI positioned 2 cm deep to liver; <6.5 cm from transducer face | Technologist or Radiologist |  |  | If a follow up study use the same depth as previous study
3.7.2.7 | Measurement Region of Interest (ROI) Placement- Position | ROI positioned near center of image, away from right/left margins and away from discrete structures | Technologist or Radiologist |  |  | Discrete structures refer to structures such as: such as liver margin, nodules, portal triads or hepatic veins
3.7.2.7 | Region of Interest (ROI) Placement- Size | ROI box diameter is 10mm or 6mm X 10mm | Technologist or Radiologist |  |  | See Appendix D. Refer to MFR specifications. If a choice is given choose an ROI equal to or larger than 10mm
3.7.2.8 | Follow-up Consistency | ROI placed consistent with baseline measurement depth for all follow up acquisitions | Technologist |  |  | Transducer position and ROI placement consistent for all measurements
3.7.2.9 | Number of Measurements | Minimum 5 qualifying measurements | Technologist or Radiologist |  |  | MFR may specify more than 5 images. Please refer to manufacturer's specific instructions (Appendix D)
3.7.2.10 | Liver Movement | Acquire when no visible liver motion | Technologist or Radiologist |  |  | Acquistion repeated if breathing or motion observed
3.7.2.11 | SWS Imaging - Color Map | Consistent between exams and patients | Technologist or Radiologist |  |  | Use institutional and/or national standards. See manufacturer guidelines
3.7.2.12 | SWS Imaging - Color Transparency | Set to visualize color changes and grayscale anatomy | Technologist or Radiologist |  |  | Refer to manufacturer guidelines - Appendix D
3.7.2.13 | SWS Imaging - Frame Averaging | Set according to preference after manufacturer recommendation | Technologist or Radiologist |  |  | After the initial setting according to manufacturer recommendations
3.7.2.14 | SWS Imaging - Frame Rate/Color Box Size | Set box as large as possible for adequate color frame rate | Technologist or Radiologist |  |  | Refer to manufacturer guidelines - Appendix D
3.7.2.15 | SWS Imaging - Point Measurement ROI location | ROI located in most homogeneous region of SWS color map (Appendix D) Reference Section 3.6.1 | Technologist or Radiologist |  |  | SEE SECTION 3.6.1. Refer to manufacturer guidelines - Appendix D
3.7.2.16 | SWS Imaging - Point Measurement ROI size | Minimum ROI size meets default 6mm X 10mm or diameter 10mm | Technologist or Radiologist |  |  | SEE SECTION 3.6.1. Refer to manufacturer guidelines - Appendix D
3.7.2.17 | SWS Imaging - Point Measurement Data Transfer | Successful Image and Data Transfer to PACS or other storage and confirmation | Technologist or Radiologist
Note on C4: possible diseases include Hepatitis C; Hepatitis B; NASH, ETOH
Note on C6: to allow positioning of ROI 2cm deep to liver capsule based on BMI
Note on C7: Wide intercostal space required. Evaluate intercostal space by manual palpation
Note on C8: Refer to appendix D for manufacturers acquisition time KA: times not listed
Note on C9: Confirm subject has right lobe of liver if positive for previous surgery
Note on C10: ICF and HIPAA forms per institutional policy
Note on C14: Query patient on fasting. Offer to acquire data on a later date or later in the day if patient is not in a fasting state.
Note on C15: ICF and HIPAA forms per institutional policy
Note on C22: Space must be wide enough to accommodate transducer & image upper Rt lobe (segments 5, 7, or 8)
Note on C24: • Clarity of visualization of normal liver parenchymal structures such as portal triads and hepatic veins on grayscale image. • No acoustic shadowing from the ribs. • Absence of grayscale artifacts.
Note on C25: if follow up study use the previous depth.
Note on C26: Away from discrete structures such as liver margin, nodules, portal triads or hepatic veins
Note on C27: Refer to Appendix D for MFR specifications. The standard ROI default size should conform to 6mm x 10mm or 10mm Diameter
Note on C28: Place ROI at constant depth for all acquisitions, and follow up acquisitions
Note on C29: MFR may specify more than 10 images. Please refer to vendor specific instructions (Appendix D)
Note on C30: Shall acquire only when there is no visible liver motion

## Quality Assurance & Processing
QIBA SHEAR WAVE SPEED PROFILE EXECUTION CHECKLIST V5.0 Processing & Analysis
ITEM# | TASK/PARAMETER | CRITERION | RESPONSIBILITY/ACTOR | COMPLETED BY [1= MFR; 2 = MFR Engineer; 3 = Site Op Manager; 4 = Site QA Manager; 5 = Operator | COMPLETED? [YES, NO, NA] | NOTES/EXPLANATIONS
3.8.2.0 | Image Related QA
3.8.2.1 | Suboptimal SWS Acquistion handling | SWS estimates excluded if associated are suboptimal due to motion or other factors | Technologst or Radiologist |  |  | See Profile Section 3.6 for rules of acquisition
3.8.2.2 | User training on image display | User has been trained on image interpretation and optimal ROI placement | MFR
Note on C4: Exclude SWS numbers that were acquired suboptimally either by observation during acquisition, inspection of saved images,or pre/post liver images not matching. • liver movement during acquisition, • patient talking during acquisition, • transducer slippage during acquisition and • inadvertent shift of ROI to a deeper or shallower depth, • placement of the ROI near to a vessel or other discrete structure

## Profile Conformance
COMPLIANCE CHECKLIST (enter: 1 if self attestation, 2 if independent review):
ITEM# | TASK/PARAMETER | SCORE Full Compliance = 3 pts; Non-compliant one task = 2 pts; Non-compliant 2-3 tasks = 1 pts; Non-compliant 3+ tasks = 0 pts | NOTES/EXPLANATIONS
3.1.2.0 | Pre-delivery: Acoustic Output (SWS Mode); Acoustic Transmit Focusing; SWS Measurement Consistency; Ultrasound Imaging Performance; SWS Imaging Performance; Software verification; Hardware & transducer MFR specified parameters
3.2.2.0 | Installation: Hardware Damage Check; Software verification; SWS Measurement Concordance
3.3.2.0 | Staff Qualification: Operator Training; Operator Qualification
3.4.2.0 | US Imaging QA Check: SWS Measurement Consistency & SWS system QA Testing using SWS Phantom; Operator Training & Qualification Testing; US Imaging / SWS Phantom Characterization & Stability Testing
3.5.2.0 | Subject Selection: Clinical Indication; Approach; Body Wall Thickness; Intercostal Space; Breathing; Prior Surgery; Informed Consent
3.6.2.0 | Subject Handling: Patient Instructions; Fasting State; Informed Consent; Patient Information
3.7.2.0 | SWS Image Acquistion and Point SWS Measurement: SWS Measurement Acquisition Device; Patient Position; Respiration ; Transducer Position; Transducer Pressure; Ultrasound Image Location Confirmation; Measurement Region of Interest (ROI) Placement-Depth; Measurement Region of Interest (ROI) Placement-Position; Region of Interest (ROI) Placement-Size; Follow-up Consistency; Number of Measurements; Liver Movement; SWS Imaging-Color Map; SWS Imaging-Color Transparency; SWS Imaging-Frame Averaging; SWS Imaging-Frame Rate/Color Box Size; SWS Imaging-Point Measurement ROI location; SWS Imaging-Point Measurement ROI size; SWS Imaging-Point Measurement Data Transfer
3.8.2.0 | Image Related QA: Suboptimal SWS Acquistion handling; User training on image display
4.1.0 | Assessment: Imaging Performance - Site has assessed imaging performance according to procedure 4.1.2 (scoring: yes = 3, other imaging assessment = 2, no assessment = 1)
Corrective Action Taken if Performance Issue Identified (scoring: yes successful = 3, yes partially successful = 2, yes unsuccessful = 1, no corrective action)
4.2.0 | Assessment: SWS Performance - Site has assessed ACTOR performance according to ACTOR assessment procedure 4.2.1 (scoring: yes, periodic = 3, yes, once = 2, other assessment performed = 1, no assessment = 0)
Corrective Action Taken if Performance Issue Identified (scoring: yes- successful, assessed actors as in 4.2.3 = 3, yes - successful, after other assessment = 2, yes - partially successful = 1, no corrective action = 0)
TOTAL SCORE: