## Introduction
Appendix E: Checklists for Implementation of the SWS Profile
The checklist steps are arranged in approximate chronological order rather than by actor or task to make them more usable in the clinic.
The checklists are in excel to make data analysis and retrieval easier and to allow convenient conformance scoring
The final TAB contains a checklist to be used by a site doing self attestation of conformance or by an auditor to check conformance. It allows for scoring of conformance. A passing score has yet to be determined.
In the future, on-line checklist forms (such as in google docs) may be created for user convenience
To start move to the next TAB "Pre Acquisition", then to "Subjects and Data Acquisition" etc.

## Pre Acquisition & QA
QIBA SHEAR WAVE SPEED PROFILE EXECUTION CHECKLIST V5.0 Scanner QA Pre & Post Installation
ITEM# | PARAMETER | REQUIREMENT | RESPONSIBILITY/ACTOR | COMPLETED BY [ 1= Radiologist; 2 = Technologist / Sonographer; 3 = MFR; 4 = QA Manager] | Rule | CONFORMS? [YES, NO, N/A] | NOTES/EXPLANATIONS
3 | Site Conformance Check
3.0.2.1 | Ultrasound Scanner | Shall confirm requirements for performance and installation are met, and US scanners conform to Profile | QA Manager |  | Required |  | Typically fulfilled by conformance letter/statement from manufacturer. Vendor/software rev verificaiton may also be verified in Appendix.
3.0.2.1 | Technologist / Sonographer / Radiologist | Shall confirm that each potential operator (technologist / sonographer and Radiologist) conforms to this Profile with respect to training, documented acquistion performance, and proper SWS acquistions. If a site has ultrasound systems from multiple manufacturers, conformance shall be confirmed for each model / scanner type at the site. | QA Manager |  | Required |  | QA Manager verifying all operators meet qualifications for performing SWS acquisitions. See 3.1 for Operator Training. Additional specific requirements may be site-specific though should be documented. Examples could include use of phantom (goal wCV ≤ .05), documentation of formal training, central review of practice cases, and formal approval of competency
3.1 | Staff Qualification
3.1.2.1 | Operator Training | Shall be trained and approved for SWS acquisition, meeting performance requirements on subjects based on site-specific protocol. Patient case review with IQR/median ≤ 0.30 if in kPa, or < 0.15 if in m/s. | Technologist / Sonographer or Radiologist |  | Required |  | Additional specific requirements may be site-specific though should be specifically defined and documented. Examples could include use of phantom (goal wCV ≤ .05), documentation of formal training, central review of practice cases, and formal approval of competency. See section 4.2.1.1.C for detailed description of in-vivo assessment testing.
3.3 | Pre-delivery
3.3.2.1 | System Conformance | Shall certify system specifications conform to the Profile, including acoustic output, focusing, and measurement consistency. | MFR |  | Required |  | Manufacturer specification and certification that scanner and software conforms to the Profile. Typically fulfilled by conformance letter/statement from manufacturer. Vendor/software rev verificaiton may also be verified in Appendix.
3.3.2.2 | Hardware & transducer MFR specified parameters | Shall ensure the equipment intended for use is listed in Appendix D as a compliant combination of System, Software Revision and Transducer. | MFR |  | Required |  | Required Scanner components are present and found in Appendix D; scanner specific instructions for performing SWS measurements are included
3.3.2.3 | SWS Measurement Consistency | Shall confirm SWS Measurement Consistency of the Ultrasound Scanner is within +/- 5% | MFR |  | Required |  | See 4.2 Assessment Procedure: SWS Measurement Consistency
3.4 | Installation (including software and hardware upgrades)
3.4.2.1 | Software verification | Shall confirm S/W version equals the version specified in products QIBA Conformance Statement or one listed in Appendix D | QA Manager |  | Required |  | Required software specifications found in Appendix D. For newest software versions that may not be listed in Appendix D, a conformance letter/statement from manufacturer that software is conformant to the Profile is sufficient.
3.4.2.2 | SWS Measurement Concordance | Shall confirm SWS measurements are within +/- 5% in a calibrated phantom | QA Manager |  | Recommended |  | See 3.5.2 Assessment Procedure: SWS Measurement Consistency. If site has resources, SWS Measurement Consistency Testing is expected. May not be feasible for all site. Refer to section 4.1.1 of the Profile for phantom specifications. See section 4.3 for detailed description of concordance testing.
3.5 | Periodic Quality Assurance (QA)
3.5.2.1 | US Imaging QA | Shall perform standard ultrasound system QA checks and conform to quality criteria specified in ACR, AIUM, or other accreditation guidelines | QA manager |  | Required |  | On arrival then annually, unless potential problem found during operations, and after major software/hardware change, repair, or major upgrade.
3.5.2.2 | SWS Measurement Consistency & System QA Testing | Shall confirm that measurements of SWS on a QIBA elastic phantom using standard instrument settings and acquisition procedures annually, and after any software change, are within ± 5% of the values of the Elastic SWS phantom specifications as determined by testing with a Verasonics system. If the system is already known to give results more than 5% different from the phantom values, the system should give values within ± 5% of the previously obtained results. | QA manager |  | Recommended |  | Annual testing, independent of testing after any software/hardware change, repair, or major upgrade. If site has resources, SWS Measurement Consistency Testing is expected. May not be feasible for all site. See Measurement Concordance Test Procedure in section 3.4.2.2 above. See section 4.2.1.1.B for detailed description of assessment testing.
3.5.2.3 | Operator Training & Qualification Testing | Shall confirm that each operator continues to meet qualification criteria (as listed in 3.1. Staff Qualification) | QA manager |  | Required |  | Phantom testing: wCV ≤ .05, and/or patient case review with IQR/median ≤ 0.30 if in kPa, or < 0.15 if in m/s. Additional specific requirements may be site-specific though should be documented. Examples could include documentation of formal QC process, central review of randomly selected cases for quality, and formal re-approval of competency
3.5.2.4 | US Imaging / SWS Phantom Characterization & Stability Testing | Shall confirm SWS Phantom Acoustic and Mechanical Properties at Independent Test Site using QIBA procedures after construction and if a weight change of >0.5% has occurred | QA Manager |  | Recommended |  | If site has resources, SWS Phantom Stability Testing is expected. May not be feasible for all site. Refer to section 4.1.1 of the Profile for phantom specifications.
Note on A19: Garra, Brian: Consider moving all QA activities to a separate tab since the QA activities do not fit in termporal order with the other components.

## Subjects & Data Acquisition
QIBA SHEAR WAVE SPEED PROFILE EXECUTION CHECKLIST V5.0 Patient & Acquisition
ITEM# | PARAMETER | REQUIREMENT | RESPONSIBILITY/ACTOR | COMPLETED BY [ 1= Radiologist; 2 = Technologist / Sonographer] | Rule | COMPLETED? [YES, NO, N/A] | NOTES/EXPLANATIONS
3.7 | Subject Selection
3.7.2.1 | Prior Surgery | Shall confirm the presence of the right liver lobe & the absence of surgical/other scars that could cause shadowing. Lack of optimal acoustic window should be communicated to interpreting radiologist for case review. | Radiologist or Technologist / Sonographer |  | Recommended |  | Confirm absence of relevant intra-abdominal surgery, or major abdominal/chest wall abnormalities or dressing, that would precude imaging of the right lobe (ie. confirm intercostal imaging of the right lobe of liver is possible).
3.7.2.2 | Breathing | Shall confirm the ability of the patient to follow breath hold instructions. Inability of patient to follow instructions should be communicated to interpreting radiologist for case review. | Radiologist or Technologist / Sonographer |  | Recommended |  | Refer to appendix D for manufacturers acquisition time needed for each breath-hold
3.8 | Subject Handling
3.8.2.1 | Fasting State | Shall confirm patient has fasted for at least 4 hours. Non-fasting status must be communicated to the interpreting radiologist. | Radiologist or Technologist / Sonographer |  | Required |  | Opererator should offer to acquire data on a later date or later in the day if patient is not in a fasting state.
3.8.2.2 | SWS Acquisition System | Shall perform measurements on the same equipment/Manufacturer when successive measurements are made on the same patient | Radiologist or Technologist / Sonographer |  | Required |  | To ensure longitudinal differences in measurement are due to patient-related factors and not system differences
3.9 | SWS Image Acquisition (SWEI) and Point SWS Measurement
3.9.2.1 | Patient Position | Shall ensure patient is comfortable, positioned in supine or 30 degree left lateral decubitus | Radiologist or Technologist / Sonographer |  | Required |  | Patient should NOT be full left lat decubitus which has been shown to produce a small but statistically significant increase in SWS
3.9.2.2 | Transducer Position | Shall position transducer at intercostal space wide enough to accommodate transducer and at level to image mid to upper right liver lobe. Shall position transducer face long axis parallel to the intercostal space and check for correct positioning by inspecting the image for shadowing at the image edges. Shall position the transducer parallel to liver capsule so acoustic waves travel perpendicular to capsule. | Radiologist or Technologist / Sonographer |  | Required |  | Transducer positioned at an intercostal space with optimized grayscale image to minimize artifacts and ensure maximum transmission of push pulse into the upper right liver lobe
3.9.2.3 | Ultrasound Image | Shall optimize B-mode image. Shall confirm absence of focal structures near image center; no acoustic shadowing from the lung or ribs | Radiologist or Technologist / Sonographer |  | Required |  | Grayscale image optimized to minimize artifacts, and help ensure measurement area is free of discrete structures such as: liver margin, nodules, portal triads or hepatic veins.
3.9.2.4 | Measurement Region of Interest (ROI) Use and Placement | Shall use the standard ROI size specified by the ultrasound vendor as the default for their system; ROI center positioned 2 cm deep to liver; <7.0 cm from transducer face; near center of image sector, away from liver margins and from discrete structures | Radiologist or Technologist / Sonographer |  | Required |  | Discrete structures refers : liver margin, nodules, portal triads or hepatic veins. For follow up study use the same depth as previous study.
3.9.2.5 | Respiration | Shall confirm patient has successfully suspended tidal respiration during acquisition. Inability of patient to follow instructions shall be communicated to interpreting radiologist. | Radiologist or Technologist / Sonographer |  | Required |  | Patient should suspend respiration at end expiration of mid respiratory cycle. Measurements should not be taken after larger inhalation or during Valsalva
3.9.2.6 | Liver Movement | Shall acquire only when minimized visible liver motion. Presence of liver motion during measurements should be communicated to interpreting radiologist. | Radiologist or Technologist / Sonographer |  | Required |  | No other liver movement is observed during acquistion process. Acquistion should be repeated if significant liver or transducer motion observed.
3.9.2.7 | Measurement Location Consistency | Shall make follow-up acquisitions and ROI placements consistent with the baseline measurement in terms of the Transducer Position and Measurement Region of Interest (ROI) Placement. ROI placement and measurement shall be at consistent depth and location from baseline across all subsequent acquisitions | Radiologist or Technologist / Sonographer |  | Required |  | All SWS measurements should be obtained from the same area of liver parenchyma
3.9.2.8 | Number of Measurements | Shall take the number of measurements specified by the Manufacturer (see Appendix D) or if not specified, 10 measurements. | Radiologist or Technologist / Sonographer |  | Required |  | Please refer to manufacturer's specific instructions (Appendix D). 10 measurements is the most common used and recommended number.
3.9.2.9 | SWS Imaging - Quality Assessment | Shall consider successful color map acquisition as at least 50% color fill. Shall measure in regions of high consistency, quality, and/or confidence, following manufacturer specifications | Radiologist or Technologist / Sonographer |  | Required |  | Refer to manufacturer guidelines - Appendix D
3.9.2.10 | SWS Imaging - Measurement ROI Placement | Shall place the measurement ROI location in most consistent, reproducible region of SWS color map or other images related to SWS variability as specified by MANUFACTURER (Appendix D). | Radiologist or Technologist / Sonographer |  | Required |  | Refer to manufacturer guidelines - Appendix D

## Profile Conformance-by section
| QIBA SHEAR WAVE SPEED PROFILE CHECKLIST V5.0 Conformance-by Section |  |  | Instructions: For each section, first fill in the answer (y or n) for each task. Then fill in the conformance score for the section in the next column. The overall site score as a percent of the maximum possible will then be calculated automatically.
SECTION CONFORMANCE CHECKLIST (enter: 1 if self attestation, 2 if independent review):
ITEM# | PARAMETER | Conforms Y/N | SECTION CONFORMANCE SCORE Full Conformance = 3 pts; Non-conformant one task = 2 pts; Non-conformant 2-3 tasks = 1 pts; Non-conformant 3+ OR ALL tasks = 0 pts | NOTES/EXPLANATIONS
3.1 | Staff Qualification
3.1.2.1 | Operator Training
3.1.2.2 | Operator Qualification
3.3 | Pre-delivery
3.3.2.1 | Acoustic Output (SWS Mode)
3.3.2.2 | Acoustic Transmit Focusing
3.3.2.3 | SWS Measurement Consistency
3.3.2.4 | Ultrasound Imaging Performance
3.3.2.5 | SWS Imaging Performance
3.3.2.6 | Software verification
3.3.2.7 | Hardware & transducer MFR specified parameters
3.4 | Installation
3.4.2.1 | Hardware Damage Check
3.4.2.2 | Software verification
3.4.2.3 | SWS Measurement Concordance
3.5 | Periodic Quality Assurance (QA)
3.5.2.1 | US Imaging QA
3.5.2.2 | SWS Measurement Consistency & System QA Testing Using SWS Phantom
3.5.2.3 | Operator training and qualification testing
3.5.2.4 | US Imaging and SWS Phantom Characterization and Stability Testing
3.7 | Subject Selection
3.7.2.1 | Clinical Indication
3.7.2.2 | Approach
3.7.2.3 | Body Wall Thickness
3.7.2.4 | Intercostal Space
3.7.2.5 | Breathing
3.7.2.6 | Prior Surgery
3.7.2.7 | Informed Consent
3.8 | Subject Handling
3.8.2.1 | Patient Instructions
3.8.2.2 | Fasting State
3.8.2.3 | Informed Consent
3.8.2.4 | Patient Information
3.9 | SWS Image Acquistion and Point SWS Measurement
3.9.2.1 | SWS Measurement Acquisition Device
3.9.2.2 | Patient Position
3.9.2.3 | Respiration
3.9.2.4 | Transducer Position
3.9.2.5 | Transducer Pressure
3.9.2.6 | Ultrasound Image Location Confirmation
3.9.2.7 | Measurement Region of Interest (ROI) Placement
3.9.2.8 | Follow-up Consistency
3.9.2.9 | Number of Measurements
3.9.2.10 | Liver Movement
3.9.2.11 | SWS Imaging-Color Map
3.9.2.12 | SWS Imaging-Color Transparency
3.9.2.13 | SWS Imaging-Frame Averaging
3.9.2.14 | SWS Imaging-Frame Rate/Color Box Size
3.9.2.15 | SWS Imaging-Point Measurement ROI location
3.9.2.16 | SWS Imaging-Point Measurement ROI size
3.9.2.17 | SWS Imaging-Point Measurement Data Transfer
3.11 | Image QA
3.11.2.1 | Suboptimal SWS Acquistion handling
3.11.2.2 | User training on image display
4.1 | Assessment Procedure: Imaging Performance
4.1.1 | Assessment: - Site has assessed imaging performance according to procedure 4.1.2 (scoring: yes = 3, other imaging assessment = 2, no assessment = 1)
4.1.2 | Corrective Action Taken if Performance Issue Identified (scoring: yes successful = 3, yes partially successful = 2, yes unsuccessful = 1, no corrective action)
4.1.3 | Assessment: SWS Performance - Site has assessed ACTOR performance according to ACTOR assessment procedure 4.1.2 (scoring: yes, periodic = 3, yes, once = 2, other assessment performed = 1, no assessment = 0)
4.1.4 | Corrective Action Taken if Performance Issue Identified (scoring: yes- successful, assessed actors as in 4.2.3 = 3, yes - successful, after other assessment = 2, yes - partially successful = 1, no corrective action = 0)
| CONFORMANCE SCORE (% OF MAX): |  | 0
TO BE UPDATED AFTER FINAL VOTE; BEFORE DISTRIBUTION

## Profile Conformance by Actor
| QIBA SHEAR WAVE SPEED PROFILE CHECKLIST V5.0 Conformance by Actor |  |  |  | Instructions: For each Actor, first fill in the answer (y or n) for each task. Then fill in the conformance score for the actor in the next column. The overall site score as a percent of the maximum possible will then be calculated automatically.
ACTOR CONFORMANCE CHECKLIST (enter: 1 if self attestation, 2 if independent review):
Item # | ACTOR | ACTIVITY / PARAMETER | Conforms Y/N | ACTOR CONFORMANCE SCORE Full Conformance = 3 pts; Non-conformant one task = 2 pts; Non-conformant 2-3 tasks = 1 pt; Non-conformant 3+ OR ALL tasks= 0 pts | NOTES / EXPLANATIONS
3.5 | Ultrasound Scanner | Periodic Quality Assurance
3.5.2.1 | Scanner | SWS Measurement Consistency & System QA Testing Using SWS Phantom
3.3 | Manufacturer | Pre-Delivery
3.3.2.1 | Mfr | Acoustic Output (SWS Mode)
3.3.2.2 | Mfr | Acoustic Transmit Focusing
3.3.2.3 | Mfr | SWS Measurement Consistency
3.3.2.4 | Mfr | US Imaging Performance
3.3.2.5 | Mfr | SWS Imaging Performance
3.3.2.6 | Mfr | Software verification
3.3.2.7 | Mfr | Hardware and transducer
3.4 |  | Installation
3.4.2.1 | Mfr | Hardware Damage
3.1 | Radiologist | Site Qualitification
3.1.2.1 | Rad | Operator Training
3.1.2.2 | Rad | Operator Qualification
3.7 |  | Subject Selection
3.7.2.1 | Rad | Clinical Indication
3.7.2.2 | Rad | Approach
3.7.2.3 | Rad | Body Wall Thickness
3.7.2.4 | Rad | Intercostal space
3.7.2.5 | Rad | Breathing
3.7.2.6 | Rad | Prior Surgery
3.8 | Rad | Subject Handling
3.8.2.1 | Rad | Patient Instructions
3.8.2.2 | Rad | Fasting State
3.8.2.3 | Rad | Informed Consent
3.8.2.4 | Rad | Patient Information
3.9 |  | Image Data Acquisition
3.9.2.1 | Rad | SWS Acquisition
3.9.2.2 | Rad | Patient Position
3.9.2.3 | Rad | Respiration
3.9.2.4 | Rad | Transducer Position
3.9.2.5 | Rad | Transducer Pressure
3.9.2.6 | Rad | Ultrasound Image – location confirmation
3.9.2.7 | Rad | Measurement Region of Interest (ROI) Placement
3.9.2.8 | Rad | Follow-up Consistency
3.9.2.9 | Rad | Number of Measurements
3.9.2.10 | Rad | Liver Movement
3.9.2.11 | Rad | SWS Imaging Color Map
3.9.2.12 | Rad | SWS Imaging Color Transparency
3.9.2.13 | Rad | SWS Imaging Frame Averaging
3.9.2.14 | Rad | SWS Imaging Frame Rate/Color Box Size
3.9.2.15 | Rad | SWS Imaging Point Measurement ROI location
3.9.2.16 | Rad | SWS Imaging Point Measurement ROI size
3.9.2.17 | Rad | SWS Imaging Point Measurement Data Transfer
3.11 |  | Image QA
3.11.2.1 | Rad | Suboptimal SWS Acquisition
3.1 | Technologist / Sonographer | Site Qualification
3.1.2.1 | Sono Tech | Operator Training
3.1.2.2 | Sono Tech | Operator Qualification
3.7 |  | Subject Selection
3.7.2.2 | Sono Tech | Approach
3.7.2.3 | Sono Tech | Body Wall Thickness
3.7.2.4 | Sono Tech | Intercostal space
3.7.2.5 | Sono Tech | Breathing
3.7.2.6 | Sono Tech | Prior Surgery
3.8 |  | Subject Handling
3.8.2.1 | Sono Tech | Patient Instructions
3.8.2.2 | Sono Tech | Fasting State
3.8.2.3 | Sono Tech | Informed Consent
3.8.2.4 | Sono Tech | Patient Information
3.9 |  | Image Data Acquisition
3.9.2.1 | Sono Tech | SWS Acquisition
3.9.2.2 | Sono Tech | Patient Position
3.9.2.3 | Sono Tech | Respiration
3.9.2.4 | Sono Tech | Transducer Position
3.9.2.5 | Sono Tech | Transducer Pressure
3.9.2.6 | Sono Tech | Ultrasound Image – location confirmation
3.9.2.7 | Sono Tech | Measurement Region of Interest (ROI) Placement
3.9.2.8 | Sono Tech | Follow-up Consistency
3.9.2.9 | Sono Tech | Number of Measurements
3.9.2.10 | Sono Tech | Liver Movement
3.9.2.11 | Sono Tech | SWS Imaging Color Map
3.9.2.12 | Sono Tech | SWS Imaging Color Transparency
3.9.2.13 | Sono Tech | SWS Imaging Frame Averaging
3.9.2.14 | Sono Tech | SWS Imaging Frame Rate/Color Box Size
3.9.2.15 | Sono Tech | SWS Imaging Point Measurement ROI location
3.9.2.16 | Sono Tech | SWS Imaging Point Measurement ROI size
3.9.2.17 | Sono Tech | SWS Imaging Point Measurement Data Transfer
3.11 |  | Image QA
3.11.2.1 | Sono Tech | Suboptimal SWS Acquisition
3 | QA Manager | Site Conformance
| QA Mgr | Ultrasound Scanner Conformance
| QA Mgr | Manufacturer Conformance
| QA Mgr | Technologist/Sonographer Conformance
| QA Mgr | Radiologists Conformance
3.4 |  | Installation
| QA Mgr | Software Verification
| QA Mgr | SWS Measurement Concordance
3.5 |  | Periodic Quality Assurance (QA)
| QA Mgr | US Imaging QA
| QA Mgr | SWS Measurement Consistency & System QA Testing Using SWS Phantom
| QA Mgr | Operator training and qualtification testing
| QA Mgr | US Imaging & SWS Phantom Characterization & Stability Testing
ALL | SITE | SITE CONFORMANCE SCORE (% OF MAX) |  | 0
TO BE UPDATED AFTER FINAL VOTE; BEFORE DISTRIBUTION