# IV. Compliance

Relation of this Profile to Expectations for QIBA Profile Compliance

Definitions

Qualified: The imaging site is formally approved by an appropriate body (i.e. ACRIN, NCRI, an imaging laboratory or CRO) for a specific clinical research study.

Accredited: Approval by an independent body or group for broad clinical usage (requires ongoing QA/QC)

Compliant: The imaging site and equipment meet all the requirements described herein, which are necessary to meet the QIBA Profile claim.

The requirements included here are intended to establish a baseline level of capabilities. Providing higher levels of performance or advanced capabilities is both allowed and encouraged. Furthermore the QIBA Profile is not intended to limit equipment suppliers in any way with respect to how they meet these requirements. Institutions meeting the stated criteria are considered to be QIBA Compliant.

At this stage of development QIBA profile compliance will be assessed by the respective actors as outlined in the sections below. There is currently no formal QIBA or other third party compliance testing and certification. This might change over time.

Proper utilization of the DWI-MRI Profile requires that each site adhere to standardized DWI-MRI Profile parameters.  This is an essential component of clinical trials in which imaging plays a central role in the research endpoints.  This is of particular importance in multicenter trials where equipment, personnel, and imaging acquisition protocols can vary significantly.  Thus, the use of standardized imaging guidelines helps control the inter- and intra-variability inherent in multicenter imaging trials.

The purpose of any imaging qualification process is to help ensure the trial imaging is of high quality and performed per the trial-standardized acquisition protocol.  To participate in the utilization of this profile, each site must qualify by first scanning the DWI-MRI phantoms based on the scanning protocols in Appendix.  Phantom scanning provides an opportunity to evaluate compliance with sample imaging acquisition protocols prior to participant recruit and actual trial-specific protocols.

Suboptimal image quality and/or imaging not performed per the trial-standardized protocol can result in exclusion of the imaging exams and/or the entire case from analysis.  Therefore, routine QC and adherence to the protocol are of great importance, and sites will be asked to re-scan using the phantoms should there be any substantive changes in hardware or software to the scanner during the conduct of the trial.

4.1  Performance Assessment:  By Actor

4.1.1.  Image Acquisition Site

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large patient population under consideration. For DCE-MRI use as quantitative imaging biomarker it is essential to put some effort into an imaging capability assessment prior to final site selection for a specific trial. For imaging it is important to consider the availability of:

- appropriate imaging equipment and quality control processes,

- experienced MR technologists for the imaging procedure, and

- processes that assure imaging protocol compliant image generation at the correct point in time.

- Appropriately trained image analysts, with oversight by a Radiologist

- Medical Physics support to ensure appropriate scanner and equipment calibration

- Processes that assure imaging QIBA Profile-compliant image generation in appropriate time window

A QA/QC program for MRI scanners and ancillary devices must be in place to achieve the goals of the clinical trial.  The minimum requirements are specified above.  This program shall include (a) elements to verify that imaging facilities are performing imaging studies correctly and (b) elements to verify that additional MRI phantom testing that address issues relating to image quality (which may include issues relating to spatial resolution, signal to noise ratio, contrast to noise ratio, B0 and B1 uniformity) and constancy.  There is agreement that some performance testing (e.g. constancy phantom) adds value: however, acceptable performance levels, frequency of performance, triggers for action and mitigation strategies need further definition before these can be required.  This phantom testing may be done in addition to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.

| Parameter | Entity/Actor | Specification |
| --- | --- | --- |
| MRI scanner | Acquisition Facility | This profile shall only address 1.5T and 3.0 T MRI scanners |
| MRI Scanner Calibration | Vendor/Engineer | Shall perform routine QA and preventative maintenance |
| MRI Scanner Calibration | Technologist | Shall perform routine (e.g. phantom measurements to ensure that SNR, CNR, RF are stable and uniform) |
| Training | Technologist | Should be MR-certified according to local regulations and should have prior experience in conducting dynamic contrast enhance imaging. Should be experienced in clinical study related imaging and should be familiar with good clinical practices (GCP). |
| Training | Lead Technologist | Should be MR-certified according to local regulations and should have prior experience in conducting dynamic contrast enhance imaging. Should be experienced in clinical study related imaging and should be familiar with good clinical practices (GCP). Should also have backup personnel need that should fulfill similar requirements as lead research technologists. Contact information for lead technologist and research technologists should be kept up to date and readily available in case questions may arise. |
| DWI-MRI Phantom | Technologist | A phantom that contains a range of ADC values (generally obtained via different concentrations of polymers or other non-toxic materials) should be scanned using the dynamic protocol on each scanner that will be used for the study. This scanning should be performed on at least a quarterly basis and results documented in a data transmittal form that is properly signed by the technologist performing the scanning. Phantom testing will also be done and results documented after any hardware or software upgrade. |

## 4.1.2   Acquisition Scanner

Distinct from the performance specifications and frequency of testing described in Section 4.1.1, which apply to QC of the Acquisition Device at the imaging facility, this Section defines performance specifications of the Acquisition Device to be met upon leaving the manufacturing facility.  In order to be compliant with this Profile, the Acquisition Device should be held to the same standard whether a mobile utility or a fixed installation; a mobile scanner may require additional calibration to achieve this performance.

The MRI scanner should use DICOM attributes to follow version numbers for the following software components : 1 Acquisition, 2 Reconstruction, 3 Post-processing, 4 Display/ROI analysis, 5 Dynamic Analysis.  Performance requirements regarding software version identification, documentation and tracking across time are described below.

The DICOM format used by the MRI scanner should meet the Conformance Statement written by manufacturer of the MRI system.  MRI data shall be encoded in the DICOM MRI SOP Class, and in SI units with additional parameters in public DICOM fields to calculate ADC. DICOM data shall be transferred using the DICOM Part 8 network protocol or as offline DICOM 10 files for media storage including CD’s and DVDs.  They shall be transferred without any form of lossy compression.

The meta-information is the information that is separate, or in addition to, the image values (in units of SI, e.g. b-values in s/mm2).  The meta-information may also include other information beyond that needed for calculation of ADC, (i.e. the type or sequencing of therapy, concomitant chemotherapies, gender, organ, cancer type). The actual mechanism of capturing the information is not specified in this Profile.  The intent here is to list what information should be captured rather than the mechanism of capture.  The mechanism can range from paper notes, to scanned forms or electronic data records, to direct entry from the measurement equipment into pre-specified DICOM fields.  Ideally all of the specified meta-data will be captured by direct electronic entry to DICOM fields, after suitable modification of the DICOM format for MRI.

The concept endorsed here is that the needed meta-data is identified.  Through revisions of this Profile, the DICOM standard, and technology the meta-data is inserted into the analysis stream in a more direct manner and technology and accepted standards evolve.

Practically speaking sites are encouraged to perform longitudinal treatment trials on one instrument.

| Parameter | Entity/Actor | Specification |
| --- | --- | --- |
| MRI scanner | Acquisition Facility | This profile shall only address 1.5T and 3.0 TMRI scanners with 55-70 cm bores |
| MRI Scanner Calibration | Vendor/Engineer | Shall maintain standard requirements for routine clinical imaging and for MR facility accreditation purposes per ACR guidelines (http://www.acr.org/Quality-Safety/Standards-Guidelines/Practice-Guidelines-by-Modality/MRI). It is recognized that not all scanners that meet ACR accreditation guidelines will be acceptable for quantitative DWI-MRI studies, and therefore phantom testing such as described in Section 4.1.5 below is advised |
| MRI Scanner Calibration | Lead Technologist | Will track scanner software version across time. Will identify and qualify a second scanner at the site, if available. If this is done prior to the study start there will be no difficulties later on in case the first scanner is temporarily unavailable. |
| MRI Scanner Calibration | Technologist | Shall perform routine (daily, weekly, monthly QA and preventative maintenance as per ACR guidelines (ref or link provided) for routine clinical imaging and for MR facility accreditation purposes. Data will be kept in data log and data transmittal form that will be readily available for inspection with oversight by the local site PI. Technologist should alert PI of any irregularities or inconsistencies in the regular QA results, as these may indicate variation in scanner performance and necessitate need for more rigorous performance evaluation with the DCE-MRI phantom |

## 4.1.3   Software Analysis

For multi-institutional trials a central reading site is assumed.

| Parameter | Entity/Actor | Specification |
| --- | --- | --- |
| Image analysis | PI/Radiologist/Image analyst/Core lab | Shall ensure that image analysis is performed with a DWI-MRI tool that complies with section 4.2.4. Software should be re-evaluated after every software and hardware upgrade. |
| Image interpretation | PI/Radiologist/Core lab | Shall ensure that image interpretation is performed with a DWI-MRI tool that complies with 4.2.4. |
| Image analysis (DRO) | PI/Radiologist/Image Analyst/Core lab | Work is ongoing for strict software qualification process that will likely include a digital reference object (DRO). Central analysis sites should include this in their quotidian rubric. Values from analysis utilizing the DRO should be recorded. |

4.1.5   Imaging qualification process:

Prior to site qualification, a DWI phantom test scan must be run to assess MRI system performance.  Sites will be provided a standardized multi-compartment ADC phantom having inserts with ADCs spanning the ranges typically obtained across the major organs.  Sites are to retain the DWI phantom, until the core lab has approved the submitted imaging.

| Parameter | Entity/Actor | Specification |
| --- | --- | --- |
| Checklist | Radiologist | Site and scanner qualification can be obtained using a simple questionnaire as a pre-qualification step. |
| Site Visit | PI | If appropriate equipment and personnel are available, a site visit is recommended. During the site visit, study related imaging protocols are discussed and, ideally, all scan parameters are entered at the MR scanner. |
| Phantom scanning | Technologist | After initial qualification, phantoms should be rescanned at regular intervals (e.g. every 6-12 months) to ensure continued scanner performance within acceptable levels. Repeat phantom scanning should also be performed after any major hardware or software upgrade, or as necessary due to deviations of scanner performance from that expected during more frequent routine QC checks. |
| Phantom scanning interval | Technologist | The phantom scans should be repeated on a regular interval (e.g 3 months) during the course of the study. |
| Image Quality Inspection | Radiologist/Core Lab | Ongoing image quality inspection on a per scan basis is essential to high quality studies |
| Hardware Software upgrades/changes | Lead Technologist | Any changes to scanner equipment, including major hardware changes or any software version change, need to be documented and will result in the need for imaging qualification renewal |

4.1.6   Site Analysis qualification:

Significant variations in any of the parameters tested (e.g. ADC, signal variance across time) during the course of an ongoing longitudinal study can affect the resulting imaging biomarker determinations, in the case of this specific claim ADC, and such changes can readily occur if there are major changes in the scanner hardware or software, e.g., an update to the pulse sequence used for the DWI-MRI or to the gradient subsystem.

| Parameter | Entity/Actor | Specification |
| --- | --- | --- |
| Phantom images | Radiologist/Core Lab technician | The qualification exam will be evaluated for image quality (e.g. artifacts, distortion, and SNR) and compliance with the standardized DWI-MRI protocol |
| ADC variance | Radiologist/Core Lab Technician | Data analysis procedures to be used in the DWI-MRI application should be used to analyze the ADC mapping data and results compared to the known ADC values of the various compartments. As uncertainty in the measurement of ADC is an important contributor to concentration measurement bias, the measured values should compare within 5 % of the known values over an ADC range of approximately 0.12-1.1 x 10-9 mm2/s. Variance measures obtained from software analysis utilizing the DRO should also adhere the same variance tolerance measures. |
| Temporal resolution | Radiologist/Core Lab Technician | The DWI-MRI data obtained from the phantom should be analyzed to provide SNR measurements and signal intensity vs. ADC characteristics for the specific DWI-MRI acquisition protocol. |
| Test documentation | Technologist/PI | All results shall be documented and, if they pass the established acceptance values, will constitute the site qualification documentation for the DWI-MRI procedure. This process ensures study specific training of the site personnel and needs to be documented and signed. |

4.1.8   Image Analysis Workstation

The image analysis workstation shall have the ability to receive and propagate data output (imaging and metadata) collected from the prior activities (Subject Handling, Image Acquisition, Reconstruction and Post-processing).  With the input data, the analysis workstation (and software analysis tools) will be able to make use of certain attribute values to perform certain measurements and computational analysis.  The analysis workstation and software may be coupled to the MRI scanned or provided by a 3rd party vendor.

| Parameter | Entity/Actor | Specification |
| --- | --- | --- |
| Metadata | Image Analysis Workstation | Shall be able to accurately propagate the information collected at the prior stages and extend it with those items noted in the Image analysis workstation section. |

Input for Image analysis is considered output of Reconstruction and Post processing software activity.  If the Image Analyst alters input data this is considered part of Image Analysis activity.  If this occurs, the original input data will be maintained as a separate file, both to be stored, including description of manipulation in an audit trail file or in a dedicated DICOM tag section.

4.1.9     ROI definition

The methods outlined and detailed in section III 5.2 will be strictly adhered to by the following personnel:

| Parameter | Entity/Actor | Specification |
| --- | --- | --- |
| ROI definition | Image analyst/core lab | ROI definition will be performed as detailed in Section III, subsection 5.2 of the Profile. Furthermore, in order to derive meaningful data, strong attention by the image analyst (singular or core lab analyst) to subsections 5.2b,c and d, (Registration of segmentations and parameter maps, Extraction of values for statistical comparison, and choice of time point for segmentation) will be adhered to. |

4.1.10 	 Calculation of ADC

The methods outlined and detailed in section III 4.0 will be strictly adhered to by the following personnel:

| Parameter | Entity/Actor | Specification |
| --- | --- | --- |
| Parametric image formation | Image analyst/core lab | If a multi-institutional trial is utilizing the profile, than a core lab will perform these functions. For single site analysis, an image analyst will perform these functions. Regardless of the actor, analysis of DWI-MRI data is to be carried out in a series of distinct steps detailed within Section III, Sub-section 4. • Generate a native tissue ADC map using the DWI data. |

4.1.11   Image Analysis Workstation Performance Specification

The calculation of quantitative imaging biomarkers Ktrans and IAUGCBN will be performed on an image analysis workstation of sufficient power to perform with mathematical rigor the modeling and image analysis prescribed in the previous subsections.

| Parameter | Entity/Actor | Specification |
| --- | --- | --- |
| Workstation Performance | Image analyst/core lab | Workstation type and performance will be logged as part of the data log and transmittal form. Software analysis including phantom analysis will need to be performed and logged after any software upgrade. |
| Software version tracking | Image analyst/core lab | Software versions will be tracked and logged as part of the data transmittal form. |

# Comments

Need suggestions here

No DRO currently exists for DWI to my knowledge.

May not be necessary

Adjust accordingly