## Sheet1
QIBA 1C Performance Based Protocol Worksheet
Step 1 - scanner desiption
Scanner Manufacturer | Siemens
Scanner Model | Sensation 64
Site id for scanner | MP100
Software version
Date of Testing | 2009-11-01
Steps 2-4: Scanning Parameters that Meet Performance Metrics
Step 2: kVp selected | 120
Step 3: Scan Field of View (SFOV) or Patient Size | body
Step 4: Reconstructed Slice with Selected (in mm) | 1
Step 5: Collimation Rotation Time and Pitch
Number of Data Channels (N) | 32
Width of Each Data Channel (T, in mm) | 0.6
Gantry Rotation Time in seconds | 0.5
Pitch (close to 1.0 is preferred) | 1
Time to cover a 35 cm long thorax (desired < 15 seconds) | 9.1146
Step 6: Iterating on Recon Kernels to obtain 7 lp/cm
If using mA, then mA setting
(if using mA, then effective mAs equals this value): | 0
if using mAs/rotation (philips) or effective mAs (Siemens), the record setting here | 200
Recontruction Filter setting (name) - Iteration 1 | B30
Size of line pairs seen with this filter | 6
Recontruction Filter setting (name) - Iteration 2 | B45
Size of line pairs seen with this filter | 7
Recontruction Filter setting (name) - Iteration 3
Size of line pairs seen with this filter
Name of Filter that meets 7 lp/cm performance criteria | B45
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter
Step 7: Iterating on mA (or mAs) settings to obtain std dev of 12 +/1-1 HU
mA or mAs or Eff. mAs setting - Iteration 1 | 25
Standard deviation measured on image of ACR phantom | 60
mA or mAs or Eff. mAs setting - Iteration 2 | 250
Standard deviation measured on image of ACR phantom | 12.2
mA or mAs or Eff. mAs setting - Iteration 3
Standard deviation measured on image of ACR phantom
mA or eff. mAs setting that meets Standard Deviation Performance criteria | 250
Scanner reported CTDIvol (in mGy) using this mA or eff. mAs setting | 18
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter