# QIBA FDG_PET Action Item – Document procedures for computing SUV from each manufacturer’s images

2008/11/03.

Dear vendor representative

As part of the Quantitative Imaging Biomarker Alliance, the FDG-PET sub-group has an initiative to describe and document the computation of quantitative values from FDG-PET images obtained from different vendors’ devices. As part of this initiative, your company has volunteered to contribute information, and we would greatly appreciate your detailed responses to the questions asked below.

Please return your response to myself via email (mailto:dclunie@radpharm.com) with a copy to Joe Koudelik at RSNA (mailto:jkoudelik@rsna.org).

David A. Clunie, Chair, FDG-PET Quantitation Computation Technical Subcommittee

## Assumptions.

- Images will be transverse, attenuation compensated, scatter corrected, decay corrected, reconstructed, i.e., not raw data.

- Images will be encoded in DICOM PET Image Storage or Enhanced PET Image Storage SOP Class, i.e., not in proprietary formats like ECAT

## Questions for vendors.

- Can the modality operator save images with appropriate DICOM header attributes to allow the stored pixel values to be rescaled directly to SUV, and if so, is it SUVbw, SUVlbm or SUVbsa, or does the operator have a choice?

  - Yes, SUVbw. SUVlbm or SUVbsa can be displayed via Philips viewer but not stored

- Does rescaling to SUV use the standard DICOM Rescale Slope attribute to units specified in the standard DICOM Units attribute (such as GML for g/l for SUVbw), or is use of a private DICOM element required, and if so, what is it ?

  - No, private attribute only

- Can the modality operator save images with appropriate DICOM header attributes to allow the stored pixel values to be rescaled to Activity Concentration in MBq/ml?

  - Yes, in two different ways: directly from the viewer or from private attribute. This will be a non-conditional yes with a new software release (Q1 2009)

- Does rescaling to Activity Concentration use the standard DICOM Rescale Slope attribute to units specified in the standard DICOM Units attribute, or is use of a private DICOM element required, and if so, what is it?

  - Yes from viewer or new release (as above), no if from private attribute

- From saved Activity Concentration images, are the following standard DICOM attributes always populated with values that allow computation of SUVbw, SUVlbm or SUVbsa:

  - Patient’s Weight (0010,1030)	- YES

  - Patient’s Size (0010,1020), i.e., height	-YES

  - Patient’s Sex (0010,0040)	-YES

  - Radiopharmaceutical Start Time (0018,1072)	-YES

  - Decay Correction (0054,1102)	-YES

  - Decay Factor (0054,1321)	-YES

  - Frame Reference Time (0054,1300)	-YES

  - Radionuclide Half Life (0018,1075)	-YES

  - Series Time (0008,0031)	-YES

  - Radionuclide Total Dose (0018,1074)	-YES

- Are the Radionuclide Total Dose (0018,1074) and Radiopharmaceutical Start Time (0018,1072) recorded as the activity in the syringe and time measured at the start of injection, or is the modality operator also able to enter the activity dose left in the syringe at the end of injection, the time at the end of the injection, and is the decay during that time accounted for, and if so:

  - what is recorded in the standard DICOM attributes?

    - Activity at start of injection

  - what is recorded in any relevant private DICOM attributes?

    - N/A

  - what is entered and display on the console?

    - Activity at start time or injection time

  - what should be used in the calculation of SUV?

    - Not clear what you are asking

- With respect to decay correction that has already been applied to these images:

  - which of the standard terms (NONE, START, ADMIN) are encoded for Decay Correction (0054,1102)?

    - Any of the above

  - which standard DICOM attribute is any decay correction performed relative to (i.e., Series Time (0008,0031), Acquisition Time (0008,0032), Frame Reference Time (0054,1300))?

    - It corresponds to the value of a) above. Therefore, it is relative to Acquisition Time (0008,0032) or Radiopharmaceutical Start Time (0018,1072).

  - how is decay correction and the time it is relative to affected by sequential bed positions in a whole body scan (i.e., is the time the start of each bed position or a single common start time)?

    - All slices are corrected relative to the same time

  - are there any private DICOM attributes that describe this and if so what are they and what do they mean?

    - No

- How is transition across a 24-hour boundary (i.e. midnight) accounted for with respect to the times specified for dose administration, acquisition start and decay correction and the corresponding DICOM header attributes?

  - On existing systems, a heuristic is applied to determine if the scan start and pharmaceutical injection times ‘make sense’. If they do not, there is an assumption that the 24-hour barrier was crossed and decay corrections are calculated accordingly. In the upcoming software release (Q1 2009) all times are stored internally as UTC, are included in DICOM headers as date/time attributes: Acquisition Datetime (0008,002A), Radiopharmaceutical Start Datetime (0018,1078). Calculations are done accordingly.

- Are you willing to provide some reference image sets to allow verification of derivation of the same numbers with other software?

  - YES to all the below

  - of the whole body, with multiple table positions, not just one slice.

  - with example regions of interest and SUV that indicate what the “correct” (expected) result should be

  - for each permutation of

    - acquisition type

    - reconstruction method

    - correction methods