## SFOV 1
QIBA 1C Performance Based Protocol Worksheet
TEST SITE
Step 1 - scanner description
Scanner Manufacturer | GE
Scanner Model | Discovery CT750 HD sct3
Site id for scanner | Duke
Software version
Date of Testing | Nov 9th, 2010
Steps 2-4: Scanning Parameters that Meet Performance Metrics
Step 2: kVp selected | 120
Step 3: Scan Field of View (SFOV) or Patient Size | 32 cm
Step 4: Reconstructed Slice width Selected (in mm) | 1.25
Step 5: Collimation Rotation Time and Pitch
Number of Data Channels (N) | 64
Width of Each Data Channel (T, in mm) | 0.625
Gantry Rotation Time in seconds | 1
Pitch (close to 1.0 is preferred) | 0.984
Time to cover a 35 cm long thorax (desired < 15 seconds) | yes
Step 6.0: Set mAs and iterate on Recon Kernels to obtain 6 lp/cm
If using mA, then mA setting | 95
(if using mA, then effective mAs equals this value):
if using mAs/rotation (philips) or effective mAs (Siemens), record the setting here
Recontruction Filter setting (name) - Iteration 1 | standard
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 18.19
Recontruction Filter setting (name) - Iteration 2 | detail
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Noise - standard deviation measured on image of ACR phantom | 23.17
Recontruction Filter setting (name) - Iteration 3 | ASiR
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 12.32
Recontruction Filter setting (name) - Iteration 4 – IF NEEDED
Size of line pairs seen with this filter (Level =1100, Win=100)
Noise - standard deviation measured on image of ACR phantom
Name of Filter that meets 6 lp/cm performance criteria (if any) | Standard, detail, ASiR
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
for the above mAs and recon filters
Step 6.1: As needed, repeat Step 6.0 with other mA (or mAs) settings to obtain std dev of 17 +/-1 HU and 6 lp/cm
If using mA, then mA setting | 150
(if using mA, then effective mAs equals this value):
if using mAs/rotation (philips) or effective mAs (Siemens), record the setting here
Recontruction Filter setting (name) - Iteration 1 | standard
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 14.91
Recontruction Filter setting (name) - Iteration 2 | detail
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Noise - standard deviation measured on image of ACR phantom | 18.46
Recontruction Filter setting (name) - Iteration 3 | ASiR
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 10.27
Recontruction Filter setting (name) - Iteration 4 – IF NEEDED
Size of line pairs seen with this filter (Level =1100, Win=100)
Noise - standard deviation measured on image of ACR phantom
Name of Filter that meets 6 lp/cm performance criteria (if any) | Standard, detail, ASiR
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
for the above mAs and recon filters
Step 6.2: As needed, repeat Step 6.1 with other mA (or mAs) settings to obtain std dev of 17 +/-1 HU and 6 lp/cm
If using mA, then mA setting | 200
(if using mA, then effective mAs equals this value):
if using mAs/rotation (philips) or effective mAs (Siemens), record the setting here
Recontruction Filter setting (name) - Iteration 1 | standard
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 12.68
Recontruction Filter setting (name) - Iteration 2 | detail
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Noise - standard deviation measured on image of ACR phantom | 15.53
Recontruction Filter setting (name) - Iteration 3 | ASiR
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 8.97
Recontruction Filter setting (name) - Iteration 4 – IF NEEDED
Size of line pairs seen with this filter (Level =1100, Win=100)
Noise - standard deviation measured on image of ACR phantom
Name of Filter that meets 6 lp/cm performance criteria (if any) | Standard, detail, ASiR
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
for the above mAs and recon filters
Step 6.2: As needed, repeat Step 6.1 with other mA (or mAs) settings to obtain std dev of 17 +/-1 HU and 6 lp/cm
If using mA, then mA setting | 80
(if using mA, then effective mAs equals this value):
if using mAs/rotation (philips) or effective mAs (Siemens), record the setting here
Recontruction Filter setting (name) - Iteration 1 | standard
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 19.1
Recontruction Filter setting (name) - Iteration 2 | detail
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Noise - standard deviation measured on image of ACR phantom | 24.86
Recontruction Filter setting (name) - Iteration 3 | ASiR
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 14.1
Recontruction Filter setting (name) - Iteration 4 – IF NEEDED
Size of line pairs seen with this filter (Level =1100, Win=100)
Noise - standard deviation measured on image of ACR phantom
Name of Filter that meets 6 lp/cm performance criteria (if any) | Standard, detail, ASiR
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
for the above mAs and recon filters
FURTHER ITERATIOONS MAY BE REQUIRED – PLEASE EXPAND THIS SHEET AS NEEDED
mA or eff. mAs setting that meets Standard Deviation Performance criteria | 150,200,80
Scanner reported CTDIvol (in mGy) using this mA or eff. mAs setting | 10.68, 14.24, 5.69
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter

## SFOV 2
QIBA 1C Performance Based Protocol Worksheet
TEST SITE
Step 1 - scanner description
Scanner Manufacturer | GE
Scanner Model | Discovery CT750 HD sct3
Site id for scanner | Duke
Software version
Date of Testing | Nov 9th, 2010
Steps 2-4: Scanning Parameters that Meet Performance Metrics
Step 2: kVp selected | 120
Step 3: Scan Field of View (SFOV) or Patient Size | 50 cm
Step 4: Reconstructed Slice width Selected (in mm) | 1.25
Step 5: Collimation Rotation Time and Pitch
Number of Data Channels (N) | 64
Width of Each Data Channel (T, in mm) | 0.625
Gantry Rotation Time in seconds | 1
Pitch (close to 1.0 is preferred) | 0.984
Time to cover a 35 cm long thorax (desired < 15 seconds) | yes
Step 6.0: Set mAs and iterate on Recon Kernels to obtain 6 lp/cm
If using mA, then mA setting | 95
(if using mA, then effective mAs equals this value):
if using mAs/rotation (philips) or effective mAs (Siemens), record the setting here
Recontruction Filter setting (name) - Iteration 1 | standard
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 18.52
Recontruction Filter setting (name) - Iteration 2 | detail
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Noise - standard deviation measured on image of ACR phantom | 23.14
Recontruction Filter setting (name) - Iteration 3 | ASiR
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 12.91
Recontruction Filter setting (name) - Iteration 4 – IF NEEDED
Size of line pairs seen with this filter (Level =1100, Win=100)
Noise - standard deviation measured on image of ACR phantom
Name of Filter that meets 6 lp/cm performance criteria (if any) | Standard, detail, ASiR
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
for the above mAs and recon filters
Step 6.1: As needed, repeat Step 6.0 with other mA (or mAs) settings to obtain std dev of 17 +/-1 HU and 6 lp/cm
If using mA, then mA setting | 150
(if using mA, then effective mAs equals this value):
if using mAs/rotation (philips) or effective mAs (Siemens), record the setting here
Recontruction Filter setting (name) - Iteration 1 | standard
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 15.28
Recontruction Filter setting (name) - Iteration 2 | detail
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Noise - standard deviation measured on image of ACR phantom | 18.06
Recontruction Filter setting (name) - Iteration 3 | ASiR
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 9.82
Recontruction Filter setting (name) - Iteration 4 – IF NEEDED
Size of line pairs seen with this filter (Level =1100, Win=100)
Noise - standard deviation measured on image of ACR phantom
Name of Filter that meets 6 lp/cm performance criteria (if any) | Standard, detail, ASiR
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
for the above mAs and recon filters
Step 6.2: As needed, repeat Step 6.1 with other mA (or mAs) settings to obtain std dev of 17 +/-1 HU and 6 lp/cm
If using mA, then mA setting | 200
(if using mA, then effective mAs equals this value):
if using mAs/rotation (philips) or effective mAs (Siemens), record the setting here
Recontruction Filter setting (name) - Iteration 1 | standard
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 12.63
Recontruction Filter setting (name) - Iteration 2 | detail
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Noise - standard deviation measured on image of ACR phantom | 16.09
Recontruction Filter setting (name) - Iteration 3 | ASiR
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 8.56
Recontruction Filter setting (name) - Iteration 4 – IF NEEDED
Size of line pairs seen with this filter (Level =1100, Win=100)
Noise - standard deviation measured on image of ACR phantom
Name of Filter that meets 6 lp/cm performance criteria (if any) | Standard, detail, ASiR
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
for the above mAs and recon filters
Step 6.2: As needed, repeat Step 6.1 with other mA (or mAs) settings to obtain std dev of 17 +/-1 HU and 6 lp/cm
If using mA, then mA setting | 80
(if using mA, then effective mAs equals this value):
if using mAs/rotation (philips) or effective mAs (Siemens), record the setting here
Recontruction Filter setting (name) - Iteration 1 | standard
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 20.54
Recontruction Filter setting (name) - Iteration 2 | detail
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Noise - standard deviation measured on image of ACR phantom | 25.55
Recontruction Filter setting (name) - Iteration 3 | ASiR
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Noise - standard deviation measured on image of ACR phantom | 13.25
Recontruction Filter setting (name) - Iteration 4 – IF NEEDED
Size of line pairs seen with this filter (Level =1100, Win=100)
Noise - standard deviation measured on image of ACR phantom
Name of Filter that meets 6 lp/cm performance criteria (if any) | Standard, detail, ASiR
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
for the above mAs and recon filters
FURTHER ITERATIOONS MAY BE REQUIRED – PLEASE EXPAND THIS SHEET AS NEEDED
mA or eff. mAs setting that meets Standard Deviation Performance criteria | 150,200,80
Scanner reported CTDIvol (in mGy) using this mA or eff. mAs setting | 12.22, 16.30, 6.52
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter