## GE
3.0 CT Scanning Protocol | Inspiratory Lowdose CT | Expiratory CT | Inspiratory Lowdose CT | Expiratory CT | Inspiratory Lowdose CT | Expiratory CT
Scanner Make | GE | GE | GE | GE | GE | GE
Scanner Model | VCT 64 | VCT 64 | Discovery CT 750 HD 64 | Discovery 750 HD 64 | Helical | Revolution 256
Scan Type | Helical | Helical | Helical | Helical | Large | Helical
Scan FOV | Large | Large | Large | Large | 0.5 | Large
Rotation Time (s) | 0.5 | 0.5 | 0.5 | 0.5 | 64 x 0.625 | 0.5
Det. Configuration | 64 x 0.625 | 64 x 0.625 | 64 x 0.625 | 64 x 0.625 | 0.992 | 64 x 0.625
Pitch | 0.984 | 0.984 | 0.984 | 0.984 | 120 | 0.992
kVp | 120 | 120 | 120 | 120 |  | 120
mA Setting |  | 100 |  | 100 |  | 100
Dose Modulation | Auto-mA, Smart-mA ON | Auto-mA OFF | Auto-mA, Smart-mA ON | Auto-mA OFF | Auto-mA, Smart-mA ON | Auto-mA OFF
HD Mode | OFF | OFF | OFF | OFF | OFF | OFF
Target noise index | NI=66.5 |  | NI=66.5 |  | NI=50.6
Max mA limit | 100 |  | 100 |  | 100
Scan Time(s) 30cm Length | <10 | <10 | <10 | <10 | <10 | <10
Additional Image Filters | No Selection | No Selection | No Selection | No Selection | No Selection | No Selection
Thickness (mm) | 0.625 | 0.625 | 0.625 | 0.625 | 0.625 | 0.625
Interval (mm) | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Iterative Reconstruction (Noise Reduction) | ASIR | Do Not Use ASIR | ASIR | Do Not Use ASIR | ASIR | Do Not Use ASIR
Recon Mode | plus |  | plus |  | plus
Reconstructions
Insp Low 1 | Standard |  | Standard |  | Standard
Insp Low 2 | Bone |  | Bone |  | Bone
Insp Low 3i | Standard ASIR 40% |  | Standard ASIR 40% |  | Standard ASIR 40%
Insp Low 4i | Standard ASIR 100% |  | Standard ASIR 100% |  | Standard ASIR 100%
Exp Recon 1 |  | Standard |  | Standard |  | Standard
Exp Recon 2 |  | Bone |  | Bone |  | Bone
DFOV | No more than a 1 cm margin outside of the lungs on either side of the scan where the lungs are the widest in the chest. DFOV is the same for all inspiratory and expiratory reconstructions. Guidance may be provided on the CT acquisition form.

## Toshiba
3.0 CT Scanning Protocol | Inspiratory Lowdose CT | Expiratory CT
Scanner Make | Toshiba | Toshiba
Scanner Model | Acquilion Genesis | Acquilion Genesis
Scan Type | Helical | Helical
Scan FOV | Large | Large
Rotation Time (s) | 0.5 | 0.5
Det. Configuration | 80 x 1.0 | 80 x 1.0
Pitch | 0.95 | 0.95
kVp | 120 | 120
Eff. mAs |  | Effective 50 mAs
Reference mAs | Reference 50 mAs | Sure Exposure Off
Max mA limit | 100
Scan Time(s) 30cm Length | <10 | <10
Additional Image Filters | No Selection | No Selection
Thickness (mm) | 1 | 1
Interval (mm) | 0.5 | 0.5
Iterative Reconstruction (Noise Reduction) | AIDR Std & Mild | AIDR Std & Mild
Reconstructions
Insp Low 1 | FC17 (No AIDR)
Insp Low 2 | FC17
Insp Low 3i | FC52
Insp Low 4i | FC17 (AIDR Mild)
Exp Recon 1 |  | FC17 (AIDR Std)
Exp Recon 2 |  | FC52 (AIDR Mild)
DFOV | No more than a 1 cm margin outside of the lungs on either side of the scan where the lungs are the widest in the chest. DFOV is the same for all inspiratory and expiratory reconstructions. Guidance may be provided on the CT acquisition form.

## Philips
3.0 CT Scanning Protocol | Inspiratory Lowdose CT | Expiratory CT
Scanner Make | Philips | Philips
Scanner Model | Brilliance 64/256 | Brilliance 64/256
Scan Type | Spiral Helix | Spiral Helix
Scan FOV | No Selection | No Selection
Rotation Time (s) | 0.5 | 0.5
Det. Configuration | 64 x 0.625 | 64 x 0.625
Pitch | 0.923 | 0.923
kVp | 120 | 120
mAs |  | mAs 50
Reference mAs | 46 mAs (128mA)
Dose Modulation | Z-DOM or DoseRight ON | Dose Right (ACS) OFF
HD Mode | n/a | n/a
Scan Time(s) 30cm Length | <10 | <10
Additional Image Filters | Adaptive Filtering OFF | Adaptive Filtering OFF
Thickness (mm) | 0.67 | 0.67
Interval (mm) | 0.5 | 0.5
Iterative Reconstruction (Noise Reduction) | iDose ON | Do Not Use iDOSE
Reconstructions
Insp Low 1 | B
Insp Low 2 | D (YB if D not available)
Insp Low 3i | B iDose L3
Insp Low 4i | B iDose L6
Exp Recon 1 |  | B
Exp Recon 2 |  | D (YB if D not available)
DFOV | No more than a 1 cm margin outside of the lungs on either side of the scan where the lungs are the widest in the chest. DFOV is the same for all inspiratory and expiratory reconstructions. Guidance may be provided on the CT acquisition form.

## Siemens1
3.0 CT Scanning Protocol | Inspiratory Lowdose CT | Expiratory CT | Inspiratory Lowdose CT | Expiratory CT | Inspiratory CT | Expiratory CT
Scanner Make | Siemens | Siemens | Siemens | Siemens | Siemens | Siemens
Scanner Model | Biograph 40 slice | Biograph 40 slice | Sensation 64 slice | Sensation 64 slice | Definition (DS and AS) 64 slice | Definition (DS and AS) 64 slice
Scan Type | Spiral | Spiral | Spiral | Spiral | Spiral Single Source | Spiral Single Source
Scan FOV | No Selection | No Selection | No Selection | No Selection | No Selection | No Selection
Rotation Time (s) | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Det. Configuration | 40 x 0.6 | 40 x 0.6 | (64) 32 x 0.6 | (64) 32 x 0.6 | (64) 32 x 0.6 | (64) 32 x 0.6
Pitch | 1.1 | 1.1 | 1 | 1 | 1 | 1
kVp | 120 | 120 | 120 | 120 | 120 | 120
Eff. mAs | Siemens: Effective mAs 50; Siemens: Effective mAs 50; Siemens: Effective mAs 50
Reference mAs | Siemens: Reference 35mAs; Siemens: Reference 35mAs; Siemens: Reference 35mAs
Dose Modulation | Care Dose ON | Care Dose OFF | Care Dose ON | Care Dose OFF | Care Dose ON | Care Dose OFF
HD Mode | n/a | n/a | n/a | n/a | n/a | n/a
Scan Time(s) 30cm Length | <10 | <10 | <10 | <10 | <10 | <10
Additional Image Filters | No Selection | No Selection | No Selection | No Selection | No Selection | No Selection
Thickness (mm) | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75
Interval (mm) | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Iterative Reconstruction (Noise Reduction) | SAFIRE | Do Not Use IRIS or SAFIRE | SAFIRE | Do Not Use IRIS or SAFIRE | SAFIRE | Do Not Use IRIS or SAFIRE
Reconstructions
Insp Low 1 | Siemens: B31f; Siemens: B31f; Siemens: B31f
Insp Low 2 | Siemens: B45f; Siemens: B45f; Siemens: B45f
Insp Low 3i | Siemens: I31f SAFIRE 2; Siemens: I31f SAFIRE 2; Siemens: I31f SAFIRE 2
Insp Low 4i | Siemens: I31f SAFIRE 5; Siemens: I31f SAFIRE 5; Siemens: I31f SAFIRE 5
Exp Recon 1 | Siemens: B31f; Siemens: B31f; Siemens: B31f
Exp Recon 2 | Siemens: B45f; Siemens: B46f; Siemens: B45f
DFOV | Siemens: No more than a 1 cm margin outside of the lungs on either side of the scan where the lungs are the widest in the chest. DFOV is the same for all inspiratory and expiratory reconstructions. Guidance may be provided on the CT acquisition form.

## Siemens2
3.0 CT Scanning Protocol | Inspiratory Lowdose CT | Expiratory CT | Inspiratory Lowdose CT | Expiratory CT | Inspiratory Lowdose CT | Expiratory CT
Scanner Make | Siemens | Siemens | Siemens | Siemens | Siemens | Siemens
Scanner Model | Definition (AS+) 128 slice | Definition (AS+) 128 slice | Definition Flash 128 slice | Definition Flash 128 slice | Somatom Force | Somatom Force
Scan Type | Spiral | Spiral | Spiral Single Source | Spiral Single Source | Spiral Single Source | Spiral Single Source
Scan FOV | No Selection | No Selection | No Selection | No Selection | No Selection | No Selection
Rotation Time (s) | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Det. Configuration | (128) 64 x 0.6 | (128) 64 x 0.6 | (128) 64 x 0.6 | (128) 64 x 0.6 | (128) 64 x 0.6 | (128) 64 x 0.6
Pitch | bar: Siemens to Siemens
kVp | 120 | 120 | 120 | 120 | 120 | 120
Eff. mAs | Siemens: Effective mAs 50; Siemens: Effective mAs 50; Siemens: Effective mAs 50
Reference mAs | Siemens: Reference 35mAs
Dose Modulation | Care Dose ON | Care Dose OFF | Care Dose OFF | Care Dose OFF | Care Dose OFF | Care Dose OFF
HD Mode | n/a | n/a | n/a | n/a | n/a | n/a
Scan Time(s) 30cm Length | <10 | <10 | <10 | <10 | <10 | <10
Additional Image Filters | No Selection | No Selection | No Selection | No Selection | No Selection | No Selection
Thickness (mm) | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75
Interval (mm) | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Iterative Reconstruction (Noise Reduction) | SAFIRE | Do Not Use IRIS or SAFIRE | Do Not Use IRIS or SAFIRE | Do Not Use IRIS or SAFIRE | Do Not Use ADMIRE | Do Not Use ADMIRE
Reconstructions
Insp Low 1 | Siemens: B31f; Siemens: B31f; Siemens: Bf40d
Insp Low 2 | Siemens: B45f; Siemens: B46f; Siemens: Qr44d
Insp Low 3i | Siemens: I31f SAFIRE 2; Siemens: I31f SAFIRE 2; Siemens: Bf40d ADMIRE 2
Insp Low 4i | Siemens: I31f SAFIRE 5; Siemens: I31f SAFIRE 5; Siemens: Bf40d ADMIRE 5
Exp Recon 1 | Siemens: B31f; Siemens: B31f; Siemens: Bf40d
Exp Recon 2 | Siemens: B45f; Siemens: B46f; Siemens: Qr44d
DFOV | Siemens: No more than a 1 cm margin outside of the lungs on either side of the scan where the lungs are the widest in the chest. DFOV is the same for all inspiratory and expiratory reconstructions. Guidance may be provided on the CT acquisition form.