Volumetric Image Analysis for Cancer using X-Ray Computed Tomography

Version 1.6

28 February 2011

Table of Contents

I. Executive Summary	3

II. Clinical Context and Claims	3

Claim 1:  Manage Individual Patients	3

Claim 2:  Response Assessment in Clinical Trials	4

III. Profile Details	5

0. Reserved (included above)	6

1. Reserved (relevance restricted to Protocol)	6

2. Reserved (relevance restricted to Protocol)	6

3. Subject Scheduling	6

4. Subject Preparation	7

5. Imaging-related Substance Preparation and Administration	9

6. Individual Subject Imaging-related Quality Control	12

7. Imaging Procedure	12

8. Image Post-processing	21

9. Image Analysis	23

10. Image Interpretation	27

11. Archival and Distribution of Data	30

12. Quality Control	33

13. Imaging-associated Risks and Risk Management	37

IV. Compliance	38

Acquisition Devices	38

References	39

Appendices	39

Appendix A: Acknowledgements and Attributions	39

Appendix B: Background Information	40

Appendix C: Conventions and Definitions	41

Appendix D: Documents included in the imaging protocol (e.g., CRFs)	48

Appendix E: Associated Documents	48

Appendix F: TBD	48

Appendix G: Model-specific Instructions and Parameters	48

# I. Executive Summary

This Profile sets out performance claims for measuring the volumes of malignant lesions. It describes the requirements placed on human and computer-controlled actors in the following contexts: (1) managing individual patients in medical settings, and (3) quantitatively evaluating therapeutic responses in clinical trials.

Summary of Clinical Trial Usage as described in assimilated protocol "Volumetric Image Analysis for Cancer using X-Ray Computed Tomography"

This protocol describes image acquisition, quality control, processing, and analysis for cancer. The context of use is to assess longitudinal measurements of change in tumor volume over relatively short time-intervals to predict treatment response in clinical trials.

# II. Clinical Context and Claims

The clinical context sets out the utilities and endpoints for clinical trial usage and then proceeds to identify targeted levels of quality for named measurement read-outs that may be used in the relevant clinical indications.

Utilities and Endpoints for Clinical Trials

This protocol is appropriate for quantifying the volumes of malignant lesions in a variety of organ systems throughout the body except the brain, and measuring their longitudinal changes within subjects. The primary objective is to evaluate their growth or regression with serially acquired CT scans and advanced image processing techniques. The information about volumetric change will drive management decisions in diagnostic settings as well as clinical trials in patients with known malignancies. Secondary objectives may include changes in other, yet-to-be defined, image features, such as changes in mass density, vascularization, degree of spiculation, etc. In many translational research settings, there will also be cross analysis of different types of trial-derived data including biochemical, pathological and molecular biomarkers with the goal of optimizing the yield of information gleaned from early clinical trials.

Additional trial design may also include establishing the presence of certain progression events for determining time to progression (TTP) or progression free survival (PFS).

## Claim 1:  Manage Individual Patients

Measurements of tumor volume are more precise (reproducible) than uni-dimensional measurements of its longest, in-plane diameter (LD). Longitudinal changes in whole tumor volume predict health outcomes such as progression free survival (PFS) and overall survival earlier than corresponding uni-dimensional measurements of LD. Therefore, tumor response or progression as determined by changes in tumor volume will be able to serve as the primary endpoint in well-controlled efficacy studies of cytotoxic and selected targeted therapies. Changes in tumor volume can serve as the endpoint for regulatory drug approval in registration trials.

Profile specified for use with: , for the following indicated biology: , and to serve the following purpose: .

This claim identifies performance characteristics for continuous and categorical endpoints. The performance for the categorical outcome measures are expressed both as absolute performance levels as well as in comparison with using LD as the basis for Response Evaluation Criteria In Solid Tumors (RECIST).

| Measurement or Categoric Result | Performance Levels Achieved under Bull's Eye Conditions |
| --- | --- |
| Measurements of tumor volume (continuous) | If Activities are Performed at Acceptable Level Intra- and inter-rater reproducibility of >70% If Activities are Performed at Target Level Intra- and inter-rater reproducibility of >80% If Activities are Performed at Ideal Level Intra- and inter-rater reproducibility of >90% |
| Longitudinal change in tumor volume (continuous) | If Activities are Performed at Acceptable Level Intra- and inter-rater reproducibility of >80% If Activities are Performed at Target Level Intra- and inter-rater reproducibility of >90% If Activities are Performed at Ideal Level Intra- and inter-rater reproducibility of >95% |
| Tumor response or progression (categoric) | If Activities are Performed at Acceptable Level Predict Survival with coef. of corr. 85% If Activities are Performed at Target Level Predict Survival with coef. of corr. 90% If Activities are Performed at Ideal Level Predict Survival with coef. of corr. 95% |
| Tumor response or progression (categoric) | If Activities are Performed at Acceptable Level Coef. of corr. == corresponding uni-dimensional result If Activities are Performed at Target Level Coef. of corr. > corresponding uni-dimensional result If Activities are Performed at Ideal Level Can predict response with twice the sensitivity as corresponding uni-dimensional result |

## Claim 2:  Response Assessment in Clinical Trials

Measurements of tumor volume are more precise (reproducible) than uni-dimensional tumor measurements of tumor diameter.  Longitudinal changes in whole tumor volume during therapy predict clinical outcomes (i.e., OS or PFS) earlier than corresponding uni-dimensional measurements.  Therefore, tumor response or progression as determined by tumor volume will be able to serve as the primary endpoint in well-controlled Phase II and III efficacy studies of cytotoxic and selected targeted therapies (e.g., antiangiogenic agents, tyrosine kinase inhibitors, etc.) in several solid, measurable tumors (including both primary and metastatic cancers of, e.g., lung, liver, colorectal, gastric, head and neck cancer,) and lymphoma.  Changes in tumor volume can serve as the endpoint for regulatory drug approval in registration trials.

Profile specified for use with: , for the following indicated biology: , and to serve the following purpose: .

The clinical trial setting buuilds on the individual patient management setting by including those results and adding the following.

| Measurement or Categoric Result | Performance Levels Achieved under Bull's Eye Conditions |
| --- | --- |
| Make Proper GO or NO GO Decisions About New Drug Candidates or Combinations | If Activities are Performed at Acceptable Level Failure to terminate an ineffective new treatment <30% If Activities are Performed at Target Level Failure to terminate an ineffective new treatment <25% If Activities are Performed at Ideal Level Failure to terminate an ineffective new treatment <20% |

# III. Profile Details

A technical description of tests for the biomarker, identifying measurement activities and read-outs, is provided:

The following sections provide details for the various components.

## 0. Reserved (included above)

## 1. Reserved (relevance restricted to Protocol)

## 2. Reserved (relevance restricted to Protocol)

## 3. Subject Scheduling

The following sections describe requirements and considerations for the physician when scheduling imaging and other activities.

3.1. Timing Relative to Index Intervention Activity

(Moz: gotta fix this whole section. Does not apply to screening context, and seems wrong in neoadjuvant settings, when treatment must come after biopsy.) The pre-treatment CT scan must be obtained prior to administration of the study intervention and general subject to protocol requirements prior to performance of percutaneous needle biopsy. This allows for several options as to when the scan can be performed. It can occur as a dedicated CT scan of the chest performed prior to the patient being referred for needle biopsy, or alternatively, a set of images can be obtained at the time of needle biopsy before the needle is placed in the lesion. The post treatment scan should be obtained per protocol but generally on the day after completion of treatment with the study drug. The study drug must be administered for the length of time specified in the protocol subject to patient tolerance. In those cases where there is a dedicated chest CT scan and the patient then has a needle biopsy, generally the DICOM images of both studies should ideally be sent, as well as the post-treatment scan. In this instance there would be three scans for the patient.

| Index Intervention Activity | Timing |
| --- | --- |
| Pre-treatment CT scan | Acceptable Prior to any intervention on the patient, including percutaneous needle biopsy Target Ideal |
| Dedicated chest CT scan | Acceptable Immediately after biopsy Target Ideal |
| Post-treatment scan | Acceptable Day after completion of treatment with the study drug Target Ideal |

3.2. Timing Relative to confounding Activities (to minimize “impact”)

This protocol does not presume any timing relative to other activities. Fasting prior to a contemporaneous FDG PET scan or the administration of oral contrast for abdominal CT is not expected to have any adverse impact on this protocol.

| Confounding Activity | Timing |
| --- | --- |
| None noted | Acceptable Target Ideal |

3.3. Scheduling Ancillary Testing

If associated biopsy/resection is expected to be performed during the same visit as the imaging procedure, consider describing that association here. If not, it can be covered in the Trial Calendar.

| Ancillary Test | Scheduling |
| --- | --- |
|  | Acceptable Target Ideal |

## 4. Subject Preparation

The following sections describe how subjects are prepared.

4.1. Prior to Arrival

Preparation needed in addition to the local standard of care for CT with contrast.

| Preparation Step | Compliance Levels |
| --- | --- |
| Consent | Acceptable Informed consent Target Consent to share de-identified imaging data Ideal Consent to share all de-identified trial data |

4.2. Upon Arrival

The following sections describe steps taken upon arrival.

4.2.1. Confirmation of subject compliance with instructions

No preparation is specified beyond the local standard of care for CT with contrast.

| Instruction | Compliance Levels |
| --- | --- |
| None noted | Acceptable N/A Target Ideal |

4.2.2. Ancillary Testing

Biopsy Performance Guidelines: The person performing the tumor sampling procedure (fine needle aspiration biopsy (FNA) or core biopsy) will either 1) already know that a particular patient is being enrolled into the study, or 2) understand that the patient may be a candidate for the study. In either case, informed consent must be obtained prior to performance of the biopsy to enable some of the aspirated material to be used once it is determined that the patient actually does meet the criteria for enrollment on the study. A example consent form is provided (Fine Needle Aspiration Lung Biopsy Registry)

Biopsy should be performed according to the sites standard of care.  Typically, biopsy should be performed with CT guidance to allow for pre-biopsy images to be obtained. If images are already available pre-biopsy that meet the study criteria, then fluoroscopic biopsy is acceptable. Either aspirated material for cytology, or specimens from a core biopsy are acceptable subject to the process of the local institution.

When possible, the pre-treatment low dose CT scan of the chest can be performed at the same time as the biopsy, in lieu of performing them separately (see CT Imaging Guidelines). The extent of the scan can be limited. Ideally an onsite pathologist (cytopathologist) will be present during the biopsy. Subject to IRB-approval, once confirmation of adequate sample for a diagnosis has been made, additional material that has already been obtained can be used for the purposes of the study. A separate description of the preparation of the cytology material is provided (see Specimen Preparation Guidelines)

| Ancillary Test | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

4.2.3. Preparation for Exam

Beyond a clear, simple language description of the image acquisition procedure, no exam preparation is specified beyond the local standard of care for CT with contrast.

| Preparation Step | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

## 5. Imaging-related Substance Preparation and Administration

The following sections describe imaging-related substance preparation and administration.

5.1. Substance Description and Purpose

The use of contrast is generally not a requirement for this protocol. However, the use of intravenous contrast material may be medically indicated for the diagnosis and staging of lung cancer in defined clinical settings.  Contrast characteristics influence the appearance and quantification of the tumors; therefore, a given subject must be scanned on the follow-up exam using the same conditions as the baseline scan which means that if no contrast is given at baseline, then the follow-up scan would also be done without contrast to ensure accurate volume change comparison.

| Parameter | Compliance Levels |
| --- | --- |
| Brand of contrast agent | Acceptable Another brand or switch of contrast agent type may be used if medically indicated, e.g., a switch from ionic to non-ionic contrast media Target A subject should be scanned with the same brand of contrast agent for each scan Ideal All subjects should be scanned with equivalent contrast media |
| Use of contrast in follow-up scans | Acceptable If used at baseline, continue using it. If not used, do not use in follow-up scans. Target Do not use contrast at baseline or other scans. Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Whether contrast was used Target Contrast media brand Ideal |

5.2. Dose Calculation and/or Schedule

Site-specific sliding scales that have been approved by local medical staffs and regulatory authorities should be used for patients with impaired renal function (e.g., contrast dose reduction based on creatinine clearance).

| Parameter | Compliance Levels |
| --- | --- |
| Dose calculation for a given subject | Acceptable If a different brand or type of contrast is used, the dose may be adjusted to ensure comparability as indicated and as documented by peer-reviewed literature and/or the contrast manufacturers’ package inserts Target For a given subject, the same contrast dose should be used for each scan subject to the medical condition of the patient Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Contrast media dose calculations and schedule Target Ideal |

5.3. Timing, Subject Activity Level, and Factors Relevant to Initiation of Image Data Acquisition

Andy all this stuff about contrast is confusing—what it you pulled it out of here and put it into an appendix for reference for those studies requiring contrast—as they will be unusual Typically a volumetric CT will not be performed using contrast. However if forsome reason contrast is used, ...

Scan delay after contrast administration is dependent upon the both the dose and rate of administration, as well as the type of scanner being used. Contrast administration should be tailored for both the vascular tree as well as optimization of lesion conspicuity in the solid organs. (These guidelines do not refer to perfusion imaging of single tumors.) Generally, since there are multiple concentrations of contrast as well as administration rates and scanning speeds, it is difficult to mandate specific values. Generally, institutional guidelines should be followed so as to optimize reproducibility of the scan technique.  <Should we discuss adjustment of imaging delay and/or timing on the basis of cardiac output as determined by some sort of pre-imaging bolus protocol?>>

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Image acquisition should start at the same time after contrast administration for each scan Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Timing and subject activity level Target All factors relevant to initiation of image acquisition Ideal |

5.4. Administration Route

The following requirements are noted.

| Parameter | Compliance Levels |
| --- | --- |
| Administration route | Acceptable Intravenous bolus injection in any vein via butterfly catheter Target Injection via butterfly in a large antecubital vein Ideal Injection in a large antecubital vein known to be patent from observation of intravenous saline drip |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Actual administration details Target Ideal |

5.5. Rate, Delay and Related Parameters / Apparatus

**Describe the rate, delay, and related parameters or apparatus. Place needed requirements.

| Parameter | Compliance Levels |
| --- | --- |
| Contrast administration | Acceptable Manually Target At the same rate for each scan Ideal Via a power injector |
| If a different brand or type of contrast is used | Acceptable The rate may be adjusted to ensure comparability if appropriate and as documented by peer-reviewed literature and/or the contrast manufacturers’ package inserts Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Rate and delay of contrast media Target Related parameters and apparatus utilized Ideal |

5.6. Required Visualization / Monitoring, if any

No particular visualization or monitoring is specified beyond the local standard of care for CT with contrast.

| Parameter | Compliance Levels |
| --- | --- |
| None noted | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Actual events observed Target Ideal |

5.7. Quality Control

See 12.2

## 6. Individual Subject Imaging-related Quality Control

See 12.3

## 7. Imaging Procedure

A set of scout images should be initially obtained. Next, in a single breath hold, contiguous thin section slices from the thoracic inlet to the adrenal glands are obtained. Pitch should be chosen so as to allow completion of the scan in a single breath hold. In some cases two or more breaths may be necessary. In those cases, it is important that the target lesion be fully included within one of the sequences. The use of contrast material is not involved. Once the scan is complete, if possible, a targeted image should be created with a small field of view (FOV), through the target lesion. This should be retrospectively performed on the same day that the scan is obtained (so as to prevent loss of raw data), as there is no need for an additional acquisition. The targeted images should cover the entire lesion, with no cutoff at the top or bottom. They should be reconstructed with approximately 50% overlapping images. This should be saved as a separate series and sent with the original scan to the coordinating center.

All efforts should be made to have the second scan performed with identical parameters as the first. This should be inclusive of as many of the scanning parameters as possible, and preferably be performed on the same scanner. This also includes the same FOV for the targeted series.

7.1. Required Characteristics of Resulting Data

This section describes characteristics of the acquired images that are important to this protocol. Characteristics not covered here are left to the discretion of the participating site.  Additional details about the method for acquiring these images are provided in section 7.2.

7.1.1. Data Content

These parameters describe what the acquired images should contain/cover.

Field of View affects pixel size due to the fixed image matrix size used by most CT scanners. If it is clinically necessary to expand the field of view to encompass more anatomy, the resulting larger pixels are acceptable.

<<would be value to have 1024x1024 images allowing for encoding of greater in-plane resolution while maintaining a large field of view, however there are many equipment issues to deal with along this path>>

<<Note tradeoffs with dose and the option of doing targeted scans at higher dose to keep lower dose elsewhere>>

| Parameter | Compliance Levels |
| --- | --- |
| Anatomic Coverage | Acceptable (need to express consistent with this being a whole-body document) Target Ideal |
| Field of view | Acceptable (need to express consistent with this being a whole-body document) Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Anatomic coverage and field of view Target Ideal |

7.1.2. Data Structure

These parameters describe how the data should be organized/sampled.

Collimation Width (defined as the total nominal beam width) is often not directly visible in the scanner interface. Wider collimation widths can increase coverage and shorten acquisition, but can introduce cone beam artifacts which may degrade image quality.

<<Discuss single slice collimation width in here somewhere>>

Slice intervals (a.k.a. "reconstruction intervals" that result in discontiguous data are unacceptable as they may “truncate” the spatial extent of the tumor, degrade the identification of tumor boundaries, confound the precision of measurement for total tumor volumes, etc.

Pitch impacts dose since the area of overlap results in additional dose to the tissue in that area. Overlaps of greater than 20% have insufficient benefit to justify the increased exposure.

Slice Width directly affects voxel size along the subject z-axis. Smaller voxels are preferable to reduce partial volume effects and (likely) provide higher precision due to higher spatial resolution.

Pixel Size directly affects voxel size along the subject x-axis and y-axis. Smaller voxels are preferable to reduce partial volume effects and (likely) provide higher measurement precision.

Isotropic Voxels are expected to improve the reproducibility of tumor volume measurements, since the impact of tumor orientation (which is difficult to control) is reduced by more isotropic voxels.

<<Smoothing/filtering in the different dimensions is somewhat relevant to this concept as well. Not all scanners can do the same things here>>

<<Note that we don’t want people to throw away resolution to match the worse to the better>>

<<Clarify that pixel size in each dimension is not the same as resolution in each dimension, have to recognize that inherent resolution is different than how the data happens to be sliced and diced.>>

Scan Plane may differ for some subjects due to the need to position for physical deformities or external hardware, but should be constant for each scan of a given subject.

Faster Rotation Speed reduces the breath hold requirements and reduces the likelihood of motion artifacts.

| Parameter | Compliance Levels |
| --- | --- |
| Collimation Width | Acceptable 5 to 125mm Target 10 to 80mm Ideal 20 to 40mm |
| Slice Interval | Acceptable Contiguous or up to 50% overlap Target Ideal |
| Slice Width | Acceptable <= 3mm Target <= 2.5mm Ideal <= 1.0mm |
| Pixel Size | Acceptable <1.0mm Target 0.75 to 1.0mm Ideal <0.75mm |
| Isotropic Voxels | Acceptable (5:1) slice width <= 5 x pixel size Target (1:1) slice width = pixel Size Ideal |
| Scan Plane | Acceptable Same for each scan of subject Target 0 azimuth Ideal |
| Rotation Speed | Acceptable Sufficient for single breath-hold acquisition Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Collimation width, slice interval, slice width, pixel size, voxel dimensionality, scan plane, rotation speed Target Ideal |
| <actual testing, e.g., to see if actual slice width is what we think it is> | Acceptable Target Ideal |

7.1.3. Data Quality

These parameters describe the quality of the images:

Motion Artifacts may produce false targets and distort the size of existing targets. “Minimal” artifacts are such that motion does not degrade the ability of image analysts to detect the boundaries of target lesions.

Noise Metrics quantify the level of noise in the image pixel values. The procedure for obtaining the noise metric for a given acquisition protocol on a given piece of equipment is described in section XX. Greater levels of noise may degrade segmentation by image analysis operators or automatic edge detection algorithms.  Noise can be reduced by using thicker slices for a given mAs. A constant value for the noise metric might be achieved by increasing mAs for thinner slices and reducing mAs for thicker slices.

Spatial Resolution Metric quantifies the ability to resolve spatial details. It is stated in terms of the number of line-pairs per cm that can be resolved in a scan of resolution phantom (such as the synthetic model provided by the American College of Radiology and other professional organizations). Lower spatial resolution can make it difficult to accurately determine the borders of tumors, and as a consequence, decreases the precision of volume measurements.  Spatial resolution is mostly determined by the scanner geometry (which is not usually under user control) and the reconstruction algorithm (which is under user control).

| Parameter | Compliance Levels |
| --- | --- |
| Motion Artifacts | Acceptable Minimal artifact Target No artifact Ideal |
| Noise Metrics | Acceptable Std. dev. in 20cm water phantom < 40 HU Target Ideal |
| Spatial Resolution | Acceptable >= 6 lp/cm Target >= 7 lp/cm Ideal >= 8 lp/cm |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Motion artifacts, noise metrics, spatial resolution settings Target Ideal |
| <actual testing, e.g., to see if actual result is what settings would suggest they should be> | Acceptable Target Ideal |

7.2. Imaging Data Acquisition

The following sections describe the acquisition of imaging data.

7.2.1. Subject Positioning

Consistent positioning is required to avoid unnecessary variance in attenuation, changes in gravity induced shape, or changes in anatomical shape due to posture, contortion, etc. Careful attention should be paid to details such as the position of their upper extremities, the anterior-to-posterior curvature of their spines as determined by pillows under their backs or knees, the lateral straightness of their spines, and, if prone, the direction the head is turned.

Factors that adversely influence patient positioning or limit their ability to cooperate (breath hold, remaining motionless, etc.) should be recorded in the corresponding DICOM tags and case report forms, e.g., agitation in patients with decreased levels of consciousness, patients with chronic pain syndromes, etc.

If the previous positioning is unknown, the subject should be positioned Supine/Arms Up/Feet First if possible. This has the advantage of promoting consistency, and reducing cases where intravenous lines, which could introduce artifacts, go through gantry.

| Parameter | Compliance Levels |
| --- | --- |
| Subject positioning | Acceptable They may be placed in a different position if medically unavoidable due to a change in clinical status Target Same positioning should be used for each scan Ideal Supine/Arms Up/Feet First |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Subject positioning shall be recorded, manually by the staff Target In the image dataset header Ideal |

7.2.2. Instructions to Subject During Acquisition

Breath holding reduces motion which might degrade the image. Full inspiration inflates the lungs which is necessary to separate structures and make lesions more conspicuous.

| Parameter | Compliance Levels |
| --- | --- |
| Breath hold | Acceptable At least near the high end inspiration Target Subjects should be instructed to hold a single breath at full inspiration Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable Patient compliance Target Ideal |

7.2.3. Timing/Triggers

For each subject, the time-interval between the administration of intravenous contrast and the start of the image acquisition should be determined in advance, and then maintained as precisely as possible during all subsequent examinations.  For lung masses, image acquisition should be timed to coincide with visualization of the thoracic arteries. For sub-diaphragmatic acquisitions, timing should coincide with opacification of the portal-venous blood vessels.

| Parameter | Compliance Levels |
| --- | --- |
| Timing | Acceptable Use a standard time Target Evaluate “manually” Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| DICOM recording | Acceptable timing/triggers Target Ideal |

7.2.4. Model-specific Parameters

Appendix G.1 lists acquisition parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

7.2.5. Archival Requirements for Primary Source Imaging Data

See 11.3

7.2.6. Quality Control

See 12.3

7.3. Imaging Data Reconstruction

The following sections describe image reconstruction.

7.3.1. Input Data to Be Used

See "Data Quality" specifications from Acquisition section.

| Parameter | Compliance Levels |
| --- | --- |
| No additional | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| <actual testing, e.g., to see if input data is actually what the DICOM fields suggests it should be> | Acceptable Target Ideal |

7.3.2. Methods to Be Used

These parameters describe general characteristics of the reconstruction:

Reconstruction Kernel Characteristics should be defined to optimize the analysis for each study. The same kernel must be used for each scan of a given subject and should be consistent across all study participants. A softer kernel can reduce noise at the expense of spatial resolution. An enhancing kernel can improve resolving power at the expense of increased noise. Moderation on both fronts is recommended with a slight bias towards enhancement.

Reconstruction Interval should be the same for each scan of a given subject.

Reconstruction Overlap should be the same for each scan of a given subject. Decisions about overlap should consider the technical requirements of the clinical trial, including effects on measurement, throughput, image analysis time, and storage requirements. Reconstructing datasets with overlap will increase the number of images and may slow down throughput, increase reading time and increase storage requirements.  It should be noted that for multidetector row CT (MDCT) scanners, creating overlapping image data sets has NO effect on radiation exposure; this is true because multiple reconstructions having different kernel, slice thickness and intervals can be reconstructed from the same acquisition (raw projection data) and therefore no additional radiation exposure is needed.

| Parameter | Compliance Levels |
| --- | --- |
| Reconstruction Kernel Characteristics | Acceptable Target Ideal |
| Reconstruction Interval | Acceptable Reconstruction intervals that are less than half the axial extent of the mass and typically ≤2.5 mm Target Reconstruction intervals that are less than one third the axial extent of the mass and typically ≤1.25 mm Ideal Reconstruction intervals that are less than one tenth the axial extent of the mass and typically ≤.625 mm |
| Reconstruction Overlap | Acceptable Contiguous (e.g., 5mm thick slices, spaced 5mm apart or 1.25mm spaced1.25 mm apart) Target 20% overlap (e.g. 5mm thick slices, spaced 4mm apart or 1.25mm spaced 1mm apart) Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.3.3. Required Characteristics of Resulting Data

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.3.4. Platform-specific Instructions

Appendix G.2 lists reconstruction parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.2.

7.3.5. Archival Requirements for Reconstructed Imaging Data

See 11.4

7.3.6. Quality Control

See 12.4

## 8. Image Post-processing

**Introduce the Image Post-processing section.

No post-processing shall be performed on the reconstructed images sent for image analysis.  Such processing, if performed, has the potential to disrupt the consistency of the results.

8.1. Input Data to Be Used

See "Required Characteristics of Resulting Data" specifications from Reconstruction section.

| Parameter | Compliance Levels |
| --- | --- |
| None noted | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| <actual testing, e.g., to see if input data is actually what the DICOM fields suggests it should be> | Acceptable Target Ideal |

8.2. Methods to Be Used

**Describe how the analysis should be performed. (e.g. algorithms to be used; where measurements should be taken; definition of key anatomical points or pathology boundaries; related annotations)

| Parameter | Compliance Levels |
| --- | --- |
| None noted | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

8.3. Required Characteristics of Resulting Data

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
| None noted | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

8.4. Platform-specific Instructions

Appendix G.3 lists post-processing parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 8.

8.5. Archival Requirements

See 11.5

8.6. Quality Control

See 12.5

## 9. Image Analysis

Each lung lesion shall be characterized as described in this section.  Lesions of interest include: a) small pulmonary Masses surrounded by air; b) small to medium pulmonary Masses surrounded by air and/or with adjacent normal and abnormal (non-neoplastic) anatomic structures; c) large pulmonary masses surrounded by air and/or with adjacent normal and abnormal (non-neoplastic) anatomic structures and/or confluent with mediastinum, chest wall, and diaphragm.

Neoplastic tissue volume? necrotic tissue volume?

Is there a gold standard structured reported format that we want to recommend?

(Pulmonary hemorage, hemorage into a cystic lesion, etc.)Fluid, blood, necrotic debris, and the like should not be included in the measurement of tumor volume. Procedures for segmenting tissue types within a mass are not described by this UPICT protocol but should be implemented when technically feasible.

9.1. Input Data to Be Used

The reconstructed images may be used directly since no post-processing is specified.  No other data is required for this Analysis step.

| Parameter | Compliance Levels |
| --- | --- |
| Reconstruction Results | Acceptable Reconstruction results conforming to Acceptable specifications Target Reconstruction results conforming to Target specifications Ideal Reconstruction results conforming to Ideal specifications |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| <actual testing, e.g., to see if input data is actually what the DICOM fields suggests it should be> | Acceptable Target Ideal |

9.2. Methods to Be Used

Each lesion shall be characterized by determining the boundary of the lesion (referred to as segmentation) and taking certain measurements of the segmented lesion.  Segmentation may be performed automatically by a software algorithm, manually by a human observer, or semi-automatically by an algorithm working with human guidance/intervention.  Measurements may be performed automatically by a software algorithm, manually by a human observer with “e-calipers”, or semi-automatically by an algorithm working with human guidance/intervention.

While all measurement metrics are proxies for tumor burden, it is still uncertain which measurement metric is optimal for assessing changes in health status. Accordingly, multiple overlapping measurements are specified here. The annotation and image markup shall be defined to be optimal for each clinical study but generally should include measurements for each target lesion as well as aggregate measures across target lesions as set out in the parameter table below.

Longest diameter (LD), in mm (uni-dimensional metric) which is defined as the longest continuous, in-plane line-length that can be placed within a non-nodal tumor mass on a transaxial image. The LD should correspond to the greatest distance between two in-plane voxels anywhere in the stack of images on which the mass can be visualized. It is expected that the axial level on which the LD will be derived will vary from time-point to time-point.  Short Axis: The short axis is defined as the longest line that is perpendicular to the longest, in-plane line length (see nearby diagram). In contrast to extra-nodal masses, the length of the short-axis is the RECIST 1.1 outcome measure for lymph nodes.

In this illustration showing the longest axis and the short axis of a lymph node, the horizontal or east-west axis is the longest.

The longest perpendicular (LP), in mm (bi-dimensional metric), that is, the longest, in-plane line that can be placed at a 90 degree angle to the LD on the one slice containing the LD for a tumor.

Regarding manual volume measurement, contours drawing is subject to the same constraints as defined on RECIST 1.1, section « Measurement of lesions »: Delineation should be performed in the plane in which the images were acquired.

Sum of the Diameters (SOD): A value computed by adding up all of the longest diameters (LDs) of all of the non-nodal Target Lesions and the Short Axis of each Target lymph node.

| Parameter | Compliance Levels |
| --- | --- |
| Lesion Volume | Acceptable mm3 or mL (volumetric metric) which is defined as the sum of all the voxel volumes within the boundaries of a discrete tumor mass on all the tomographic slices on which it is visible. Target Actual volume as would be calculated without regard to spatial sampling loss Ideal Actual volume measure for irregularly shaped lesions that would be deemed non-measurable using unidimensional axial diameters |
| Sum of Target Lesion Volumes | Acceptable A value computed by adding up all of the target lesion volumes calculated using Acceptable approach above Target A value computed by adding up all of the target lesion volumes calculated using Target approach above Ideal A value computed by adding up all of the target lesion volumes calculated using Ideal approach above |
| LD, LP, SOD | Acceptable As defined in RECIST 1.1 Target Ideal |

The following set of requirements extends what has been stated in the protocol.

<Technical parameters such as algorithm developers need, e.g., how well it finds boundaries when the tumor to background contrast is such and such, when timor is attached, when ...>

| Parameter | Compliance Levels |
| --- | --- |
| Automated boundary segmentation | Acceptable Can provide starting segmentation of various types of lung tumors on CT images; including a) small pulmonary Masses surrounded by air b) small to medium pulmonary Masses with adjacent normal and abnormal (non neoplastic) anatomic strucutres c) large pulmonary masses confluent with mediastinum, chest wall, diaphragm as productivity tool but with many lesions requiring reader correction Target Can segment various types of lung tumors on CT images; including a) small pulmonary Masses surrounded by air b) small to medium pulmonary Masses with adjacent normal and abnormal (non neoplastic) anatomic strucutres c) large pulmonary masses confluent with mediastinum, chest wall, diaphragm automatically most of the time without reader correction Ideal Can segment various types of lung tumors on CT images; including a) small pulmonary Masses surrounded by air b) small to medium pulmonary Masses with adjacent normal and abnormal (non neoplastic) anatomic strucutres c) large pulmonary masses confluent with mediastinum, chest wall, diaphragm fully automatically without reader correction |
| Automatically computed read-outs | Acceptable Can produce unidimensional (greatest maximal diameter), bidimensional (greatest maximal diameter and largest perpendicular) and volumetric quantification of the segmented tumor Target Additional measurements which would be acceptable to obtain are WHO measurement: maximum 3D diameter; shape parameters like roundishness or others; HU mean, max, min, standard deviation within the segmentation Ideal Provide error margins for each measurement Provide a HU-histogram of the segmented voxels |
| Compute speed | Acceptable The software can process an algorithm in xx (a specified period of time) Target Ideal |
| Number of mouse clicks and screen operations for reader | Acceptable The maximum number of mouse clicks to perform segmentation on a lesion is xx Target Ideal |
| Image markup | Acceptable Markups of lesion segmentation boundary and calculated unidimensional and volumetric measurement will be available for audit purposes Target Ideal |
| Recording | Acceptable DICOM fields as indicated in various sections above Target Software can output image and segmentation results in the following output formats: DICOM SR; DICOM RTSS; upcoming DICOM segmentation format; STL-Format; DICOM secondary capture; XLS, CSV, XML Ideal |

9.3. Required Characteristics of Resulting Data

It is expected that automated boundary detection algorithms will place segmentation edges with greater precision, accuracy and speed than an operator can draw by hand with a pointing device. It is also expected that automated algorithms for finding the Longest Diameter (LD) and Longest Perpendicular (LP) within each ROI will have greater speed and precision of measurement than an operator using electronic calipers. The performance of the algorithms will, however, depend on the characteristics of the lesions may be challenged by complex lung tumors.   Operator assisted semi-automatic segmentation should produce the same level of intra- and inter-rater reliability for the volume measurements of each target lesion.

| Parameter | Compliance Levels |
| --- | --- |
| Intra-rater reliability of fully automated read-outs | Acceptable Greater than 80% Target Greater than 90% Ideal Greater than 95% |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| Annotation and image markup | Acceptable Compliant to BRIDG information model Target Compliant to Study Data Tabulation Model (SDTM) Ideal |
| Case report form | Acceptable ... Target Ideal |

9.4. Platform-specific Instructions

Appendix G.4 lists image analysis parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 9.

9.5. Archival Requirements

See 11.6

9.6. Quality Control

See 12.6

## 10. Image Interpretation

While Analysis is primarily about computation; Interpretation is primarily about judgment. Interpretation may be performed at both the lesional / target level and in the aggregate at the subject level (e.g., in an oncology study each index lesion may be measured in longest diameter during the analysis phase, but in this phase a judgment may be made as to whether there is a new “non-index” lesion; the aggregation of the measured lesions with comparison to previous studies coupled with the judgment as to the presence or absence of a new lesion will result in the RECIST classification at the subject level).

The following definitions apply to the interpretation of resulting data from volumetric image analysis:

Baseline: The measurements based on the pre-treatment scan set acquired most closely to the start of treatment.

Nadir: The lowest value for the sum of the longest diameters or sum of the volumes of all target lesions. In effect, the nadir is the new "baseline equivalent" value for assessing progression.

Progressive Disease (PD) is defined as an increase in the sum of the longest diameters by 20% from nadir, if but only if, the new sum of the diameters (SOD) exceeds 5 mm. For volumes, PD is defined as an increase of more than 20% or two times the Standard Deviation of measurement, whichever is greater.

Non-Measurable Lesions: Neoplastic masses that, in their longest uninterrupted diameter at baseline, are too small to measure because the greatest distance between any two in-plane pixels is less than two times the axial slice thickness.

Non-Target Lesions (NTL): Additional neoplastic masses that meet the criteria of target lesions but are not selected for quantitative assessment, neoplastic masses that either do not meet the minimum size criteria or are not suitable for repeat measurement, and any truly non-measurable lesions, such as bone metastases, leptomeningeal metastases, malignant ascites, pleural/pericardial effusion, inflammatory breast disease, lymphangitis cutis/pulmonis, cystic lesions, ill defined abdominal masses, etc. Non-target lesions must be followed qualitatively.

10.1. Input Data to Be Used

See "Required Characteristics of Resulting Data" specifications from Image Analysis section.

| Parameter | Compliance Levels |
| --- | --- |
| Analysis Read-outs | Acceptable Read-outs conforming to Acceptable specifications Target Read-outs conforming to Target specifications Ideal Read-outs conforming to Ideal specifications |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| <actual testing, e.g., to see if input data is actually what the annotation and image markup fields and/or eCRFs suggests it should be> | Acceptable Target Ideal |

10.2. Methods to Be Used

**Describe how the interpretation should be performed. (e.g. definition of key anatomical points or pathology boundaries; scoring scales and criteria such as BIRADS, interpretation schema such as RECIST, related annotations)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

10.3. Required Characteristics of Resulting Data

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

10.4. Platform-specific Instructions

Appendix G.5 lists image interpretation parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 9.

10.5. Archival Requirements

See 11.7

10.6. Quality Control

See 12.7

## 11. Archival and Distribution of Data

The following sections describe the archival and distribution of data.

11.1. Reserved

11.2. Reserved

11.3. Primary Source Imaging Data

This protocol presumes no archiving the pre-reconstruction image data.

| Parameter | Compliance Levels |
| --- | --- |
| None noted | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| search fields, whther set up for content-based retrieval, ... | Acceptable Target Ideal |

11.4. Reconstructed Imaging Data

Reconstructed images shall be archived locally, formatted as either DICOM CT image objects or DICOM Enhanced CT image objects.  Retention period and policy is left to the Clinical Trial Protocol author and requirements of the FDA or other relevant study sponsor or regulatory body.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| search fields, whther set up for content-based retrieval, ... | Acceptable Target Ideal |

11.5. Post-Processed Data

No post processing is specified, however if post-processing is performed, the images shall be archived the same as 11.4.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| search fields, whther set up for content-based retrieval, ... | Acceptable Target Ideal |

11.6. Analysis Results

Segmentation results may be recorded as DICOM Segmentation Objects, or STL Model Files.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| search fields, whther set up for content-based retrieval, ... | Acceptable The data described in 9.3 may be provided in any of the following formats: DICOM SR, DICOM RTSS, DICOM secondary capture, XLS, CSV, XML Target Ideal |

11.7. Interpretation Results

The following requirements are placed on interpretation results.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| search fields, whther set up for content-based retrieval, ... | Acceptable Target Ideal |

## 12. Quality Control

This is the section where all aspects of Quality Control are identified and described.

12.1. QC Associated with the Site

The following sections consider various aspects of quality control.

12.1.1. Quality Control Procedures

**Describe required procedures and documentation for routine and periodic QC for the site and various pieces of equipment.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.1.2. Reserved

12.1.3. Reserved

12.2. QC Associated with Imaging-related Substance Preparation and Administration

The following requirements are placed on QC associated with imaging-related substance preparation and administration.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.3. QC Associated with Individual Subject Imaging

**Introduce the section

The following sections describe aspects of QC associated with individual subject imaging.

12.3.1. Phantom Imaging and/or Calibration

**Introduce the section

The following requirements are placed on phantom imaging and/or calibration.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.3.2. Quality Control of the Subject Image and Image Data

The following requirements are placed on QC of the subject image and image data.

| Parameter | Compliance Levels |
| --- | --- |
| Acquisition System Calibration | Acceptable Site staff shall conform to the QA program defined by the device manufacturer. Target A protocol specific calibration and QA program shall be designed consistent with the goals of the clinical trial. This program may include (a) elements to verify that sites are performing the specified protocol correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address a limited set of issues primarily relating dose and image quality (such as water calibration and uniformity). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial. Ideal A protocol specific calibration and QA program shall be designed consistent with the goals of the clinical trial. This program shall include (a) elements to verify that sites are performing the specified protocol correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address issues relating to both radiation dose and image quality (which may include issues relating to water calibration, uniformity, noise, spatial resolution -in the axial plane-, reconstructed slice thickness z-axis resolution, contrast scale, CT number calibration and others). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial. |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.4. QC Associated with Image Reconstruction

The following requirements are placed on QC associated with image reconstruction.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.5. QC Associated with Image Processing

The following requirements are placed on QC associated with image processing.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.6. QC Associated with Image Analysis

The following requirements are placed on QC associated with image analysis.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.7. QC Associated with Interpretation

The following requirements are placed on QC associated with interpretation.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

## 13. Imaging-associated Risks and Risk Management

The following sections consider various imaging-associated risks and risk management.

13.1. Radiation Dose and Safety Considerations

**Introduce the section.

It is recognized that X-ray CT uses ionizing radiation and this poses some small, but non-zero risk to the patients in any clinical trial. The radiation dose to the subjects in any trial should consider the age and disease status (e.g. known disease or screening populations) of these subjects as well as the goals of the clinical trial. These should inform the tradeoffs between desired image quality and radiation dose necessary to achieve the goals of the clinical trial.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13.2. Imaging Agent Dose and Safety Considerations

The following requirements are placed regarding imaging agent dose and safety considerations.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13.3. Imaging Hardware-specific Safety Considerations

The following requirements are placed regarding hardware-specific safety considerations.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13.4. Reserved

13.5. Reserved

# IV. Compliance

## Acquisition Devices

In addition to specifications as set out in the Image Acqusition section above:

• while multi-slice is not required, it will produce better results.  Acceptable: 16-slice, Target: 64-slice, Ideal: 64 or greater

o <<Consider specifying the point spread function?>>

• conforms to the Medical Device Directive Quality System and the Essential Requirements of the Medical Device Directive

• designed and tested for safety in accordance with IEC 601-1, as well as for ElectroMagnetic Compatibility (EMC) in accordance with the European Union’s EMC Directive, 89/336/EEC

• Labelled for these requirements, as well as ISO 9001 and Class II Laser Product, at appropriate locations on the product and in its literature

• CSA compliant

# References

**a reference

# Appendices

## Appendix A: Acknowledgements and Attributions

This imaging protocol is proffered by the Radiological Society of North America (RSNA) Quantitative Imaging Biomarker Alliance (QIBA) Volumetric Computed Tomography (v-CT) Technical Committee. The v-CT technical committee is composed of scientists representing the imaging device manufacturers, image analysis software developers, image analysis laboratories, biopharmaceutical industry, academia, government research organizations, professional societies, and regulatory agencies, among others. All work is classified as pre-competitive. A more detailed description of the v-CT group and its work can be found at the following web link: http://qibawiki.rsna.org/index.php?title=Volumetric_CT.

The Volumetric CT Technical Committee (in alphabetical order):

- Athelogou, M. Definiens AG

- Avila, R. Kitware, Inc.

- Beaumont, H. Median Technologies

- Borradaile, K. Core Lab Partners

- Buckler, A. Buckler Biomedical Sciences LLC

- Clunie, D. Core Lab Partners

- Cole, P. Imagepace

- Dorfman, G. Weill Cornell Medical College

- Fenimore, C. Nat Inst Standards & Technology

- Ford, R. RadPharm, Inc.

- Garg, K. University of Colorado

- Gustafson, D. Intio, Inc.

- Gottlieb, R. Roswell Park Cancer Center

- Hayes, W. Bristol Myers Squibb

- Hillman, B. Metrix, Inc.

- Kim, HG. University of California Los Angeles

- Kohl, G. Siemens AG

- Lehner, O. Definiens AG

- Lu, J. Nat Inst Standards & Technology

- McNitt-Gray, M. University California Los Angeles

- Mozley, PD. Merck & Co Inc.

- Mulshine, JL. Rush

- Nicholson, D. Definiens AG

- O'Donnell, K. Toshiba

- O'Neal, M. Core Lab Partners

- Petrick, N. US Food and Drug Administration

- Reeves, A. Cornell University

- Richard, S. Duke University

- Rong, Y. Perceptive Informatics, Inc.

- Schwartz, LH. Columbia University

- Saiprasad, G. University of Maryland

- Samei, E. Duke University

- Siegel, E. University of Maryland

- Sullivan, DC. RSNA Science Advisor and Duke University

- Thorn, M. Siemens AG

- Yankellivitz, D. Mt. Sinai School of Medicine

- Yoshida, H. Harvard MGH

- Zhao, B. Columbia University

The Volumetric CT Technical Committee is deeply grateful for the support and technical assistance provided by the staff of the Radiological Society of North America.

## Appendix B: Background Information

Anatomic imaging using computed tomography (CT) has been historically used to assess tumor burden and to determine tumor response (or progression) to treatment based on uni-dimensional or bi-dimensional measurements. The original WHO response criteria were based on bidimensional measurements of the tumor and defined response as a decrease of the sum of the product of the longest perpendicular diameters of measured lesions by at least 50%. The rationale for using a 50% threshold value for definition of response were data evaluating the reproducibility of measurements of tumor size by palpation and on planar chest x-rays [23][24].The more recent RECIST criteria introduced by the National Cancer Institute (NCI) and the European Organisation for Research and Treatment of Cancer (EORTC) standardized imaging techniques for anatomic response assessment by specifying minimum size thresholds for measurable lesions and considered other imaging modalities beyond CT. As well, the RECIST criteria replace longest bidirectional diameters with longest unidimensional diameter as the representation of a measured lesion [8]. RECIST defines response as a 30% decrease of the largest diameter of the tumor. For a spherical lesion, this is equivalent to a 50% decrease of the product of two diameters. Meta-analyses combining the results of several large Phase II and Phase III studies have shown that tumor response according to WHO or RECIST criteria is correlated with patient survival for some tumor types [25]. However, there is considerable variability between individual studies, and the same response rate can be associated with completely different survival rates in different studies [26]. For some tumor types, meta-analyses found no or only a very weak correlation with patient survival [27][28]. Given the history of response criteria outlined above, these observations are probably not unexpected. Current response criteria were designed to ensure a standardized classification of tumor shrinkage after completion of therapy. They have not been developed on the basis of clinical trials correlating tumor shrinkage with patient outcome.

Technological advances in signal processing and the engineering of multi-detector row computed tomography (MDCT) devices have resulted in the ability to acquire high resolution images rapidly, resulting in volumetric scanning of anatomic regions in a single breath-hold. Volumetric image analysis (VIA) may be a substantially more sensitive technique for detecting longitudinal changes in tumor masses than reliance on lines representing tumor diameters as defined by RECIST. Comparative analyses in the context of real clinical trial data have found VIA more reliable and often more sensitive to longitudinal changes in response than the use of diameters in RECIST. As a result of this increased detection sensitivity and reliability, VIA may improve the predictability of clinical outcomes during therapy compared with RECIST. VIA could also benefit patients who need alternative treatments when their diseases stops responding to their current regimens.

The rationale for volumetric approaches to accessing assessing longitudinal changes in tumor burden is multi-factorial. First, most cancers may grow and regress irregularly in three dimensions. Measurements obtained in the axial plane fail to account for growth or regression in the longitudinal axis, whereas volumetric measurements incorporate changes in all dimensions. Secondly, changes in volume are less subject to either reader error or interscan variations. For example, partial response using the RECIST criteria requires a greater than 30% decrease in tumor diameter, which corresponds to greater than 50% reduction in volume of tumor. If one assumes a 21 mm diameter lesion (of 4850 mm3 volume), partial response would result require that the tumor shrink to a in a diameter of less than158 mm, but which would correspond to a decrease in volume all the way down to 17702145 mm3. The much greater magnitude of volumetric changes is less prone to measurement error than changes in diameter, particularly if the lesions are irregularly shaped or spiculated . As a result of the increased sensitivity and reproducibility, VIA may be more suited than uni-dimensional measurements to identify early changes in patients undergoing treatment.

The specific goal of RSNA/QIBA Quantitative CT Technical Committee in this effort is to produce alternative methods of response assessment, based on volumetric image acquisition and analysis, which will be accepted through appropriate regulatory pathways as predictors of clinical benefits, such as overall survival (OS). The working hypothesis is that volumetric imaging is an effective method for quantifying treatment-induced changes in tumor volume, and ultimately, changes in the health status of patients with cancer. The committee is conducting work to identify and evaluate methods of VIA in collaboration with FDA Division of Applied Math/ Office of Science and Engineering Laboratories (OSEL)/ Center for Devices and Radiological Health (CDRH), National Cancer Institute (NCI), National Institute of Standards and Technology (NIST), American College of Radiology Imaging Network (ACRIN), major imaging equipment manufacturers (such as Philips, GE, Siemens, and Toshiba), the Extended Pharmaceutical Research and Manufacturers of America (PhRMA) Imaging Group, scientists from academia (Columbia University, UCLA, etc.) and others.

## Appendix C: Conventions and Definitions

Acquisition vs. Analysis vs. Interpretation: This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be computational steps that transform the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge. (The transformation of knowledge into wisdom is beyond the scope of this document.)

Bulls-eye Compliance Levels Acquisition parameter values and some other requirements in this protocol are specified using a “bullseye” approach. Three rings are considered from widest to narrowest with the following semantics:

ACCEPTABLE: failing to meet this specification will result in data that is likely unacceptable for the intended use of this protocol.

TARGET: meeting this specification is considered to be achievable with reasonable effort and equipment and is expected to provide better results than meeting the ACCEPTABLE specification.

IDEAL: meeting this specification may require unusual effort or equipment, but is expected to provide better results than meeting the TARGET.

An ACCEPTABLE value will always be provided for each parameter. When there is no reason to expect better results (e.g. in terms of higher image quality, greater consistency, lower dose, etc.), TARGET and IDEAL values are not provided.

Some protocols may need sites that perform at higher compliance levels do so consistently, so sites may be requested to declare their “level of compliance”. If a site declares they will operate at the TARGET level, they must achieve the TARGET specification whenever it is provided and the ACCEPTABLE specification when a TARGET specification is not provided. Similarly, if they declare IDEAL, they must achieve the IDEAL specification whenever it is provided, the TARGET specification where no IDEAL level is specified, and the ACCEPTABLE level for the rest.

Other Definitions:

Anonymization is the process of de-identification and further removal or ambiguation of information to reduce the probability of re-identification of the image despite access to other information sources

Adjudication is the processes of decision making that involves an independent party with the authority to determine a binding resolution through a prespecified process. In the standard read design once a primary analysis of all time point data for each patient has been completed, the findings are compared in order to identify any discrepancies in overall assessments of Best Overall Response, Date of Progression, Date of Response and Date of Response Confirmation. So as to ascertain the final determination for discrepant cases, a third Reviewer will review the patient data and choose to agree with the findings of one of the two prior analyses. The third Reviewer evaluates the same set of images used by Reviewers 1 and 2 and will not have seen the images before. A variation of this read design is to have a third Reviewer who is blinded to the previous findings and reviews the cases in exactly the same fashion as the initial Reviewers. If agreement is not reached in three separate analyses, a consensus of Reviewers is required in order to verify the final determination for that patient. Read designs will be outlined in detail in the Independent Review Charter in advance of eCRF design, or Adjudication is a means of deciding an outcome where two different opinions are formed, or Generally referring to a blinded independent read designed to resolve discrepancies between two primary readers.

The Adjudication Rate is the number of cases that are adjudicated divided by the total number of cases evaluated.

Baseline Followed by Randomized Temporal Image Presentation is the sequence of image presentation such that the baseline (earliest) time point is shown to the reviewer for the purpose of identifying regions of interest, such as selecting neoplastic masses as target lesions. Subsequent time points are presented in a random order with respect to the date.

Blinding is a procedure in which one or more parties to the trial are kept unaware of the treatment assignments and other information that might introduce bias. Single blinding usually refers to the subjects being unaware, and double-blinding usually refers to the subjects, investigators, monitor, and, in some cases, data analysts being unaware of the treatment assignments, or Blinding is the outcome of all processes and procedures that are deployed to prevent image analysis operators or reviewers from becoming aware of or informed about the circumstances surrounding a case, such as the information in the following abbreviated list: investigational (test diagnostic or test therapeutic) arm (or any data that might reveal the investigational arm); subject demographics; site assessments (including site choice of lesions); situational specific descriptions of the scans (such as “confirmation” or “end of treatment scans”); results or assessments of other reviewers participating in the reading process (except during some adjudication scenarios); clinical data other than that which has been pre-specified described in the imaging charter.

A Blinded Read is the analysis of images to determine results of the testing in which the radiologist is unaware of any subject or site information.

Burned-in Information is information that is part of the actual pixel data as opposed to present in the image header.

Comment: A comment in this instance is generally referring to a text field that can capture additional reader insight into the read process or reader thought processes. Comments are generally required when the reader indicates an image in Not Evaluable or their opinion differs from the derived response.

Clinical Read or Site Read: Generally referring to an independent read that combines the result of the imaging review with pre-defined clinical information to come to a final outcome associated with the efficacy endpoint. Readers are generally blind to treatment groups but may be provided a variety of clinical information appropriate to the disease assessment. Or Image interpretation done at the investigational site

Computer Generated Quantitative Image Analysis is an analysis performed automatically by a computer with little or no human interaction using signal processing algorithms to quantify an imaging outcome measure. This type of analysis should be deterministic (always produce identical output from the same input) or have low variability. A Confirmation Read is generally referring to a central read that occurs based on an “on-site” event. Confirmation reads are associated with eligibility criteria, disease progression or other events that may benefit from a third party confirmation.

Data Lock: The Data Lock is the point and method when the results of an assessment or imaging outcome variable are considered final and are protected. This must be pre-defined in the analysis. Locking must not be construed to mean an assessment cannot be overturned as indicated by emerging data as long as: (1) the process is pre-defined in the Imaging Charter; (2) the process is driven by data that, by design, emerges after the initial assessment; and (3) there are adequate audit trails that can substantiate the changes.

De-identification is the process of removing real patient identifiers or the removal of all subject demographics from imaging data for anonymization  De-personalization is the process of completely removing any subject-related information from an image, including clinical trial identifiers.

A Derived Response is an outcome measure algorithmically derived based on information from the blinded reader assessments.

In the End of Review Data Lock, the data are locked when the reviews of all the time points for the subject have been completed.

Evaluable: Generally referring to image quality. Based on presence or absence of necessary imaging and the associate image quality. Evaluable is the response generated when an image and/or time point can be interpreted.

In the Exam Level Data Lock, the data are "locked" in "final form" after each Exam (medical imaging procedure) is assessed. The purpose of the Exam Lock is to assess the differential contribution of each Exam to the overall assessment.

Human Interfaced Image Analysis is image analysis that is driven primarily by a human reviewer who may use computer generated analysis tools to quantify an imaging outcome measure.

Hybrid Randomized Image Presentation. In this paradigm, the first stage of the assessment is fully randomized or the post-baseline scans are randomized. Once the results have been locked for each time point, the images are re-presented in known chronological order for reconsideration. Changes in any of the randomized assessments are tracked and highlighted in the final assessment. In within-patient-control trials (e.g. comparative imaging) images obtained before and after the test agent should be presented in randomized un-paired fashion. The minimum size of the randomization block necessary to minimize recall should be considered.

Image Analysis, Image Review, and/or Read: Procedures and processes that culminate in the generation of imaging outcome measures, such as brain volume, cardiac output, or tumor response criteria. Reviews can be performed for eligibility, safety or efficacy. The review paradigm may be context specific and dependent on the specific aims of a trial, the imaging technologies in play, and the stage of drug development, among other parameters.

Imaging Data are variables derived from the image review or, Imaging Data are quantitative or qualitative variables resulting from image review. Such variables may be used to assess eligibility for study and treatment response, or Imaging Data is information that results from or is produced by the image analysis or review processes (such as lesion selection and their associated spatial measurements), or from algorithmically derived assessments specified in the protocol. In this context, the term also refers to "marks" placed on images, such as regions of interest boundaries, annotations such as "Target Lesion 4", etc.

Imaging Endpoint: Endpoint based on objective image features chosen to evaluate the activity of a study treatment (e.g. retardation of joint destruction in patients with rheumatoid arthritis)

Imaging Examination or Imaging Exam or Exam: An Examination or Exam is a single set of intimately related images acquired contemporaneously with a single technology, such as a CT scan of the chest, a whole body bone scintigram, or an echocardiogram.

Image Header: The Image Header is that part of the file or dataset containing the image other than the pixel data itself

Imaging Case Report Forms are IRC-specific forms designed to capture elements of image acquisition, and/or image interpretation and/or derived responses for enrollment and/or eligibility review and/or confirmation reads and/or efficacy assessment.

Imaging Phantoms: Devices used for periodic testing and standardization of image acquisition. This testing must be site specific and equipment specific and conducted prior to the beginning of a trial (baseline), periodically during the trial and at the end of the trial.,

Image Review Plan: The Image Review Plan or Radiology Technical Manual is a that document that summarizes the plan for the acquisition of imaging data. Imaging endpoint that is correlated with a clinical outcome but is not sufficient to show clinical benefit Use definitions in GFI Developing Med Imaging Drug Part 3 2004

Independent Review Charter (IRC): The Image Review Charter is a document submitted to a regulatory agency to document and support the use of imaging to support an IND, NDA or BLA. The document identifies and lists imaging resources, imaging surrogate criteria, processes for receipt, handling, preparation and archive of images, the process steps for review and assessment of images and the various methodologies for quality assurance and quality control, or The Image Review Charter is a detailed technical document governing the acquisition, processing, display, interpretation, transfer, and integrity of imaging data in efficacy trials that use imaging for assessment of efficacy outcomes. IRCs are an integral component of the clinical trial protocol and promote quality and verifiability of imaging data. The IRC is designed for use by the clinical investigators, by the central image laboratory and by regulatory agencies. The IRC might be submitted to a regulatory agency for review and comment and to reach agreement on the use of imaging in trials intended to support an NDA or BLA. The IRC contains a summary of the clinical protocol including study design, study population, efficacy endpoints and primary efficacy analysis. The IRC describes imaging resources, processes for receipt, handling, preparation and archive of images, process steps for review and assessment of images and various methodologies for quality assurance and quality control. The IRC needs to cross reference relevant portions of the clinical protocol including enrollment criteria, outcome measures, and statistical analysis plan (including primary efficacy analysis, procedures for handling missing or uninterpretable data etc.) For ease of regulatory review the IRC should include a protocol synopsis that is sufficiently detailed to permit verification of consistency of the IRC with the clinical protocol and statistical analysis plan.

Individually Identifiable Information is data that alone or in combination may be used to identify an individual.

Interpretable: Generally referring to image quality assessed by the blinded reader. Based on presence or absence of necessary imaging and the associate image quality. Grounds for the assessment are commonly captured. For example, not Optimal but Evaluable is the response generated when an image and/or time point is of questionable quality but can be interpreted.

Intra-Observer Variability or Intra-Reader Variability is the variability in the interpretation of a set of images by the same reader after an adequate period of time inserted to reduce recall bias.

Inter-Observer Variability or Inter-Reader Variability is the variability in the interpretation of a set of images by the different readers.

Not evaluable: Generally referring to image quality. Based on presence or absence of necessary imaging and the associate image quality. Not Evaluable is the response generated when an image and/or time point cannot be interpreted. May be assessed in real time by a blinded third party quality assessor independently of image reader. Provision for reimaging (where feasible) should be prespecified. Listing of criteria is provided and responses are captured in the CRF.

In the “N” Time Point Data Lock, a variable number of time points “N” can be combined and shown together at a particular stage of the review process. For example, the baseline/screening and the first subsequent time point after baseline/screening may be reviewed together to establish the baseline extent of disease.

Off-Protocol imaging is imaging that may have been performed during a trial and should not be reviewed by the IRC or Imaging which is done during a trial but not required by the protocol.

On-Protocol imaging is imaging that is performed during a trial as required by and defined in the protocol, or On-Protocol imaging is imaging that is performed during a trial as required by and defined in the protocol that should be reviewed by the IRC.

The Order of Image Presentation is the sequence that images are presented to reviewers for formal review and generation of the imaging outcome measures. Sufficiently well established designs include:  Personal information is data related to person identification - see EU guidance (e.g., Age)

A Primary Read is the blinded review of imaging data in which one or more independent reviewers review images to generate the outcome information associated with the efficacy endpoint, or A Primary Review is the blinded review of data in which one or more independent reviewers review the data to generate the outcome information associated with the efficacy endpoint.

Pseudonymization is the process of de-identification and replacement of identifiers with a pseudonym that is unique to the individual and known within the context of a trial but not linked to the individual in the external world.

Randomized Independent Temporal Image Presentation is the sequence of image presentation that each time point is presented alone, in a random order with respect to the date of acquisition, and reviewed independently without access to other time points.

Reader Independence: Academic, financial, trial conduct

Scheduled imaging is imaging that is performed during a trial at one or more of the time points (or window assigned to a time point) designated for imaging assessment in the protocol. Applies to either on-protocol or off-protocol imaging, or Exams that are scheduled as routine assessments.

A Secondary Read is a blinded review of imaging data in which one or more independent reviewers review images to generate outcome data that is not part of the efficacy endpoints. An example would be a read that is part of Intra-reader analysis.

Sensitive Personal Information is data related to personal preferences and disposition. - see EU guidance (e.g., Ethnicity).

Sequential Chronologic Image Presentation is the sequence of image presentation that images are shown to the reviewer in the order in which they were actually acquired. In this format, the reviewer should not know the total number of time points to be assessed unless that information has been pre-specified in the imaging charter. (For example, pre-specification is usual and customary in imaging studies of neurodegenerative disorders, arthritis, osteoporosis, and congestive heart failure, among others.)

Simultaneous Chronological Image Presentation is the sequence of image presentation that all images associated with a subject are shown to the reviewer at the same time without blinding the date or sequence or total number.

Simultaneous Randomized Temporal Image Presentation is the sequence of image presentation that all images associated with a subject are shown to the reviewer at the same time in a random order with respect to the date but without blinding to total number.

Sequential Unblinding Sequential interpretation of images with and without clinical information (e.g. demography, clinical assessments).

Statistical Analysis Plan for Medical Imaging: Analysis plan focused on primary efficacy analysis and including statement of null hypothesis, study power, statistical test, efficacy population, and handling of missing or uninterpretable images, sensitivity analyses

A Time Point is a discrete period during the course of a clinical trial when groups of imaging exams or clinical exams are scheduled as defined in the study protocol.

In the Time Point Data Lock, the data are locked after all of the pre-specified information associated with each time point is assessed. In some paradigms, the time points are known to be presented in chronological order; in others, the time points may be randomized during the early stages of the image analysis process (vida infra).

Truth Standard: Single or multiple image modalities

Unique Identifiers (UIDs) are globally unique identifier used to identifiers images, sets of images, or components within an image.

Uninterpretable: Generally referring to image quality assessed by the blinded reader.

Unscheduled imaging is imaging that is performed during a trial at a time/date outside the window assigned to a time point designated for imaging assessment in the protocol. It may be ad hoc imaging performed to evaluate an unscheduled clinical circumstance. It may be on-protocol or off-protocol imaging.

## Appendix D: Documents included in the imaging protocol (e.g., CRFs)

**(Material the site needs to submit)

**Subject preparation

**Imaging agent dose calculation

**Imaging agent

**Image data acquisition

**Inherent image data reconstruction / processing

**Image analysis

**Interpretation

**Site selection and Quality Control

**Phantom Imaging and Calibration

## Appendix E: Associated Documents

**e.g. the Imaging Charter, Site Manual, Standard Operating Procedures, etc.

## Appendix F: TBD

## Appendix G: Model-specific Instructions and Parameters

The presence of specific product models/versions in the following tables should not be taken to imply that those products are fully compliant with the QIBA Profile. Compliance with a profile involves meeting a variety of requirements of which operating by these parameters is just one. To determine if a product (and a specific model/version of that product) is compliant, please refer to the QIBA Conformance Document for that product. G.1. Image Acquisition Parameters The following technique tables list acquisition parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

These technique tables may have been prepared by the submitter of this imaging protocol document, the clinical trial organizer, the vendor of the equipment, and/or some other source. (Consequently, a given model/version may appear in more than one table.) The source is listed at the top of each table.

Sites using models listed here are encouraged to consider using these parameters for both simplicity and consistency. Sites using models not listed here may be able to devise their own acquisition parameters that result in data meeting the requirements of Section 7.1 and conform to the considerations in Section 13.

In some cases, parameter sets may be available as an electronic file for direct implementation on the imaging platform.

Table G.1a

Generic: This represents parameters for a generic CT. Source: QIBA v-CT Cmte  Date: 2009-mm-dd

| Parameter | Compliance Level* | Generic |  |
| --- | --- | --- | --- |
| kVp | Acceptable | 110 to 140 |  |
|  | Target | 110 to 130 |  |
| mAs (medium patient) | Acceptable | 40 to 350 |  |
|  | Target | 80 to 160 |  |
| Scan Duration | Acceptable | < 30 sec. |  |
|  | Target | < 15 sec. |  |
|  | Ideal | < 10 sec. |  |
| Table Speed | Acceptable |  |  |
|  | Target |  |  |

* See Appendix C for a discussion of the Levels of Compliance

kVp and mAs should be adjusted as necessary, depending on the body habitus of individual patients. The values should be consistent for all scans of the same patient.

Scan Duration values are intended to allow completion of the scan in a single breath hold for most/a majority/nearly all subjects respectively.

Table Speed values are intended to yield an IEC Pitch Value of approximately 1 while achieving the corresponding Scan Duration.

Table G.1b: "Target" Compliant Protocols for Specific Systems

The following table provides sample parameters sets that meet the “Target” Level of Compliance for specific models.

See Appendix C for a discussion of the Levels of Compliance.

| Parameters | vCT 1A (Philips) | GE | ACRIN |
| --- | --- | --- | --- |
|  | MxIDT 8000 (Thin) | MxIDT 8000 (Thick) | Ultra | VCT-64 | 6678 |
| Data Content |  |  |  |  |  |
| Anatomic Coverage |  |  |  |  |  |
| Field of View : Pixel Size |  |  |  |  | Rib-to-rib: 0.55-.75mm |
| Data Structure |  |  |  |  |  |
| Collimation Width | 16x0.75 mm | 16x1.5 mm |  |  | (TBA) |
| Slice Interval |  |  |  |  |  |
| Slice Width | 0.8 mm | 5.0 mm |  |  | 1.0 mm |
| Pixel Size |  |  |  |  | 0.55 mm |
| Isotropic Voxels |  |  |  |  | (2:1) |
| Scan Plane |  |  |  |  |  |
| Rotation Speed |  |  |  |  | 0.5 sec |
| Data Quality |  |  |  |  |  |
| Motion Artifact |  |  |  |  |  |
| Noise Metric |  |  |  |  |  |
| Spatial Resolution Metric |  |  |  |  |  |
| Acquisition |  |  |  |  |  |
| Tube Voltage | 120 kVp | 120 kVp |  |  | 120 kVp |
| Exposure | 100 mAs | 100 mAs |  |  | 100 mAs |
| Pitch | 1.2 | 1.2 |  |  |  |
| Reconstruction |  |  |  |  |  |
| Recon. Kernel | Detailed filter | Detailed filter |  |  | Standard |
| Recon. Interval |  |  |  |  |  |
| Recon. Overlap | 50% | 50% |  |  | 20% |

G.2. Image Reconstruction Parameters

See above.

G.3. Post-Processing Instructions

None provided.

G.4. Analysis Instructions

None provided.

G.5. Interpretation Instructions

None provided.

# Text boxes

patients candidate for neoadjuvant window of opportunity trials

primary and metastatic lung cancer

to assess therapeutic response to cytotoxic and selected targeted therapies (e.g., antiangiogenic agents, tyrosine kinase inhibitors, etc.)

**target subpopulation

**biology

**biomarker use