## Sheet1
QIBA 1C Performance Based Protocol Worksheet
Step 1 - scanner desiption
Scanner Manufacturer | GE
Scanner Model | Discovery 690 (CT part is the same as VCT64)
Site id for scanner | 919684PET4A
Software version | kh_xar.30.V40_H_H64_G_GTLNI:9x6_lyso
Date of Testing | 2010-01-09
Steps 2-4: Scanning Parameters that Meet Performance Metrics
Step 2: kVp selected | 120
Step 3: Scan Field of View (SFOV) or Patient Size | Small Body (32cm)
Step 4: Reconstructed Slice with Selected (in mm) | 1.25
Step 5: Collimation Rotation Time and Pitch
Number of Data Channels (N) | 64
Width of Each Data Channel (T, in mm) | 0.625
Gantry Rotation Time in seconds | 1
Pitch (close to 1.0 is preferred) | 0.984:1
Time to cover a 35 cm long thorax (desired < 15 seconds) | #VALUE!
Step 6: Iterating on Recon Kernels to obtain 7 lp/cm
If using mA, then mA setting | 250
(if using mA, then effective mAs equals this value): | #VALUE!
if using mAs/rotation (philips) or effective mAs (Siemens), the record setting here
Recontruction Filter setting (name) - Iteration 1 | bone
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Recontruction Filter setting (name) - Iteration 2 | detail
Size of line pairs seen with this filter (Level =1100, Win=100) | 7
Recontruction Filter setting (name) - Iteration 3 | std
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Name of Filter that meets 7 lp/cm performance criteria | detail
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter
Step 7: Iterating on mA (or mAs) settings to obtain std dev of 12 +/1-1 HU
mA or mAs or Eff. mAs setting - Iteration 1 | 250
Standard deviation measured on image of ACR phantom | 14.22
mA or mAs or Eff. mAs setting - Iteration 2 | 325
Standard deviation measured on image of ACR phantom | 12.18
mA or mAs or Eff. mAs setting - Iteration 3
Standard deviation measured on image of ACR phantom
mA or eff. mAs setting that meets Standard Deviation Performance criteria | 325
Scanner reported CTDIvol (in mGy) using this mA or eff. mAs setting | 24.19
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter