## Sheet1
This spreadsheet provides Comments and Resolutions to the QIBA SLN Profile as decided by the QIBA CT Small Lung Nodule Biomarker Committtee
Document Date: | 2018-02-09
| Date Received | Submitted By | Comment Received | Profile Section | For Document Version | Status | QIBA SLN BC Resolution
51 | 2018-06-14 | Rick Avila | The SLN Profile needs a description for how to calculate the resolution aspect ratio. | 4.1.1 | 2018-06-15 | Closed | A description for calculating the resolution aspect ratio was added to section 4.1.1
50 | 2018-06-14 | Rick Avila | The pitch requirement fails for multi-detector CT systems and provides little additional value beyond the fundamental image quality metrics being measured. Consider removing pitch as a requirement. | 3.4.2 | 2018-06-15 | Closed | Pitch <= 2.0 has been removed as a requirement from section 3.4.2. A recommendation to maintin pitch <= 2.0 remains in the discussion section of the Profile.
49 | 2018-06-14 | Rick Avila | The measurement of fundamental image quality properties is performed at each of three phantom modules, the furthest being 200mm from scanner iso-center. However, even the largest human lungs extend no more than 175mm from iso-center. Consider checking fundamanetal image quality properties no further than 175 mm. | 4.1.1 | 2018-06-15 | Closed | The Profile has been modified to require that the two inner modules must pass all CT image quality requirements plus a third position at 175mm must also pass. The 175mm position can be calculated using linear interpolation between modules.
48 | 2018-06-14 | Rick Avila | Some sites want to scan smaller sized patients using model-specific scanning modes with a small field of view. In this instance, consider having sites provide two sets of conformance data, one for small sized patients and another for larger patients. | 4.1.1 | 2018-06-15 | Closed | The Profile has been modified to support the use of multiple protocols for different size patients. As long as all protocols pass the requirements the scanner and the set of protocols should pass.
47 | 2017-12-08 | Sean Fain | Initially reacted negatively to the lack of a definitive claim (i.e. claims that depend on nodule size). Assume that this information would be used to decide which nodules would be beyond the resolution of a given study and could then be adjusted according to the demands of the study protocol. I am authoring a profile for the lung density biomarker committee and found Section 4 to be particularly clear. I plan to model my approach more after this section. Nice work. | All | 2017-11-15 | Closed | You are correct. Solid lung nodule volume change performance is highly dependent on the size of the nodule and the fundamental properties of the image acquisition system.
46 | 2017-12-08 | Sean Fain | I don't know if this will overwrite my previous comment, but I was reviewing Section 4 and found that the references don't seem to line up. Reference 53 did not seem to report on the MTF50 and corresponding in-plane standard deviation of the PSF as cited. May want to followup and check references. | 4 | 2017-11-15 | Closed | The references have been reformatted and are now corrected.
45 | 2017-12-05 | Ehsan Samei | The proposed method for conformance does not meet current standards of CT quality nor proven to reflect volumetric characterization in clinical cases. | All | 2017-11-15 | Closed | We are not aware of a standard of CT quality for solid lung nodule measurement. The SLN BC has reviewed the supporting methods and study data and has determined there is sufficient evidence to move forward.
44 | 2017-11-20 | James Mulshine | This profile addresses a very challenging new clinical management and contains a number of innovative refinements. Close attention to emerging data through this process will be essential to guide further development of subsequent versions of this profile. Continuing the QIBA process improvement efforts will enable even more robust support of screening managment decision-making with the result in more people being cured of early lung cancer. | All | 2017-11-15 | Closed | The SLN BC committee is committed to improving the profile through evidence-based methods.
43 | 2017-11-20 | Kevin O'Donnell | Most of the comments were rejected, which is not a priori bad, but for two thirds of them the rationale seemed to be generic copy/pasted text that didn’t address the comment directly and sometimes not at all. Also the text references an October 30 document I couldn’t find, but I may be looking in the wrong place. | All | 2017-11-15 | Closed | The reason for the multiple entries with the same resolution text was that the povided answers were identical. In many of these cases the submitted issue had already been discussed, reviewed and documented many times in the past so comment submitters were asked to review the previously provided answers to these submissions. The Oct 30 document will be posted.
42 | 2017-11-20 | Kevin O'Donnell | The editing of section 3.1.2 got messed up somehow and there are now randomly replicated sentences in half the rows. | 3 | 2017-11-15 | Closed | This has been corrected.
41 | 2017-11-20 | Kevin O'Donnell | The Profiel text is very clear that claims only hold for nodules between 6 and 10mm. The CV table extends to 12mm to support growth above 10mm. The SLN Biomarker Committee reviewed the Profile and decided that the current Profile organization and text is appropriate." On reviewing the executive summary, the clinical context, the claim and the image QA requirements, it seems likely that the opposite is true; the limit seems to be 12mm. The only place that says 10mm is 4.3.2 1b where the requirement was replicated inside the procedures. | 4.3 | 2017-11-15 | Closed | Committee decided to change upper limit of most text to 10 mm and keep 12 mm as upper limit of table.
40 | 2017-11-20 | Kevin O'Donnell | Thes "x.y.org" links throughout the document were modified to point to the QIBA SLN Wiki which has links to the referenced resources (all 12 occurrences were changed).” Having broken links on the wiki instead of in the Profile misses the point that broken links are not ready for field testing. The CT Vol Profile made this mistake. We should avoid repeating it. | All | 2017-11-15 | Open | The SLN BC is working to prepare all materials for field testing.
39 | 2017-11-20 | Kevin O'Donnell | After fixing comments 27 & 28, the last two rows of 3.7.2 are now reworded duplicates and rows 1, 4 and 5 of 3.10.2 also seem to be duplicates. | 3.7 & 3.10 | 2017-11-15 | Closed | The duplicate entries in 3.7 have been removed.
38 | 2017-11-20 | Kevin O'Donnell | After fixing comment 33, the Warping procedure says to use Mean Squared Error and the requirement says to use Root Mean Squared Error. | 3.4.2 & 4.1 | 2017-11-15 | Closed | This has been corrected.
37 | 2017-11-13 | Nancy Obuchowski | ​Lines 164, 167, and 170: change "probability" to "confidence" Lines 918 and 931: add "with 95% confidence" | 2 and 4 | 2017-10-31 | Closed | The Profile was modified as suggested.
36 | 2017-10-26 | Kevin O'Donnell | For QIBA Profiles, conformance consists of: - Confirm each requirement assigned to you in Section 3. Some requirements reference a specific assessment procedure. A checklist of the requirements re-sorted by actor is handy.(But keep the text the same to avoid differing interpretations) | 4 | 2017-09-27 | Closed | The SLN Biomarker Committee considers the current Profile version more intuitive for Profile users. This may be considered for a future version of the Profile.
35 | 2017-10-26 | Kevin O'Donnell | This procedure states that completing it makes a site conformant despite not checking (or meeting) one or two dozen of the requirements in Section 3. Procedures for assessing individual requirements are not the same as procedures for assessing overall conformance to the Profile. If it would help, we can revise the template to rename section 4. from Assessment Procedures to Requirement Assessment Procedures, and rename section 5. from Conformance to Conformance Procedures. But intermingling them is a bit confusing. | 4.3 | 2017-09-27 | Closed | The SLN Biomarker Committee reviewed the Profile and found that all requirements are verified with the stated conformance procedures. The SLN Biomarker Committee decided that the current Profile organization and text is preferred.
34 | 2017-10-26 | Kevin O'Donnell | The current SLN draft diverges further away from the template: omitting the Conformance section, renaming the Assessment Procedures section to Conformance, and requiring that multiple assessments can only be performed as a group. If the goal is to highlight that automated software is available that can do multiple assessments at once, that can be done without locking out equivalent alternatives. | 4.1 | 2017-09-27 | Closed | The SLN Biomarker Committee examined this area and found that the current Profile has a Conformance Section 4 in keeping with QIBA conventions and that their is no requirement to do multiple assessments as a group. The SLN Biomarker Committee decided that the current Profile organization and text is preferred.
33 | 2017-10-26 | Kevin O'Donnell | The requirements on Resolution, Edge Enhancement, HU Deviation, Voxel Noise and Spatial Warping do not reference any specific assessment procedure implying sites may use whatever procedure they find reasonable. Since specific assessment procedures for these 5 requirements have been written in section 4, there should be a reference here, e.g. (See 4.X Assessment Procedure: Resolution) | 3.4.2 | 2017-09-27 | Closed | This has been addressed and fixed in the November 15, 2017 version of the SLN Profile.
32 | 2017-10-26 | Kevin O'Donnell | The Large Nodule profile dropped a lot of requirements when discussing and evaluating the comments from the field tests. See: https://docs.google.com/spreadsheets/d/1cyz0frvELb6Tq1EHjpHxnvjeUxotExnJ3OrKPdIa30U/edit?usp=sharing There is a fair chance the same arguments hold for SLN. Consider dropping them to simplify the profile. | 3 | 2017-09-27 | Closed | The SLN Biomarker Committee decided that the current Profile organization and text is preferred. This area can be considered in future versions of the Profile.
31 | 2017-10-26 | Kevin O'Donnell | The requirement that the Image Analysis Tool "Shall have appropriate scientific validation, including the properties of measurement linearity, coefficient of variation, and zero bias." may leave implementers unsure of exactly what they should do to conform. Replace with specific requirements on linearity, bias, etc, and consider specific assessment procedures for each. | 3.1.2 | 2017-09-27 | Closed | The SLN Biomarker Committee decided that the current Profile organization and text is preferred. This area can be considered in future versions of the Profile.
30 | 2017-10-26 | Kevin O'Donnell | Table 1 has been modified to state points not ranges (i.e. 6mm rather than 6-7mm). It is not clear how to derive CV values from the table. Should implementers take their measured diameter and round up, round to the nearest, or interpolate? Clarify. | 2 | 2017-09-27 | Closed | The text was modified in the November 14, 2017 version of the SLN Profile to specify using interpolation.
29 | 2017-10-26 | Kevin O'Donnell | The profile is inconsistent about whether the size range is 6-10mm or 6-12mm. Decide and revise. | 4.3.2 | 2017-09-27 | Closed | The Profiel text is very clear that claims only hold for nodules between 6 and 10mm. The CV table extends to 12mm to support growth above 10mm. The SLN Biomarker Committee reviewed the Profile and decided that the current Profile organization and text is appropriate.
28 | 2017-10-26 | Kevin O'Donnell | Throughout section 4 there are references to website x.y.org. The idea of the Technical Confirmation stage is to provide the Consensus Profile to the test sites and have them execute it. The Large Nodule Profile did not fully set up the test datasets before publishing the Consensus draft and that resulted in confusion during the Field Test. Resolve the references (including instructions on how to navigate to the referenced dataset or other information) and make sure they work. | 4 | 2017-09-27 | Closed | Thes "x.y.org" links throughout the document were modified to point to the QIBA SLN Wiki which has links to the referenced resources (all 12 occurrences were changed).
27 | 2017-09-08 | Kevin O'Donnell | Placing requirements on noise in 3 different locations seems to be more than is needed to meet the claim. Consider reducing to one. It is thought useful to measure it at three distances from isocenter. | 3.4.2 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
26 | 2017-09-08 | Kevin O'Donnell | The assessment procedure should not require noise measurements at locations beyond the requirements (but it can permit them). Consider reducing to one. | 4.1.1 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
25 | 2017-09-08 | Kevin O'Donnell | Placing requirements on HU Bias in 3 different locations seems to be more than is needed to meet the claim. Consider reducing to one. It is thought useful to measure it at three distances from isocenter. | 3.4.2 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
24 | 2017-09-08 | Kevin O'Donnell | The assessment procedure should not require HU Bias measurements at locations beyond the requirements (but it can permit them). An assessment procedure for measuring the HU Bias in 3 locations in a conventional phantom can also be added (although it might be tricky to get both air and acrylic at three distances?) | 4.1.1 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
23 | 2017-09-08 | Kevin O'Donnell | The assessment procedure should not require Edge Enhancement measurements at locations beyond the requirements (but it can permit them). Discussion about whether it must be evaluated at three distances or only one which could be off isocenter. | 4.1.1 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
22 | 2017-09-08 | Kevin O'Donnell | The Spatial Warping metric should probably be renamed to avoid confusion with MR? | 4 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
21 | 2017-09-11 | Mike McNitt-Gray | This is a non-standard phantom using non-standard methods and even non-standard metrics. That is ok, but to be widely accepted (and widely imposed as part of compliance) it does require two major steps: (a) some explanation and demonstration of the underlying methods and what the results are actually measuring. This could be done by comparison to existing methods or use some other form of demonstration; and (b) some evidence that these have some effect on nodule volumetric measurements. | 4 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
20 | 2017-09-11 | Mike McNitt-Gray | Please explain how the noise requirements in the three additional materials would be expected to improve overall volume quantification performance (e.g. segmentation performance and beyond) compared to protocols constrained by only a single material noise requirement. | 4 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
19 | 2017-09-11 | Mike McNitt-Gray | Please explain how the noise requirements in the three additional materials would be expected to improve overall volume quantification performance (e.g. segmentation performance and beyond) compared to protocols constrained by only a single material noise requirement. Placing requirements on noise in 4 different materials seems to be more than is needed to meet the claim. Consider reducing to one or two. Agree that likely don't need noise in 4 materials to meet claim. One requirement on one nodule-like material (e.g. acrylic) is adequate. | 4 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
18 | 2017-09-11 | Mike McNitt-Gray | It would be interesting to collect data about addiitional materials (particularly if there is basically no extra cost to do so) | 4.1.1 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
17 | 2017-09-11 | Mike McNitt-Gray | Please explain how the noise measured in essentially a non-attenuating environment (current phantom) would be expected to be representative of noise in a patient attenuation environment. | 4.1.1 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
16 | 2017-09-11 | Mike McNitt-Gray | What is the justification for reflecting resolution in terms of PSF as opposed to the MTF50. If equivalent, consider using MTF50 in the requirement since it is more familiar to medical physicists and would be comparable to other profiles. PSF is a common resolution metric in image processing. Want to try using it in medical imaging. The resolution requirement may be specified using the PSF metric but should note the corresponding f50 value for comparability to the large nodule spec. The assessment procedure will specify producing the PSF and will note the conversion equation. | 3.4.2 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
15 | 2017-09-11 | Mike McNitt-Gray | What is the justification for use of a 3D metric and an in-plane-vs-z-axis ratio | 3.4.2 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
14 | 2017-09-11 | Mike McNitt-Gray | Please explain how requirements on voxel bias would be expected to improve volume quantification (e.g. segmentation performance and beyond) compared to protocols where only noise and resolution were constrained. Does not seem necessary to meet the claim. Consider removing. Voxel bias would affect threshold based segmenters which are in common use. HU Bias should be required in two materials (air, acrylic). It would be interesting to collect data about addiitional materials (particularly if there is basically no extra cost to do so). The assessment procedure can describe but not require producing and collecting bias measurements in other materials. | 3.4.2 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
13 | 2017-09-11 | Mike McNitt-Gray | Please explain how requirements on edge enhancement would be expected to improve volume quantification (e.g. segmentation performance and beyond) compared to protocols where only noise and resolution were constrained. Kernel edge enhancement is non-isotropic (doesn’t enhance Z). So the performance of edge finding segmenters will depend on tumor orientation and their performance will degrade with increased edge enhancement. The requirement evaluates an air/acrylic edge. Other materials/edges are not considered. | 3.4.2 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
12 | 2017-09-11 | Mike McNitt-Gray | Please explain how requirements on spatial warping would be expected to improve volume quantification (e.g. segmentation performance and beyond) compared to protocols where only noise and resolution were constrained. Images shown from one (two?) scanners with significant spatial distortion of tape rolls located far from isocenter. Q. What did the resolution metrics for those scans look like? Please demonstrate the effects of these parameters on volume quantification under a variety of acquisition and reconstruction conditions (eg, pitch, recon filters, slice thickness, dose) using current state-of-the-art scanners. | 3.4.2 | 2017-08-28 | Closed | In addition to the Profile text, explanation of the current new phantom and conformance approach was given in a 15 page QIBA SLN TF response letter dated October 7, 2017. A 2 page response to further questions was also provided on October 30, 2017.
11 | 2017-09-05 | Kevin O'Donnell | “Technologist shall use the same compliant scanner and acquisition protocol for acquisition of all time points” is an Image Acquisition requirement, not product validation. | 3.1.2 | 2017-05-15 | Closed | Specification moved to 3.7.2
10 | 2017-09-05 | Kevin O'Donnell | “Image Analyst shall analyze all time points and shall use the same conformant image analysis tool at all analysis time points.” is an Image Analsis requirement, not product validation. | 3.1.2 | 2017-05-15 | Closed | Spcification moved to 3.10.2
9 | 2017-09-05 | Kevin O'Donnell | A Design of Experiments approach for acquisition device validation by vendors is described but not made a requirement. | 3.1.2 | 2017-05-15 | Closed | Specification added.
8 | 2017-09-05 | Kevin O'Donnell | The Radiologist has been made responsible for performance of the phantom scans. | 3.4.2 | 2017-05-15 | Closed | Technologist added to all specifications as they help radiologists achieve scanning goals. Physicst was added to all phantom scans.
7 | 2017-09-05 | Kevin O'Donnell | Some of the requirements must use specific assessment procedures but they don't say so. | 3.4.2 | 2017-05-15 | Closed | "See section 4.1. Assessment Procedure: Image Quality" was added to all relevant specifications.
6 | 2017-09-05 | Kevin O'Donnell | 3D Spatial Warping is a bit redundant. | 3.4.2 | 2017-05-15 | Closed | Parameter name was changed to "Spatial Warping"
5 | 2017-09-05 | Kevin O'Donnell | The spatial warping is required to be "less than Z MSE". | 3.4.2 | 2017-05-15 | Closed | An initial value of 0.5 mm was set for this specification but review of new scan data of phantom resulted in changing this limit down to 0.3 mm.
4 | 2017-09-05 | Kevin O'Donnell | The CT Image Quality Assessment Procedure jumps around a bit and, for example, doesn't really address phantom positioning. Putting the text in order of performance (get phantom, position it, scan it, take measurements and calculate metrics) would be easier to understand/follow. Specific proposed text provided (See May 2017 document). | 4.1.1 | 2017-05-15 | Closed | The SLN Biomarker Committee prefers the current technical description followed by a simple set of clinical steps in Section 4.3. No change is being performed at this time. If significant issues arise with this during testing, this will be revisited.
3 | 2017-09-05 | Kevin O'Donnell | The algorithm for calculating Edge Enhancement seems to leave out details that an implementer would need to achieve results consistent with the reference implementation. Specific proposed text provided (see May 2017 document). (Although it would be good to get more eyes on this). Nick observed that he found the text a bit confusing. He got the broad idea but thought the current text might not be executed consistently. | 4.1.1 | 2017-05-15 | Closed | The SLN Biomarker Committee decided that the current text is preferred. If significant issues arise with this during testing, this will be revisited.
2 | 2017-09-05 | Kevin O'Donnell | The Nodule Volume Bias and Variation assessment procedure doesn't seem to take into account the work done in large nodules. Consider replicating 4.4 and 4.5 from the Large Nodule profile rather than reinventing essentially the same procedure. | 4.1.2 | 2017-05-15 | Closed | The current procedure has been reviewed by Nancy Obuchowski. The SLN Biomarker Committee decided that the current procedure description is preferred. If significant issues arise with this during testing, this will be revisited.
1 | 2017-09-05 | Kevin O'Donnell | The assessment procedures replicates some of the requirements from Section 3 and state others with different words. This makes it unclear which to conform to, which will likely result in confusion and variability in conformance. Don't replicate requirements. Kevin can help describe the use of the automated assessment tools without undercutting the validity of the profile. | 4.2 & 4.3 | 2017-05-15 | Closed | The SLN Biomarker Committee decided that the current procedure description is preferred. If significant issues arise with this during testing, this will be revisited.