## Results and Resolutions
Commenter | Comment No. | Line # | Priority | Issue | Proposal | Committee Discussion | Owner (opt.) | Resolution (w Rationale if rejected) | Status
[8 empty] | 1 | TBD | To be decided
See http://qibawiki.rsna.org/index.php/Public_Comment_Process for more guidance on the comment resolution process | [7 empty] | 0 | OK | No action requested
[8 empty] | 0 | Discuss | Need to decide resolution
[8 empty] | 0 | TODO | Resolution decided
[8 empty] | 0 | Done | Profile update completed
|  | Use PDF lines, they're stable |  | Describe your issue. Don’t write a book, but do include enough to indicate what you see as a problem. | Suggest new wording or describe a way to address the issue. The committee may simply accept your suggested text. Even if they don’t, it gives a good sense of what you’re looking for. Leaving this blank means you can’t imagine how to resolve the issue | <Cmte Notes as needed.> <Optionally, use the Owner column to divide up the work and assign rows to a committee member who will lead discussion and resolution> | <Delete this column if not used> | Describe how the comment was/will be resolved. May be simply accepted & changed as proposed, may be accepted & resolved differently, or may be rejected with a rationale for why. | TBD
KOD | 1 | 0 | L | Good looking profile. | Keep up the good work. | Thanks | SBF
KOD | 2 | 75 | L | Just FYI that change log is totally optional | If you're finding it helpful, keep it. If it's obstructively tedious to maintain, feel free to drop it. | Will drop. | SBF | Dropped.
KOD | 3 | 85 | L | Closed Issues can be dropped if they don't explain anything beyond what is/isn't in the profile. | E.g. Closed Issues 4,5,6 could be dropped. | Done. | SBF | Done.
KOD | 4 | 120 | L | The text mentions the 11HU cutoff but doesn't mention that lung VA is required | Clarify | Added, "11 HU after lung volume adjustment" throughout. | SBF | Added, "11 HU after lung volume adjustment" throughout.
KOD | 5 | 158 | L | The asymetry of the two claims gives me pause. | Consider either: - having three claims (RA-950 w/o VA, HU w/o VA, HU with VA) - have an HU claim w&w/o VA, and an RA-950 claim - deciding it doesn't bother you :-) | Broke up into 3 claims and reordered to state the most repeatable claim first (Perc 15 with VA). | SBF | Broke up into 3 claims and reordered to state the most repeatable claim first (Perc 15 with VA).
KOD | 6 | 212 | M | The protocol alone is not enough to meet the claim. | Qualify by saying "support meeting the claim" or "if combined with the requirements on the other actors, will meet or exceed the claim" | Done. | SBF | Done.
KOD | 7 | 218 | L | The sentence mixes the two claims (with and without VA). | Be explicit about VA. | Done. | SBF | Done.
KOD | 8 | 254 | M | The table seems to be missing some rows. E.g. the Physicist is part of Periodic QA but that is not listed. | Review/update this overview table. | Done. | SBF | Done.
KOD | 9 | 254 | M | Combo actors creates ambiguity. (See guideline about Single Actor) | Change Physicist/Technologist to Physicist Change "Clinician and Statistician" to Statistician | Done. | SBF | Done.
KOD | 10 | 264 | L | This is referred to as both Scanner Qualification and Product Qualification at different places in the document | Since it includes Analysis SW, go with Product Qualification (and fix references to Scanner Qualification elsewhere, and add 3.1 to Table 1) | Done. | SBF | Done.
KOD | 11 | 299 | L | Spatial Resolutions: Would be nice to include the section title after each section reference like you did in row 2 for 4.1.2. | Add for references to 4.1.3, 4.1.4, etc. | Done. | SBF | Done.
KOD | 12 | 299 | M | Acquisition speed: which is the pass fail requirement? Scan time? Table Speed? Both? Either? | Clarify | Coordinated between out committee and CT Volumetry to resolve on common language | SBF | Opted for <10 s to provide flexibility for meeting the claims; further refined after coordinating with CT volumetry: "Shall set parameter values that will cover an axial field of view of 35 cm in 10 seconds or less." Then in the discussion section we can clarify that this may include table speed, rotation time, collimation, pitch, etc..
KOD | 13 | 299 | M | CT radiation exposure: Inconsistent. It's called CT Dose in 3.4.2. And there isn't a shall here. | Remove the row since it's covered by row 1 anyway. | Done: Removed this row from Table 3.1.2 | SBF | Done: Removed this row from Table 3.1.2
KOD | 14 | 299 | M | Measured HU (Bias): Any assessment procedure is permitted? | Consider referencing an assessment procedure. | Done. | SBF | Done.
KOD | 15 | 299 | M | Measured HU (Bias): Most of the * text below probably belongs in the discussion section above. | Move the * text to 3.1.1 and consider just adding "when using a phantom with uniform low density regions" to the requirement. (Our shouldn't you just required use of such a phantom in the assessment procedure then don't mention it here?) | Done. | SBF | Done.
KOD | 16 | 299 | M | HU Stability: Any assessment procedure is permitted? | Consider referencing an assessment procedure, and mention the N=5 in the procedure rather than here. | Done. | SBF | Done.
KOD | 17 | 299 | L | Lung Density Analysis: Do you want to specify how any of these are calculated? |  | Done. | SBF | Done.
KOD | 18 | 299 | L | Reproducibility of Analysis Software: First line is duplicate of prior row (missing "for") | Drop the first line | Corrected. | SBF | Corrected.
KOD | 19 | 299 | M | Reproducibility of Analysis Software: Second line is a requirement on the operator at Image Analysis time | Move second line to section 3.10.2 | Corrected. | SBF | Corrected.
KOD | 20 | 299 | H | Reproducibility of Analysis Software: Third line is in passive voice and it's not clear who has to do exactly what assessment and what the pass/fail mark is | Revise to resolve those issues. As one (perhaps not relevant) example, see the assessment of the segmentation software in the CTVol profile. | Corrected. | SBF | Corrected.
KOD | 21 | 299 | M | Rather than "more detail"… | Follow your pattern from above "as described in Section 4.2" | Done. | SBF | Done.
KOD | 22 | 313 | L | Has lack of either of these certifications commonly been the source of failing to achieve the profile claim? | If not, consider dropping the section. | Done. | CH | Done
KOD | 23 | 322 | M | Scanner Calibration: Breaks guideline on having someone clearly responsible. | Remove Tech. If you want borrow a note from CTVol that performance of the activity can be delegated, but the responsibility is with the person in the physicist role that this happen. | Done. | SBF | Removed technologist from Table 3.3.2
KOD | 24 | 322 | M | Scanner Calibration: The Acquisition Device requirement is a bit unclear. Is the ideal that the vendor service should come repeat all the 3.1.2 assessments? | Consider dropping the Acquisition Device reference here. The Physicist will check the parameters you list here. | Done. | SBF | Dropped the acquisition device reference from Table 3.3.2.
KOD | 25 | 322 | M | Scanner Calibration: The sentence is a bit ambiguous about what the physicist has to do. | Shall confirm, at least quarterly (wasn't it monthly in the first row of the table?), that <state the specific parameters and the pass/fail mark, but reference the assessment procedure in 4.x rather than duplicating some of those details here> | Done. | SBF | Changed to monthly and referenced specifications in 3.1.2 and procedures in 4.1.1.
KOD | 26 | 322 | M | HU Stability: (might be just me but) seems a bit unclear what is being subtracted from what. Scan N compared to Scan N-1 (subsequent)? Scan N compared to Scan 0 (baseline)? | Reword to clarify and perhaps have an assessment procedure that state mathematically what is being assessed. | Done. | CH | Clarified by refereing to the detailed procedures in Table 3.1.2.
KOD | 27 | 334 | L | Paragraph seems mostly duplicative | Can probably be dropped. | Done. | SBF | Moved this apragraph up to Section 3.3 and edited so as to retain unique information and references.
KOD | 28 | 418 | L | Acquisition Protocol: Having two requirements in one row makes copy-paste to construct a checklist awkward since they should really be separate checklist items. | Consider splitting into two rows. (Same in a few other places too) | Done. | SBF | Split the rows for Acquisition and Reconstruction Protocol in Table 3.4.2.
KOD | 29 | 418 | L | Technologist: Is the Tech typically central to protocol design? | If not, consider changing Tech to Radiologist. | Done. | SBF | In Table 3.4.2, changed to Radiologist actor except for slice thickness which was changed Physicist actor.
KOD | 30 | 418 | M | CT Dose: Needs a "shall" and its not clear what "with the amount of radiation adjusted based on patient size and shape according to manufacturer." adds to the requirement. Also, this doesn't affect the claim so not really a profile requirement. Could move into discussion about dose above. | Revise or remove. | KOD-30 and AAPM-6, As a committee we decided that the low dose protocol design is a recommendation and not required to meet the claims. Therefore it is removed from this table. | SBF | KOD-30 and AAPM-6, As a committee we decided that the low dose protocol design is a recommendation and not required to meet the claims. Therefore it is removed from this table. Note that the recommendation for low dose protocol design is addressed in the Discussion in this section and below. The recommended protocols in Appendix E are also designed around this spec.
KOD | 31 | 418 | L | Scan Duration: | Consider moving to Image Acquisition? | Addressed | SBF | Sated in both location and changed to <10 cm breath-hold.
KOD | 32 | 418 | L | Resolution: Why imply a difference between the 1mm requirement and the 1.0mm requirement | Use the same text. | Done. | SBF | Changed to 1.0 mm throughout.
KOD | 33 | 418 | L | Protocol validation requirements don't all reference a procedure | Reference the corresponding 4.x procedures. (And not just "Section 4") | Done. | SBF | Added references to procedures as suggested.
KOD | 34 | 473 | L | Lung Coverage: Clarify "match the display field of view between time points". Note that in CTVol we decided it was unlikely techs would pull up priors to check this. | "Shall set the display field of view of the current study to match that of the baseline study for the subject (to ensure consist spatial resolution)." | Removed per comments 41-42. | SBF | Removed.
KOD | 35 | 473 | M | Lung coverage: It's a bit confusing to have specific numerical values in a requirement but then to preface them with "e.g." | Either drop the E.g. and make the numbers normative, or drop the numbers and if people want some hints they'll see them in the discussion section above. If you choose the latter, break out a paragraph in discussion that starts with "Lung Coverage" in bold to make it easy to find. | Removed per comments 41-42. | SBF | Removed.
KOD | 36 | 473 | L | Breath hold: want to avoid slipping discussion text into being normative. | Change "specified" to "described" | Done. | SBF | Changed to "decribed."
KOD | 37 | 473 | M | Contrast: List of two actors and use of passive voice makes it nicely ambiguous. So if the Radiologist prescribes contrast for a study, the Tech is mandated to simply not administer it? | Consider a variant on the way contrast is currently handled in CTVol that steers to no contrast. | Done. | SBF | Changed to active voice and made actor Radiologist.
KOD | 38 | 491 | M | Acquisition Protocol Selection: "Same CT Scanner". Is your intention to require the same model, or the same actual device? | Clarify. Also, I'd be inclined to put machine selection and protocol selection into different parameters/rows. | Requires discussion. Is it sufficient to use the same scaner make and model, or the same exact scanner. The latter may difficult for realistic longitudinal multi-center designs. | SBF | Change to "Same CT scanner make and model."
KOD | 39 | 491 | M | Acquisition Protocol Selection: "Qualified Device". The scanner doesn't really have something to do here. | Use the phrase "Conformant Acquisition Device" in your requirement on scanner selection and then drop the third row here. | Remove. | SBF | Removed.
KOD | 40 | 491 | L | Scan Plane: Need a shall. And is this really a common source of variance? | Revise or remove. | Done. Agree this is confusing. | SBF | Removed.
KOD | 41 | 491 | L | Anatomic Coverage: Nice normative requirement in the appropriate section. | Since it's covered here, just remove the differently named "Lung Coverage" parameter up at line 473. | Done. Agreed. | SBF | Removed at line 473.
KOD | 42 | 491 | L | Axial Field of View: Ditto the Anatomic Coverage comment | Remove the corresponding parameter at 473. The text here is better. | Done. | SBF | Removed.
KOD | 43 | 491 | L | CT Dose: There's no shall here and maybe it should just come out per comment above. You've covered it in discussion. | Remove 3 mGy CTDI vol target dose from the specifications. It is a guidance and not a "shall" to meet the claims. | See Comment No. 30: As a committee we decided that the low dose protocol design is a recommendation and not required to meet the claims. Therefore it is removed from this table. | SBF/Committee | Removed 3 mGy CTDI vol target dose from the specifications. It is a guidance and not a "shall" to meet the claims.
KOD | 44 | 513 | L |  | Change grey text to black. |  | SBF | Done.
KOD | 45 | 525 | L | Don't hide "shalls" in the discussion section. | Move into the table below. |  | SBF | Moved this to the table in section 3.8 above and then deleted section 3.9.
KOD | 46 | 528 | M | It doesn't say what happens if one or more of the QA evaluations or confirmations fails (i.e. you can still use the scan if you need to, but don't assume the measurements have the performance stated at the beginning of the profile.) | Clarify. CTVol is an improvement but could probably be a bit more specific too. | Agree, added text to this effect. | CH/SBF | Added the following text, "If the image acquisition and software analysis do not conform to the above specifications, the profile performance specifications may no longer be valid."
KOD | 47 | 528 | M | CT Dose: This means that if the CTDIVol is not Per Protocol then you will disqualify the measurements. | Consider dropping as a requirement. | Done | CH | Dropped
KOD | 48 | 528 | M | CT Dose: (0018,9345) is optional in the DICOM header and it's a numeric value. Not sure how "Per Protocol" would be indicated so it's unclear what this requirement means. |  | Removed reference as helpful for safe implementation of longitudinal studies but not necessary to achieve claims. | CH | Removed reference to CT Dose
KOD | 49 | 528 | M | Conformance to Baseline: Implies that a change of 20 or 30% in FOV would degrade the measurement, but if there were substantial changes in subject size then it would not degrade the measurement. | Consider removing the exception and acknowledge in the discussion that subject size changes may make FOV changes unavoidable, but that still affects the measurement performance so would still need to disqualify. Or some other discussion clarifying the issue. | Agreed the references to 10% FOV change were arbitrary and did not address a central challenge of density, but more an issue for maintaining spatial resolution for airways. | CH | Removed reference to FOV leaving the general "shall confirm protocol is consistent with baseline."
KOD | 50 | 528 | M | Absence of IV Contrast, and Lung Volume: Seems like an unnecessary bar to mandate that software be able to detect these by itself. Could still do good measurements without these features. | Consider making these two requirements on the Technologist like the rest in this section. | Agreed. | CH | Done.
KOD | 51 | 531 | M | Has compression been a common source of problems? | If yes, make a requirements table and move the shall into it and be clear what actor the shall is on; If not, consider dropping the section. | These concerns became moot with the removal of Section 3.9 | SBF | Reworded to be more specific that "manual segmentation is labor intensive necessitating automated techniques."
KOD | 52 | 532 | L | Anonymization is presumably not used in clinical practice, and in any case is not relevant to meeting the claim | Remove reference to anonymization. | Ditto above regarding removal of Section 3.9. | SBF | Removed reference to anonymization.
KOD | 53 | 565 | M | There's quite a bit bundled into single cells. | Try to factor out into elementary parameters/features/performance. | Agreed | CH | Done
KOD | 54 | 565 | M | The list of segmentation features seems like something to check at product qualification time. | Move this requirement to 3.1 | Agreed | CH | Done
KOD | 55 | 565 | M | "same analysis software shall be identical" is a bit redundant and ambiguous. Identical software means same product, version, patch, and configuration? To what detail? Same Neural Net weighting values? | Clarify and consider wording from CTVol to reprocess baseline result if on different software. | Agreed | SBF | Removed "same" and added parenthetical language to encourage image reanalysis.
KOD | 56 | 565 | M | "shall be deterministic" - is this typical? In any case, seems like a product qualification not a runtime variation. | Move this requirement to 3.1 | Agreed | CH | Reworded to read, "The analysis software used shall be identical for each longitudinal time point measured (reanalyze images if necessary)."
KOD | 57 | 567 | M | No "shalls" and there aren't any claim relevant requirements. | Consider dropping the section. Can move the Discussion paragraph up to the "Clinical Interpretation" section below the claim if you want to keep it. | This is the image interpretation section. Most of the substance in the section is discussed below the claims. KOD suggests removing entirely. Need committee discussion. | SBF | Removed Table 3.10 but retained section to reinforce concepts.
KOD | 58 | 629 | M | This chunk of shalls belongs in an assessment procedure rather than the Section 4 preamble. | Move into 4.1.1 | This is a good suggestion. | SBF | Moved to 4.1.1 as suggested.
KOD | 59 | 629 | M | Ambgiuous if this is required. Maybe it's just for the minimum standard, so you can do anything else if you think it's better. The procedure CAN be followed but you don't have to. | Drop the opening sentence and just go with the "shalls" | We opted not to implement this suggestion to be consistent with AAPM-8 and 9. | SBF | Chose to not implement.
KOD | 60 | 630 | L | Grouping/sequencing feels a bit odd. | 1. The assessor shall warm up the x-ray tube and perform calibration scans (often called air-calibration scans) according to scanner manufacturer recommendations. 2. The assessor shall place a phantom containing regions of uniform low density (such as the COPDGene phantom 16,17) at the isocenter of the scanner. It is expected that scout (topogram, scanigram, etc.) images will be obtained to optimize positioning and coverage of the phantom. 3. The assessor shall scan the phantom using acquisition protocol and reconstruction parameters that conform to this Profile (See Section 3.6.2 and 3.7.2). <Then the opening step in 4.1.2 and 4.1.3 could be "The assessor shall scan the phantom as described in Steps 1-3 of 4.1.1"> <In principle we should perhaps leave out Step 3 so the procedure is generic and it is 3.1.2 that says you have to make these assessments with a conformant protocol> | This is redundant with changes above. We have blended these into Section 4.1.1. per KOD-58. | SBF | Resolved by previous changes.
KOD | 61 | 642 | L | Missing the intro sentences summarizing the scope of use and form of the assessment procedure. | Add based on template. | This was addressed by AAPM-8,9 and by modifications to address KOD-58 | SBF | Resolved by previous changes.
KOD | 62 | 645 | L | the standards? | "the assessments" |  | SBF | Changed to "lung equivalent foam standards".
KOD | 63 | 648 | L | The terms "repeated realizations", "repeated acquisitions", "repeated scans", "N=5 acquisitions", "unique realizations" all mean roughly the same thing. | Pick one term, maybe repeated acquisition. |  | SBF | Done.
KOD | 64 | 650 | M | "specifications in 3.1.2" The assessment procedure just explain how to generate an assessment value correctly for HU Bias, etc. It is the requirement table that sets pass/fail limits on those values. So 4.1 sections don't need to reference 3.x sections since it's redundant. | Drop the reference here and in line 661. It limits use of the procedure and duplicates information that then needs to be maintained in multiple places. Same fix in a few more places further down. |  | SBF | Done.
KOD | 65 | 652 | L | Why "approximately cylindrical ROI"? | Can we just say place a ROI since you later constrain it to have 3,000 voxels and be within the material insert? And just to be clear when we say cylindrical, we're not implying that you span multiple slices, right? | Agreed. | CH | Done.
KOD | 66 | 656 | L | An ROI shall be drawn… | The assessor shall draw the ROI |  | SBF | Done.
KOD | 67 | 657 | L | cylindrical circumference? | … of the material insert being measured? |  | SBF | Done.
KOD | 68 | 668 | L | The title is Voxel Noise and Noise Power Spectrum, but the procedure provides Voxel Noise Bias and Voxel Noise Standard Deviation. | Maybe just title it Voxel Noise, and clarify in the comment for line 670 |  | SBF | Changed to "Voxel Noise" as suggested.
KOD | 69 | 670 | M | The template second intro sentence describing "voxel noise is assessed in terms of …" is missing and would help tighten up the procedure. | Add such a sentence and highlight this procedure (as currently written) is specific to lung density equivalent. |  | SBF | Changed as suggested.
KOD | 70 | 676 | L | Says to record the ROI Mean and insure it meets the bias requirement. | Connect the dots between mean and bias a bit more explicitly. |  | SBF | Modified to make more explicit.
KOD | 71 | 683 | M | What does it mean that "The vendor shall specify that the kernel is matched"? | Usually all requirements here are on the assessor. Also, state what they are required to do in a way that there is an unamgiguous pass/fail. | This may need discussion to insure the kernel is matched, or is it sufficient to use an "appropriately smooth" kernel and be done? | SBF | Reworded this section to apply to the assesor, added a figure of a typically smooth MTF (GE Standard), and tightened up the text: "The assessor/vendor shall confirm that the reconstruction kernel for the proposed protocol design (Section 3.4 and Table 3.4.2) is matched to an appropriately smooth kernel such as the GE Standard (Figure 4), Siemens B31f or B34f, Philips B, and Cannon FC17."
KOD | 72 | 694 | M | Missing summary | Add assessment procedure summary paragraph from template. |  | SBF | Added as requested.
KOD | 73 | 694 | M | Duplicating material from elsewhere | Drop first two paragraphs Lines 694-700 |  | SBF | Replaced with intro as requested.
KOD | 74 | 700 | M | Assessment procedures should describe the prescribed way to assess and compute a metric value. The pass/fail threshold is in the profile requirement in section 3.x | State in the intro paragraph that in-plane resolution is assessed in terms of FWHM and require that a FWHM value be computed and recorded (you can state or reference the details of that computation if you think it necessary). Don't duplicate the 1.0 threshold since it's already stated in 3.x. |  | SBF | Done.
KOD | 75 | 709 | L | Avoid passive voice | "The assessor shall …" |  | SBF | Chaned throughout Section 4.1.3 to active voice.
KOD | 76 | 706 | M |  | Make this 4.1.5? And follow the pattern as described for 718 below. |  | SBF | Resolved by new section defined for through-plane resolution in comment 77 below.
KOD | 77 | 718 | M | Through-plane spatial resolution requirement in 3.1.2 references 4.1.4. | Consider making this text into 4.1.4 and add the overview paragraph. And tidy up a bit to remove duplicate references to CATphan and duplications of requirements. |  | SBF | Created new section 4.1.4 to address through-plane resolution separate from in-plane.
KOD | 78 | 736 | H | Seems like analysis software performance and repeatability is a key consideration but there are details yet to be nailed down. (Not fun but probably must be done). The text here is a start that alludes to that. Just need to continue. | Should probably come up with specific performance metrics for the analysis software and a clear assessment procedure for generating those metrics. | No more discussion necessary. I think the committee has resolved this and suggest Kevin may not be aware. | SBF/CH | Extensive evaluation procedures are in Appendices F and G.