Volumetry for Lung Cancer Tumor Assessment using X-Ray Computed Tomography

Version 1.9

5 May 2011

Table of Contents

I. Executive Summary	2

II. Clinical Context and Claims	3

Utilities and Endpoints for Clinical Trials	3

Claim:  Measure Longitudinal Change in Whole Tumor Volume	3

III. Profile Details	4

1. Subject Handling	4

2. Imaging Data Acquisition	8

3. Imaging Data Reconstruction	11

4. Image Analysis	13

IV. Compliance	15

Acquisition Devices	15

Reconstruction Software	15

Software Analysis Tool	16

Performing Site	16

References	16

Appendices	18

Acknowledgements and Attributions	18

Background Information	19

Conventions and Definitions	21

Model-specific Instructions and Parameters	22

# I. Executive Summary

X-ray computed tomography (CT) is often an effective imaging technique for assessing therapy response. In clinical practice, qualitative impressions are generally based on visual inspection of the images. Quantification is helpful when tumor masses change slowly over the course of illness. When measurements are made, they are often only done as uni-dimensional diameters on axial slices, as specified for the RECIST response criteria.  Over time limitations of this method have been reported.  Many investigators have suggested that quantifying whole tumor volumes could solve many of the limitations of depending on diameter measures, and have a major impact on patient management [1-2]. An increasing number of studies have shown that volumetry has value [3-12]. However, technical problems have delayed its adoption [13]. Historically, substantial amounts of effort were required as well as precision [14-16] and accuracy [17] when sufficient standardization is not in place have led to concerns about the risks of confusing variability with medically meaningful changes.

QIBA has constructed a systematic approach for standardizing and qualifying volumetry as a biomarker of response to treatments for a variety of medical conditions, including lung disease [18]. Several studies at various scopes are now underway to provide comparison between the effectiveness of volumetry and uni-dimensional RECIST in multi-site, multi-scanner-vendor settings. This QIBA Profile is expected to provide specifications that may be adopted by users as well as equipment developers to meet targeted levels of accuracy and clinical performance in identified settings.

This Profile sets out performance claims for measuring the volumes of malignant lesions as well as details on how to meet them. It describes the requirements placed on human and computer-controlled actors in the following contexts: (1) managing individual patients in medical settings, and (2) quantitatively evaluating therapeutic responses in clinical trials.

The intended audience of this document is:

- Technical staff of vendors planning to participate in the QIBA initiative

- Practicing clinicians at healthcare institutions considering appropriate specifications for acquiring equipment

- Experts involved in quantitative medical image analysis

- Anyone interested in the technical and clinical aspects of medical imaging

Summary of Clinical Trial Usage as described in assimilated protocol "Volumetric Image Analysis for Cancer using X-Ray Computed Tomography"

The original WHO response criteria were based on bi-dimensional measurements of the tumor and defined response as a decrease of the sum of the product of the longest perpendicular diameters of measured lesions by at least 50% [19].  The rationale for using a 50% threshold value for definition of response were data evaluating the reproducibility of measurements of tumor size by palpation and on planar chest x-rays [20]. The more recent RECIST criteria introduced by the National Cancer Institute (NCI) and the European Organisation for Research and Treatment of Cancer (EORTC) standardized imaging techniques for anatomic response assessment by specifying minimum size thresholds for measurable lesions and considered other imaging modalities beyond CT. As well, the RECIST criteria replace longest bidirectional diameters with longest uni-dimensional diameter as the representation of a measured lesion [21]. RECIST defines response as a 30% decrease of the largest diameter of the tumor. For a spherical lesion, this is equivalent to a 50% decrease of the product of two diameters.  These thresholds are dominated by the reproducibility of diameter measurements and are subject to a variety of errors as diameters seek to approximate the volume of a mass.

The variability of RECIST and WHO criteria has encouraged volumetric approaches to the anatomic measure of response assessment. The rationale for volumetric approaches is multifactorial.  First, lung cancers may grow and regress irregularly in three dimensions. Measurements obtained in the axial plane fail to account for growth or regression in the longitudinal axis, whereas volumetric measurements incorporate changes in all dimensions. Secondly, changes in volume are less subject to either reader error or interscan variations. For example, partial response using the RECIST criteria requires a greater than 30% decrease in tumor diameter, which corresponds to greater than 50% reduction in volume of tumor. If one assumes a 21 mm diameter lesion (of 4850 mm3 volume), partial response would result in a diameter of 18 mm, but a decrease in volume to 2145 mm3. The much greater magnitude of volumetric changes is less prone to measurement error than changes in diameter, particularly if the lesions are irregularly shaped or speculated.  With the ability to compute volume directly, it is increasingly demonstrated that the thresholds may tighten to get closer to biological change and/or to move from categorical to continuous response assessment.

This document describes image acquisition, quality control, processing, and analysis for cancer tumor assessment using volumetry. The context of use is to assess longitudinal measurements of change in tumor volume over relatively short time-intervals to predict treatment response in clinical trials.

# II. Clinical Context and Claims

The clinical context sets out the utilities and endpoints for clinical trial usage and then proceeds to identify targeted levels of quality for named measurement read-outs that may be used in the relevant clinical indications.

## Utilities and Endpoints for Clinical Trials

These specifications are appropriate for quantifying the volumes of malignant lesions and measuring their longitudinal changes within subjects. The primary objective is to evaluate their growth or regression with serially acquired CT scans and image processing techniques. The information about volumetric change will drive management decisions in diagnostic settings as well as clinical trials in patients with known malignancies.

## Claim:  Measure Longitudinal Change in Whole Tumor Volume

Quantitative measurements of longitudinal change in whole tumor volume are more precise (reproducible) than uni-dimensional measurements of tumor diameter. Specifically, longitudinal change of >=26% of tumors with initial diameter between 10-80mm has a 95% confidence of being due to biological change rather than technical factors provided the imaging steps are done in compliance with the details below.

# III. Profile Details

A technical description of tests for the biomarker, identifying measurement activities and read-outs, is provided:

The following sections provide details for the various components shall do in order to be in compliance:

Section 1, Subject Handling, is practiced by a Performing Site.

Section 2, Imaging Data Acquisition, is practiced by a Performing Site using an Acquisition Device.

Section 3, Imaging Data Reconstruction, is practiced by a Performing Site using Reconstruction Software.

Section 4, Image Analysis, is practiced by a Performing Site using one or more Software Analysis Tools.

The requirements included herein are intended to establish a baseline level of capabilities. Providing higher performance or advanced capabilities is both allowed and encouraged and the profile is not intended to be limiting in any way with respect to capabilities. The intention is not to dictate implementation details.

## 1. Subject Handling

1.1 Timing Relative to Index Intervention Activity

The pre-treatment CT scan shall take place prior to any intervention on the patient, including percutaneous needle biopsy. This scan is referred to as the “baseline” scan.  A dedicated chest CT scan shall be taken immediately after biopsy.  The post-treatment scan shall be taken on the day after completion of treatment with the study drug.

1.2 Timing Relative to confounding Activities (to minimize “impact”)

This protocol does not presume any timing relative to other activities. Fasting prior to a contemporaneous FDG PET scan or the administration of oral contrast for abdominal CT is not expected to have any adverse impact on this protocol.

1.3 Scheduling Ancillary Testing

If associated biopsy/resection is expected to be performed during the same visit as the imaging procedure, it shall be described in the Trial Calendar.

1.4 Subject Preparation Prior to Arrival

Local standard of care shall be followed for CT with contrast.  No requirements are stated otherwise.

1.5 Subject Preparation Upon Arrival

Beyond a clear, simple language description of the image acquisition procedure, no exam preparation is specified beyond the local standard of care for CT with contrast.

1.6 Imaging-related Substance Preparation and Administration

The use of contrast is not an absolute requirement for this protocol. However, the use of intravenous contrast material may be medically indicated for the diagnosis and staging of lung cancer in defined clinical settings.   Contrast characteristics influence the appearance and quantification of the tumors.  The following requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Type of Contrast Agent | Acceptable Another brand or contrast agent type may be used if medically indicated, e.g., a switch from ionic to non-ionic contrast media Target A subject shall be scanned with equivalent contrast agent for each scan Ideal All subjects shall be scanned with equivalent contrast media |
| Use of contrast in follow-up scans | Acceptable If used at baseline, equivalent contrast shall be used at subsequent time points. If not used at baseline, it shall not be used in follow-up scans |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Whether contrast was used and if so, Type of Contrast Agent shall be recorded. |

1.6.1 Dose Calculation and/or Schedule

Site-specific sliding scales that have been approved by local medical staffs and regulatory authorities shall be used for patients with impaired renal function (e.g., contrast dose reduction based on creatinine clearance).

| Parameter | Compliance Levels |
| --- | --- |
| Dose Calculation and Schedule | Acceptable If a different brand or type of contrast is used, the dose shall be adjusted to ensure comparability as indicated and as documented by peer-reviewed literature and/or the contrast manufacturers’ package inserts. Target For a given subject, the same contrast dose shall be used for each scan subject to the medical condition of the patient. |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Actual contrast media Dose Calculation and Schedule shall be recorded. |

1.6.2 Administration Route

The following specifications are noted.

| Parameter | Compliance Levels |
| --- | --- |
| Administration route | Acceptable Intravenous bolus injection may be in any vein but shall be via butterfly catheter. Target Injection via butterfly in a large antecubital vein. Ideal Injection in a large antecubital vein known to be patent from observation of intravenous saline drip. |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Actual contrast media administration details shall be recorded. |

1.6.3 Rate, Delay and Related Parameters / Apparatus

The following specifications are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Contrast administration | Acceptable Manually. Target At the same rate for each scan. Ideal Via a power injector. |
| If a different brand or type of contrast is used | Acceptable The rate shall be adjusted to ensure comparability if appropriate and as documented by peer-reviewed literature and/or the contrast manufacturers’ package inserts. |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Actual contrast media Rate, Delay, and Apparatus utilized shall be recorded. |

1.7 Subject Positioning

Consistent positioning is required to avoid unnecessary variance in attenuation, changes in gravity induced shape and fluid distribution, or changes in anatomical shape due to posture, contortion, etc. Careful attention shall be paid to details such as the position of their upper extremities, the anterior-to-posterior curvature of their spines as determined by pillows under their backs or knees, the lateral straightness of their spines, and, if prone, the direction the head is turned.

If the previous positioning is unknown, the subject shall be positioned Supine/Arms Up/Feet First if possible. This has the advantage of promoting consistency, and reducing cases where intravenous lines, which could introduce artifacts, go through gantry.

| Parameter | Compliance Levels |
| --- | --- |
| Subject Positioning | Acceptable May be placed in a different position if medically unavoidable due to a change in clinical status. Target Same positioning shall be used for each scan. Ideal Shall be Supine/Arms Up/Feet First. |
| Table Height | Acceptable Table height shall be adjusted to place lung field at isocenter. |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Actual Subject Positioning and Table Height shall be recorded. |

1.8 Instructions to Subject During Acquisition

Breath holding reduces motion that might degrade the image. Full inspiration inflates the lungs, which is necessary to separate structures and make lesions more conspicuous.

| Parameter | Compliance Levels |
| --- | --- |
| Breath hold | Acceptable Shall be at least near the high end inspiration. Target Subjects shall be instructed to hold a single breath at full inspiration. |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Factors that adversely influence patient positioning or limit their ability to cooperate (e.g., breath hold, remaining motionless, agitation in patients with decreased levels of consciousness, patients with chronic pain syndromes, etc.) shall be recorded. |

1.9 Timing/Triggers

For each subject, the time-interval between the administration of intravenous contrast and the start of the image acquisition shall be determined in advance, and then maintained as precisely as possible during all subsequent examinations.  For lung masses, image acquisition shall be timed to coincide with visualization of the thoracic arteries.

| Parameter | Compliance Levels |
| --- | --- |
| Timing / Triggers | Acceptable Shall use a standard time. Target Shall be automatically triggered on detection of contrast bolus arrival. |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Actual Timing and Triggers shall be recorded. |

1.10 Required Visualization / Monitoring, if any

No particular visualization or monitoring is specified beyond the local standard of care for CT with contrast.

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Shall provide means to record any Actual events observed by the technician that may have an effect on scan quality according to local standard. |

## 2. Imaging Data Acquisition

CT scans for tumor volumetric analysis will be performed on qualified equipment, and all CT scans for an individual participant shall be performed on the same platform throughout the trial. In the rare instance of equipment malfunction, follow-up scans on an individual participant can be performed on the same type of platform. All efforts shall be made to have the follow-up scans performed with identical parameters as the first.  This shall be inclusive of as many of the scanning parameters as possible, including the same field of view (FOV).

A set of scout images shall be initially obtained. Next, in a single breath hold, contiguous thin section slices from the thoracic inlet to the adrenal glands are obtained. Pitch shall be chosen so as to allow completion of the scan in a single breath hold. In some cases two or more breaths may be necessary. In those cases, it is important that the target lesion be fully included within one of the sequences.

2.1 Data Content

The image data is consistently reconstructed attenuation values (CT number) in the specified anatomic region of the lung.  The image data is acquired in less than some scan duration.  Document will use “required coverage” to mean the specified anatomic region of interest.  Field of View affects downstream reconstructed pixel size because the fixed image matrix size of most CT scanners is 512X512.  If it is necessary to expand the field of view to encompass more anatomy, the resulting larger pixels are less than ideal. Faster Rotation Speed reduces the breath hold requirements and reduces the likelihood of motion artifacts.  Scan Plane may differ for some subjects due to the need to position for physical deformities or external hardware, but shall be constant for each scan of a given subject.

The following parameters describe what the acquired images shall contain/cover.

| Parameter | Compliance Levels |
| --- | --- |
| Anatomic Coverage | Acceptable Entire lung fields, bilaterally (lung apices through bases) Target Entire lung fields, bilaterally (lung apices through adrenal glands) |
| Field of View | Acceptable Complete thorax : 0.55 to 1.0mm Target Rib-to-rib: 0.55 to 0.8mm |
| Scan Duration | Acceptable 30 sec Target 15 sec Ideal 10 sec |
| Scan Plane (Image Orientation) | Acceptable Scan Plane may differ for some subjects due to the need to position for physical deformities, but shall be constant for each scan of a given subject. Target Transverse |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Actual Anatomic Coverage, Field of View, Scan Duration, and Scan Plane shall be recorded. |

2.2 Data Structure

Collimation Width (defined as the total nominal beam width) is often not directly visible in the scanner interface. Wider collimation widths can increase coverage and shorten acquisition, but can introduce cone beam artifacts which may degrade image quality.  Pitch impacts dose since the area of overlap results in additional dose to the tissue in that area. Overlaps of greater than 20% have insufficient benefit to justify the increased exposure. Slice Width directly affects voxel size along the subject z-axis. Smaller voxels are preferable to reduce partial volume effects and provide higher precision due to higher spatial resolution.

| Parameter | Compliance Levels |
| --- | --- |
| Total Collimation Width | Acceptable 5 to 125mm Target 10 to 80mm Ideal 20 to 40mm |
| IEC Pitch | Acceptable Less than 2.0 Target Less than 1.5 Ideal Less than 1.0 |
| Tube Potential | Acceptable 110 to 140 with no tube current modulation Target 110 to 130 |
| Slice Width | Acceptable <= 3.0mm Target 1.0 to 2.5mm Ideal <= 1.0mm |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Actual Total Collimation Width, Single Collimation Width, Scan Pitch, Tube Potential, and Slice Width shall be recorded. |

2.3 Data Quality

Motion Artifacts may produce false targets and distort the size of existing targets. “Minimal” artifacts are such that motion does not degrade the ability of image analysts to detect the boundaries of target lesions.

Placeholder text from [23]: “Purpose: To provide a scientific basis for setting sinogram modeling accuracy targets based on impact of such errors on image quality. Modeling inaccuracies in photon spectrum and scatter distribution assumed by statistical image reconstruction (SIR) algorithms lead to systematic image artifacts. Methods and Materials: A synthetic two‐dimensional phantom (25×35 cm) was used to generate both noiseless and noisy sinogram data, based upon a 120 kVp spectrum filtered by 12 mm Al (66.6 keV mean energy)and variable scatter levels (4%, 20%, and 100% of the minimum primary transmission through the phantom). A third generation Siemens Somatom Plus 4 scanner geometry was assumed. The SIR algorithm was the alternating minimization (AM) algorithm [IEEE TMI 26:283]. 500 AM iterations using 22 ordered subsets were applied to the data. Various mismatches between the assumptions in the algorithm and the truth were studied, including erroneous spectra (110kVp to 130kVp, filtration from 6 mm to 18 mm Al, or 62.2 to 69.7 keV mean energy) and erroneous scatter levels (0.25 to 4.0 times the actual sinogram scatter). Result: AM image quality was evaluated in terms of bias, noise, contrast ratio, etc. To assure +/−2% accuracy in the reconstructed attenuation image, photon spectrum uncertainties corresponding to 2 keV shifts in mean energy can be tolerated. For a 30 cm thick subject, this corresponds to errors in primary transmission of 6%–8%. For 20% scatter levels, the maximum tolerated discrepancy in scatter‐to‐primary ratio (SPR) is about 5% to 8%and 30%–50% for typical MSCT scatter levels. Conclusions: This work indicates AM and other SIR algorithm image estimates are sensitive to errors in the detector response models assumed by the algorithms. For thick patients, a sinogram modeling accuracy of 6% is needed to support reconstructed images of 2% accuracy. “

Note that mAs is not specified here but is instead determined for each model and represented in Model Specific Parameters in order that the manufacturer make recommendations on how to best establish operating points for their equipment that meets all requirements simultaneously.

| Parameter | Compliance Levels |
| --- | --- |
| Motion Artifact | Acceptable Minimal artifact Target No artifact |
| Photon Spectrum Uncertainty | Acceptable 2 keV shifts in mean energy (placeholder pending group discussion) |
| Scatter to Primary Ratio | Acceptable 5% to 8%and 30%–50% for typical MSCT scatter levels (placeholder pending group discussion) |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Actual Motion Artifact, Photon Spectrum Uncertainty, Scatter to Primary Ratio as well as the model-specific Acquisition Device parameters utilized to achieve compliance with these metrics shall be recorded. |

## 3. Imaging Data Reconstruction

The image data is a uniform digital sampling of the reconstruction.  The sampling is assumed to satisfy Nyquist–Shannon sampling requirements to obtain all the information available in reconstruction.  The axial image sampling is 512X512, which may be inadequate for large subjects, 1024X1024 would be ideal.  The distances between axial image reconstructions (slice increment) should be constant.

Spatial Resolution quantifies the ability to resolve spatial details. Lower spatial resolution can make it difficult to accurately determine the borders of tumors, and as a consequence, decreases the precision of volume measurements.  Spatial resolution is mostly determined by the scanner geometry (which is not usually under user control) and the reconstruction algorithm (which is under user control).  It is stated in terms of “the number of line-pairs per cm that can be resolved in a scan of resolution phantom (such as the synthetic model provided by the American College of Radiology and other professional organizations).” –OR– “the full width at half of the line spread function”.

Noise Metrics quantify the magnitude of the random variation in reconstructed CT numbers.  Some properties of the noise can be characterized by the standard deviation of reconstructed CT numbers over a uniform region in phantom.  Noise (pixel standard deviation) can be reduced by using thicker slices for a given mAs. A constant value for the noise metric might be achieved by increasing mAs for thinner slices and reducing mAs for thicker slices.  The limitation of the standard deviation as complete measure of noise is the pixel standard deviation is modified by changes in spatial characteristics of reconstruction kernel in the absence of changes in radiation dose or slice thickness.

Acquisition Field of View affects reconstructed pixel size because the fixed image matrix size of most CT scanners is 512X512.  If it is necessary to expand the field of view to encompass more anatomy, the resulting larger pixels are less than ideal. Pixel Size directly affects voxel size along the subject x-axis and y-axis. Smaller voxels are preferable to reduce partial volume effects and provide higher measurement precision.  Pixel size in each dimension is not the same as resolution in each dimension; inherent resolution is different than how the data happens to be sliced and diced.   It is important not to throw away resolution to match the worse to the better.  Reconstruction Interval (a.k.a. Slice Thickness) that results in discontiguous data is unacceptable as they may “truncate” the spatial extent of the tumor, degrade the identification of tumor boundaries, confound the precision of measurement for total tumor volumes, etc.  Reconstruction Interval and Pixel Spacing shall be the same for each scan of a given subject.

Reconstruction Overlap shall be the same for each scan of a given subject. Decisions about overlap shall consider the technical requirements of the clinical trial, including effects on measurement, throughput, image analysis time, and storage requirements. Reconstructing datasets with overlap will increase the number of images and may slow down throughput, increase reading time and increase storage requirements.  It shall be noted that for multidetector row CT (MDCT) scanners, creating overlapping image data sets has NO effect on radiation exposure; this is true because multiple reconstructions having different kernel, slice thickness and intervals can be reconstructed from the same acquisition (raw projection data) and therefore no additional radiation exposure is needed.

Reconstruction Kernel Characteristics shall be defined to optimize the analysis for each study. The same kernel must be used for each scan of a given subject and shall be consistent across all study participants. A softer kernel can reduce noise at the expense of spatial resolution. An enhancing kernel can improve resolving power at the expense of increased noise. Moderation on both fronts is recommended with a slight bias towards enhancement.

These parameters describe general characteristics of the reconstruction:

| Parameter | Compliance Levels |
| --- | --- |
| Spatial Resolution | Acceptable >= 6 lp/cm –OR– Axial FWHM <= 2.0mm Target >= 7 lp/cm –OR– Axial FWHM <= 1.0mm Ideal >= 8 lp/cm –OR– Axial FWHM <= 0.5mm |
| Noise | Acceptable Std. dev. in 20cm water phantom < 40 HU |
| Pixel Spacing | Acceptable <1.0mm Target 0.75 to 1.0mm Ideal <0.75mm |
| Reconstruction Interval | Acceptable Reconstruction intervals shall be less than half the axial extent of the mass and typically ≤2.5 mm Target Reconstruction intervals shall be less than one third the axial extent of the mass and typically ≤1.25 mm Ideal Reconstruction intervals shall be less than one tenth the axial extent of the mass and typically ≤.625 mm |
| Reconstruction Overlap | Acceptable 20% overlap Target 50% overlap |
| Reconstruction Kernel Characteristics | Acceptable Soft to overenhancing Target Standard to enhancing Ideal Slightly enhancing |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Actual Spatial Resolution, Noise, Pixel Spacing, Reconstruction Interval, Reconstruction Overlap, Reconstruction Kernel Characteristics, as well as the model-specific Reconstruction Software parameters utilized to achieve compliance with these metrics shall be recorded. |

## 4. Image Analysis

Each lesion shall be characterized as described in this section.  Lesions of interest include: a) small pulmonary Masses surrounded by air; b) small to medium pulmonary Masses surrounded by air and/or with adjacent normal and abnormal (non-neoplastic) anatomic structures; c) large pulmonary masses surrounded by air and/or with adjacent normal and abnormal (non-neoplastic) anatomic structures and/or confluent with mediastinum, chest wall, and diaphragm.

Procedures for segmenting or excluding tissue types and fluid, blood, necrotic debris within a mass are not described by this protocol, but may be implemented when technically feasible, in addition to measuring the entire volume within the outer tumor margin.

4.1 Methods to Be Used

Each lesion shall be characterized by determining the boundary of the lesion (referred to as segmentation), then computing the volume of the segmented lesion.  Segmentation may be performed automatically by a software algorithm, manually by a human observer, or semi-automatically by an algorithm working with human guidance/intervention.  The volume of the segmented region is then computed automatically.

| Parameter | Compliance Levels |
| --- | --- |
| Common Lesion Selection | Acceptable No requirement Target The software shall allow a common set of lesions to be designated for measurement, which are then subsequently measured by all readers Ideal The software shall detect and measure all measurable lesions automatically without the need for human intervention or multiple readers |
| Lesion Volume | Acceptable Shall be calculated as the sum of all the voxels within the boundaries of a discrete tumor mass on all the tomographic slices on which it is visible, regardless of its irregular shape. Target Shall be calculated without regard to spatial sampling loss (i.e., accounting by some means of interpolation for volume averaging due to non-isotropic voxel reconstruction and finite sampling). |
| Change Assessment Workflow | Acceptable Shall be performed as “locked sequential read”. |
| Multiple Lesions | Acceptable The software shall allow multiple lesions to be measured, and each measured lesion to be associated with a human-readable identifier that can be used for correlation across time points |
| Sum of Target Lesion Volumes | Acceptable A value computed by adding up all of the target lesion volumes calculated using Acceptable approach above shall be computed. Target A value computed by adding up all of the target lesion volumes calculated using Target approach above shall be computed. Ideal A value computed by adding up all of the target lesion volumes calculated using Ideal approach above shall be computed. |

For semi-automated or automated segmentation, the analysis software shall segment (based on a starting seed point/stroke/ROI) various types of tumors on CT images. The following further requirements are placed on image analysis software:

| Parameter | Compliance Levels |
| --- | --- |
| Boundary segmentation | Acceptable With many (> 50%) lesions requiring reader correction Target With few (< 10%) lesions requiring reader correction Ideal Fully automatically without reader correction |
| Automatically computed read-outs | Acceptable Automatic computation of volume of the segmented tumor shall be provided. Target Error margins for each measurement Provide a HU-histogram of the segmented voxels shall be provided. |
| Image Header Recording | Acceptable Software shall record in (and reload for review from) lesion segmentation boundary and volumetric measurement as well as metadata about reader identity, date and time and purpose of measurement. Target Software shall record in (and reload for review from) lesion segmentation boundary and volumetric measurement as well as metadata in standard formats including one or more of the following output formats: DICOM Presentation State, DICOM Structured Report; DICOM RT Structure Set; DICOM raster or surface segmentation. Ideal Software shall record in (and reload for review from) ALL of the Target formats. |

4.2 Required Characteristics of Resulting Data

It is expected that automated boundary detection algorithms will place segmentation edges with greater precision, accuracy and speed than an operator can draw by hand with a pointing device. The performance of the algorithms will, however, depend on the characteristics of the lesions may be challenged by complex tumors. Operator assisted semi-automatic segmentation shall produce at least the same level of intra- and inter-rater reliability for the volume measurements of each target lesion as manual segmentation.

| Parameter | Compliance Levels |
| --- | --- |
| Read-outs as described in Methods section | Acceptable Same precision, accuracy and speed than an operator can draw by hand with a pointing device shall be demonstrated. Target Greater precision, accuracy and speed than an operator can draw by hand with a pointing device shall be demonstrated. |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Annotation and Markup metadata | Acceptable Actual model-specific Analysis Software set-up and configuration parameters utilized to achieve compliance with these metrics shall be recorded. |

# IV. Compliance

## Acquisition Devices

Compliance to specifications as set out in the Image Acquisition section above.  Additionally, compliant Acquisition Devices shall provide means to record the information identified in the Subject Handling section as means to document compliance of the Performing Site to the specifications noted there.

## Reconstruction Software

Compliance to specifications as set out in the Image Reconstruction section above.  Additionally, compliant Reconstruction Software shall propagate the information collected at the prior Subject Handling and Imaging Acquisition stages and extend it with those items noted in the Reconstruction section.  See the compliance procedure notes associated with Acquisition Devices above for procedural assistance to identify Model Specific Parameters for Reconstruction Software.

## Software Analysis Tool

Compliance to specifications as set out in the Image Analysis section above.  Additionally, compliant Software Analysis Tools shall propagate the information collected at the prior Subject Handling, Imaging Acquisition, and Imaging Reconstruction stages and extend it with those items noted in the Analysis section

## Performing Site

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large patient population under consideration.  For imaging it is important to consider the availability of:

- appropriate imaging equipment and quality control processes,

- appropriate injector equipment and contrast media,

- experienced CT technologists for the imaging procedure, and

- processes that assure imaging protocol compliant image generation at the correct point in time.

A protocol specific calibration and QA program shall be designed consistent with the goals of the clinical trial. This program shall include (a) elements to verify that sites are performing the specified protocol correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address issues relating to both radiation dose and image quality (which may include issues relating to water calibration, uniformity, noise, spatial resolution -in the axial plane-, reconstructed slice thickness z-axis resolution, contrast scale, CT number calibration and others). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.

# References

[] Moertel CG, Hanley JA. The effect of measuring error on the results of therapeutic trials in advanced disease. Disease 1976; 38: 388-394.

[2] Quivey JM, Castro JR, Chen GT, Moss A, Marks WM. Computerized tomography in the quantitative assessment of tumour response. Br J Disease Suppl 1980; 4:30-34.

[3] Munzenrider JE, Pilepich M, Rene-Ferrero JB, Tchakarova I, Carter BL. Use of body scanner in radiotherapy treatment planning. Disease 1977; 40:170-179.

[4] Wormanns, D., Kohl, G., Klotz, E., Marheine, A., Beyer, F., Heindel, W., and Diederich, S. Volumetric measurements of pulmonary nodules at multi-row detector CT: In vivo reproducibility. Eur Radiol 14: 86–92, 2004.

[5] Kostis WJ, Yankelevitz DF, Reeves AP, Fluture SC, Henschke CI, Small Pulmonary Nodules: Reproducibility of Three-dimensional Volumetric Measurement and Estimation of Time to Follow-up CT, Radiology, Volume 231 Number 2, 2004.

[6] Revel M-P, Lefort C, Bissery A, Bienvenu M, Aycard L, Chatellier G, Frija G, Pulmonary Nodules: Preliminary Experience with Three-dimensional Evaluation, Radiology  May 2004.

[7] Marten K, Auer F, Schmidt S, Kohl G, Rummeny EJ, Engelke C, Inadequacy of manual measurements compared to automated CT volumetry in assessment of treatment response of pulmonary metastases using RECIST criteria, Eur Radiol (2006) 16: 781–790.

[8] Goodman, L.R., Gulsun, M., Washington, L., Nagy, P.G., and Piacsek, K.L. Inherent variability of CT lung nodule measurements in vivo using semiautomated volumetric measurements. AJR Am J Roentgenol 186: 989–994, 2006.

[9] Gietema HA, Schaefer-Prokop CM, Mali W, Groenewegen G, Prokop M, Pulmonary Nodules: InterscanVariability of Semiautomated Volume Measurements with Multisection CT— Influence of Inspiration Level, Nodule Size, and Segmentation Performance, Radiology: Volume 245: Number 3 December 2007.

[10] Wang Y, van Klaveren RJ, van der Zaag–Loonen HJ, de Bock GH, Gietema HA, Xu DM, Leusveld ALM, de Koning HJ, Scholten ET, Verschakelen J, Prokop M, Oudkerk M, Effect of Nodule Characteristics on Variability of Semiautomated Volume Measurements in Pulmonary Nodules Detected in a Lung Cancer Screening Program, Radiology: Volume 248: Number 2—August 2008.

[11] Zhao, B., Schwartz, L.H., and Larson, S.M. Imaging surrogates of tumor response to therapy: anatomic and functional biomarkers. J Nucl Med 50: 239–249, 2009.

[12] Hein, P.A., Romano, V.C., Rogalla, P., Klessen, C., Lembcke, A., Dicken, V., Bornemann, L., and Bauknecht, H.C. Linear and volume measurements of pulmonary nodules at different CT dose levels: Intrascan and interscan analysis. Rofo 181: 24–31, 2009.

[13] Mozley PD, Schwartz LH, Bendtsen C, Zhao B, Petrick N, Buckler AJ. Change in lung tumor volume as a biomarker of treatment response: A critical review of the evidence. Annals Oncology; doi:10.1093/annonc/mdq051, March 2010.

[14] Petrou M, Quint LE, Nan B, Baker LH. Pulmonary nodule volumetric measurement variability as a function of CT slice thickness and nodule morphology. Am J Radiol 2007; 188:306-312.

[15] Bogot NR, Kazerooni EA, Kelly AM, Quint LE, Desjardins B, Nan B. Interobserver and intraobserver variability in the assessment of pulmonary nodule size on CT using film and computer display methods. Acad Radiol 2005; 12:948–956.

[16] Erasmus JJ, Gladish GW, Broemeling L, et al. Interobserver and intraobserver variability in measurement of non-small-cell carcinoma lung lesions: Implications for assessment of tumor response. J Clin Oncol 2003; 21:2574–2582.

[17] Winer-Muram HT, Jennings SG, Meyer CA, et al. Effect of varying CT section width on volumetric measurement of lung tumors and application of compensatory equations. Radiology 2003; 229:184-194.

[18] Buckler AJ, Mozley PD, Schwartz L, et al. Volumetric CT in lung disease: An example for the qualification of imaging as a biomarker. Acad Radiol 2010; 17:107-115.

[19] AMERICAN COLLEGE OF RADIOLOGY IMAGING NETWORK, ACRIN 6678, FDG-PET/CT as a Predictive Marker of Tumor Response and Patient Outcome: Prospective Validation in Non-small Cell Lung Cancer, August 13, 2010.

[20] Miller AB, Hoogstraten B, Staquet M, Winkler A. Reporting results of cancer treatment. Cancer 1981;47:207-214.

[21] Eisenhauer EA, Therasse P, Bogaerts J, et al. New response evaluation criteria in solid tumors: Revised RECIST guideline (version 1.1). Eur J Cancer 2009;45:228-247.

[22] McNitt-Gray MF. AAPM/RSNA Physics Tutorial for Residents: Topics in CT. Radiation dose in CT. Radiographics 2002;22:1541-1553.

[23] Xie L, O'Sullivan J, Williamson J, Politte D, Whiting B,  TU‐FF‐A4‐02: Impact of Sinogram Modeling Inaccuracies On Image Quality in X‐Ray CT Imaging Using the Alternating Minimization Algorithm, Med. Phys. 34, 2571 (2007); doi:10.1118/1.2761438.

[24] Moertel CG, Hanley JA. The effect of measuring error on the results of therapeutic trials in advanced cancer. Cancer 38:388-94, 1976.

[25] Lavin PT, Flowerdew G: Studies in variation associated with the measurement of solid tumors. Cancer  46:1286-1290, 1980.

[26] Eisenhauera EA, Therasseb P, Bogaertsc J, et a. New response evaluation criteria in solid tumours: Revised RECIST guideline (version 1.1).  Eur J Cancer 2009; 45: 228-247.

# Appendices

## Acknowledgements and Attributions

This imaging protocol is proffered by the Radiological Society of North America (RSNA) Quantitative Imaging Biomarker Alliance (QIBA) Volumetric Computed Tomography (v-CT) Technical Committee. The v-CT technical committee is composed of scientists representing the imaging device manufacturers, image analysis software developers, image analysis laboratories, biopharmaceutical industry, academia, government research organizations, professional societies, and regulatory agencies, among others. All work is classified as pre-competitive. A more detailed description of the v-CT group and its work can be found at the following web link: http://qibawiki.rsna.org/index.php?title=Volumetric_CT.

The Volumetric CT Technical Committee (in alphabetical order):

- Athelogou, M. Definiens AG

- Avila, R. Kitware, Inc.

- Beaumont, H. Median Technologies

- Borradaile, K. Core Lab Partners

- Buckler, A. BBMSC

- Clunie, D. Core Lab Partners

- Cole, P. Imagepace

- Dorfman, G. Weill Cornell Medical College

- Fenimore, C. Nat Inst Standards & Technology

- Ford, R. Princeton Radiology Associates.

- Garg, K. University of Colorado

- Gottlieb, R. Roswell Park Cancer Center

- Gustafson, D. Intio, Inc.

- Hayes, W. Bristol Myers Squibb

- Hillman, B. Metrix, Inc.

- Judy, P. Brigham and Women’s Hospital

- Kim, HG. University of California Los Angeles

- Kohl, G. Siemens AG

- Lehner, O. Definiens AG

- Lu, J. Nat Inst Standards & Technology

- McNitt-Gray, M. University California Los Angeles

- Mozley, PD. Merck & Co Inc.

- Mulshine, JL. Rush

- Nicholson, D. Definiens AG

- O'Donnell, K. Toshiba

- O'Neal, M. Core Lab Partners

- Petrick, N. US Food and Drug Administration

- Reeves, A. Cornell University

- Richard, S. Duke University

- Rong, Y. Perceptive Informatics, Inc.

- Schwartz, LH. Columbia University

- Saiprasad, G. University of Maryland

- Samei, E. Duke University

- Siegel, E. University of Maryland

- Sullivan, DC. RSNA Science Advisor and Duke University

- Thorn, M. Siemens AG

- Yankellivitz, D. Mt. Sinai School of Medicine

- Yoshida, H. Harvard MGH

- Zhao, B. Columbia University

The Volumetric CT Technical Committee is deeply grateful for the support and technical assistance provided by the staff of the Radiological Society of North America.

## Background Information

QIBA

The Quantitative Imaging Biomarker Alliance (QIBA) is an initiative to promote the use of standards to reduce variability and improve performance of quantitative imaging in medicine. QIBA provides a forum for volunteer committees of care providers, medical physicists, imaging innovators in the device and software industry, pharmaceutical companies, and other stakeholders in several clinical and operational domains to reach consensus on standards-based solutions to critical quantification issues. QIBA publishes the specifications they produce (called QIBA profiles), first to gather public comment and then for field test by vendors and users.

QIBA envisions providing a process for developers to test their implementations of QIBA profiles through a compliance mechanism. After a committee determines that a profile has undergone sufficient successful testing and deployment in real-world care settings, it is released for use.  Purchasers can specify conformance with appropriate QIBA profiles as a requirement in requests for proposal. Vendors who have successfully implemented QIBA profiles in their products can publish conformance statements (called QIBA Compliance Statements) represented as an appendix called “Model-specific Parameters.”  General information about QIBA, including its governance structure, sponsorship, member organizations and work process, is available at http://qibawiki.rsna.org/index.php?title=Main_Page.

CT Volumetry for Cancer Response Assessment

Anatomic imaging using computed tomography (CT) has been historically used to assess tumor burden and to determine tumor response (or progression) to treatment based on uni-dimensional or bi-dimensional measurements. The original WHO response criteria were based on bi-dimensional measurements of the tumor and defined response as a decrease of the sum of the product of the longest perpendicular diameters of measured lesions by at least 50%. The rationale for using a 50% threshold value for definition of response was based on data evaluating the reproducibility of measurements of tumor size by palpation and on planar chest x-rays [24][25]. The more recent RECIST criteria introduced by the National Cancer Institute (NCI) and the European Organisation for Research and Treatment of Cancer (EORTC) standardized imaging techniques for anatomic response assessment by specifying minimum size thresholds for measurable lesions and considered other imaging modalities beyond CT. As well, the RECIST criteria replace longest bi-directional diameters with longest uni-dimensional diameter as the representation of a measured lesion [26]. RECIST defines response as a 30% decrease of the largest diameter of the tumor. For a spherical lesion, this is equivalent to a 50% decrease of the product of two diameters. Current response criteria were designed to ensure a standardized classification of tumor shrinkage after completion of therapy. They have not been developed on the basis of clinical trials correlating tumor shrinkage with patient outcome.

Technological advances in signal processing and the engineering of multi-detector row computed tomography (MDCT) devices have resulted in the ability to acquire high-resolution images rapidly, resulting in volumetric scanning of anatomic regions in a single breath-hold. Volume measurements may be a more sensitive technique for detecting longitudinal changes in tumor masses than reliance on linear tumor diameters as defined by RECIST. Comparative analyses in the context of real clinical trial data have found volume measurements to be more reliable and often more sensitive to longitudinal changes in response than the use of diameters in RECIST. As a result of this increased detection sensitivity and reliability, volume measurements may improve the predictability of clinical outcomes during therapy compared with RECIST. Volume measurements could also benefit patients who need alternative treatments when their diseases stops responding to their current regimens.

The rationale for volumetric approaches to accessing assessing longitudinal changes in tumor burden is multi-factorial. First, most cancers may grow and regress irregularly in three dimensions. Measurements obtained in the transverse plane fail to account for growth or regression in the longitudinal axis, whereas volumetric measurements incorporate changes in all dimensions. Secondly, changes in volume are less subject to either reader error or inter-scan variations. For example, partial response using the RECIST criteria requires a greater than 30% decrease in tumor diameter, which corresponds to greater than 50% reduction in volume of tumor. If one assumes a 21 mm diameter lesion (of 4850 mm3 volume), partial response would result require that the tumor shrink to a in a diameter of less than 158 mm, but which would correspond to a decrease in volume all the way down to 17702145 mm3. The much greater absolute magnitude of volumetric changes is potentially less prone to measurement error than changes in diameter, particularly if the lesions are irregularly shaped or spiculated. As a result of the observed increased sensitivity and reproducibility, volume measurements may be more suited than uni-dimensional measurements to identify early changes in patients undergoing treatment.

## Conventions and Definitions

Acquisition vs. Analysis vs. Interpretation: This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be computational steps that transform the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge. (The transformation of knowledge into wisdom is beyond the scope of this document.)

Bulls-eye Compliance Levels Acquisition parameter values and some other requirements in this protocol are specified using a “bulls-eye” approach. Three rings are considered from widest to narrowest with the following semantics:

ACCEPTABLE: failing to meet this specification will result in data that is likely unacceptable for the intended use of this protocol.

TARGET: meeting this specification is considered to be achievable with reasonable effort and equipment and is expected to provide better results than meeting the ACCEPTABLE specification.

IDEAL: meeting this specification may require unusual effort or equipment, but is expected to provide better results than meeting the TARGET.

An ACCEPTABLE value will always be provided for each parameter. When there is no reason to expect better results (e.g. in terms of higher image quality, greater consistency, lower dose, etc.), TARGET and IDEAL values are not provided.

Some protocols may need sites that perform at higher compliance levels do so consistently, so sites may be requested to declare their “level of compliance”. If a site declares they will operate at the TARGET level, they must achieve the TARGET specification whenever it is provided and the ACCEPTABLE specification when a TARGET specification is not provided. Similarly, if they declare IDEAL, they must achieve the IDEAL specification whenever it is provided, the TARGET specification where no IDEAL level is specified, and the ACCEPTABLE level for the rest.

Other Definitions:

Image Analysis, Image Review, and/or Read: Procedures and processes that culminate in the generation of imaging outcome measures, such tumor response criteria. Reviews can be performed for eligibility, safety or efficacy. The review paradigm may be context specific and dependent on the specific aims of a trial, the imaging technologies in play, and the stage of drug development, among other parameters.

Image Header: The Image Header is that part of the file or dataset containing the image other than the pixel data itself

Imaging Phantoms: Devices used for periodic testing and standardization of image acquisition. This testing must be site specific and equipment specific and conducted prior to the beginning of a trial (baseline), periodically during the trial and at the end of the trial.

Intra-Rater Variability is the variability in the interpretation of a set of images by the same reader after an adequate period of time inserted to reduce recall bias.

Inter-Rater Variability is the variability in the interpretation of a set of images by the different readers.

A Time Point is a discrete period during the course of a clinical trial when groups of imaging exams or clinical exams are scheduled as defined in the study protocol.

## Model-specific Instructions and Parameters

Compliance with a profile involves meeting a variety of requirements of which operating by these parameters is just one. To determine if a product (and a specific model/version of that product) is compliant, please refer to the Compliance section above.

Sites using models listed here are encouraged to consider using these parameters for both simplicity and consistency. Sites using models not listed here may be able to devise their own settings that result in data meeting the requirements but this is outside the formal scope of QIBA compliance.

In some cases, parameter sets may be available as an electronic file for direct implementation on the imaging platform.

Table G.1: Acquisition Device Model-specific Parameters Demonstrated to Achieve Compliance

IMPORTANT NOTE with respect to this example table: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  These settings were determined by the team in the 1C study as an example of how it could be done but more strict attention to all parameters identified in the Profile are necessary in order for a company to claim any particular model is compliant.  That said, we appreciate the good will and help that the vendors represented here have provided in this early phase of QIBA.

| Acquisition Device | Product Setting to Achieve Compliance Levels |
| --- | --- |
| GE Discovery HD750 sct3 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.625 Gantry Rotation Time in seconds 1 mA 120 Pitch 0.984 Scan FoV Large Body (500mm) |
| Philips Brilliance 16 IDT mx8000 | kVp 120 Number of Data Channels (N) 16 Width of Each Data Channel (T, in mm) 0.75 Gantry Rotation Time in seconds 0.75 Effective mAs 50 Pitch 1.0 Scan FoV 500 |
| Philips Brilliance 64 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.625 Gantry Rotation Time in seconds 0.5 Effective mAs 70 Pitch 0.798 Scan FoV 500 |
| Siemens Sensation 64 | kVp 120 Collimation (on Operator Console) 64 x 0.6 (Z-flying focal spot) Gantry Rotation Time in seconds 0.5 Effective mAs 100 Pitch 1.0 Scan FoV 500 |
| Toshiba Aquilion 64 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.5 Gantry Rotation Time in seconds 0.5 mA TBD Pitch .828 Scan FoV Medium and Large |

Table G.2: Reconstruction Software Model-specific Parameters Demonstrated to Achieve Compliance

IMPORTANT NOTE: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  These settings were determined by the team in the 1C study as an example of how it could be done but more strict attention to all parameters identified in the Profile are necessary in order for a company to claim any particular model is compliant.  That said, we appreciate the good will and help that the vendors represented here have provided in this early phase of QIBA.

| Reconstruction Software | Product Setting to Achieve Compliance Levels |
| --- | --- |
| GE Discovery HD750 sct3 | Reconstructed Slice Width, mm 1.25 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel STD |
| Philips Brilliance 16 IDT mx8000 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm (contiguous) Display FOV, mm 350 Recon kernel B |
| Philips Brilliance 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm (contiguous) Display FOV, mm 350 Recon kernel B |
| Siemens Sensation 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel B30 |
| Toshiba Aquilion 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm Display FOV, mm TBD Recon kernel FC12 |

Table G.3: Image Analysis Software Model-specific Parameters Demonstrated to Achieve Compliance

IMPORTANT NOTE: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  In particular, the following example table only has placeholders for these example products which need to be replaced with product model-specific settings in order to claim compliance.

| Image Analysis Software | Product Setting to Achieve Compliance Levels |
| --- | --- |
| Siemens LunCARE | a <settings to achieve…> b <settings to achieve…> c <settings to achieve…> d <settings to achieve…> |
| GE Lung VCAR | e <settings to achieve…> f <settings to achieve…> g <settings to achieve…> h <settings to achieve…> |
| R2 ImageChecker CT Lung System | i <settings to achieve…> j <settings to achieve…> k <settings to achieve…> l <settings to achieve…> |
| Definiens (name specific product) | m <settings to achieve…> n <settings to achieve…> o <settings to achieve…> p <settings to achieve…> |
| Median (name specific product) | q <settings to achieve…> r <settings to achieve…> s <settings to achieve…> t <settings to achieve…> |
| Intio (name specific product) | u <settings to achieve…> v <settings to achieve…> w <settings to achieve…> x <settings to achieve…> |