## Results and Resolutions
Commenter | Section | Line # | Priority | Issue | Proposal | Committee Discussion | Owner (opt.) | Resolution (w Rationale if rejected) | Status
[8 empty] | 1 | TBD | To be decided
See http://qibawiki.rsna.org/index.php/Public_Comment_Process for more guidance on the comment resolution process | [7 empty] | 0 | OK | No action requested
[8 empty] | 0 | Discuss | Need to decide resolution
[8 empty] | 0 | TODO | Resolution decided
[8 empty] | 0 | Done | Profile update completed
|  | Use PDF lines, they're stable |  | Describe your issue. Don’t write a book, but do include enough to indicate what you see as a problem. | Suggest new wording or describe a way to address the issue. The committee may simply accept your suggested text. Even if they don’t, it gives a good sense of what you’re looking for. Leaving this blank means you can’t imagine how to resolve the issue | <Cmte Notes as needed.> <Optionally, use the Owner column to divide up the work and assign rows to a committee member who will lead discussion and resolution> | <Delete this column if not used> | Describe how the comment was/will be resolved. May be simply accepted & changed as proposed, may be accepted & resolved differently, or may be rejected with a rationale for why. | TBD
(Esaote) Gian Matteo Morano | Appendix D | 1760 | M | Esaote (ultrasound manufacturer) would like to contribute to model specific insturctions and parameters (Appendix D); a model draft of 4 machines included. | Model specific instructions for Appendix D are included for consideration
Riwa Kishimoto, MD, PhD | 2. Clinical Context & Claims | 215 | L | Clinical Context and Claims, is inter-operator variability (different technologists with the same scanner at the same point) regarded as claim 4 (longitudinal claim)? Our previous study showed a similar 95% limit of agreement for inter-operator reproducibility with Table 2-4, such as -21.3~23.8% for 4cm, -29.0~25.6% for 6cm and -31.7~30.8% for 8cm depth of right lobe of the liver in healthy volunteers (SWS:1.06-1.13m/s). We thought our 95% LOA were relatively large, but I think our results are consistent with this claim. Article Referenced: Kishimoto R, et al Eur Radiol. 2019 Nov;29(11):5999-6008. doi: 10.1007/s00330-019-06195-8. Intra- and inter-operator reproducibility of US point shear-wave elastography in various organs: evaluation in phantoms and healthy volunteers. | none listed
Paul Barbone | N/A |  | L | Recommend center frequency of SW pulse (and bandwidth) be reported with each SWS measurement. Individual platforms may have different center freq ie. 100 Hz, 300 Hz, it is useful to know if measured at the same center freq or not. | none listed
Karen Alton | Appendix D | 1265 | L | "See Specifications in Profile Section 3.6, 3.8, 3.10" | Section 3.10 was removed. Propose change to "…Profile Sections 3.7, 3.8"
(Canon) Todd Erpelding, PhD | Appendix D | 1250 | M | Canon “Model-Specific Instructions” are not up-to-date with the latest version. | Update Canon system and transducer information to include our latest releases (Aplio 450/550 and some new transducers). Addition line 1250: "Aplio a-series (a450/a550/a)"
| Appendix D | 1256 | M | Add new software version | Addition line 1256: "Aplio a450/a550/a V1.0 or later"
| Appendix D | 1260 | M | Add 1 machine and 2 transducers to the Canon Transducer Chart | Additon line 1260: Column 3 out of 5 "Aplio a450/a550/a", Row 3 out of 8: "PVI-574BX", Row 7 out of 8 "PVT-574BT"
Andy Milkowski | Appendix D | 1563 | M | New Equipment | Addition line 1564: "ACUSON Redwood"
| Appendix D | 1563 | M | New Equipment | Addition line 1565: "ACUSON Juniper"
| Appendix D | 1568 | M | New Software | Addition line 1568: "ACUSON Redwood: VA10A or later"
| Appendix D | 1568 | M | New Software | Addition line 1569: "ACUSON Juniper: VA10A or later"
| 3.1 Pre-delivery | 377-378 | M | "Need additional clarity regarding frequency of pre-delivery testing" | "Clarify that manufacturer needs to perform this testing and document."
(No Name) collegue of Riwa Kishimoto, MD, PhD | Closed Issues | 105 | L | "There are many places without space between the number and the unit." "Page 7, Q. Give stiffness in m/sec or kPa? A. M/sec -> m/sec" | (I think he's suggesting that the Capital M should be a lowercase in the Answer line ??? -KA)
David T. Fetzer, MD | Closed Issues | 150 (& 620) | L | "Tidal respiration" may be confusing | "Consider end inspiration or end expiration"
| Closed Issues - # Measurements | 150 | M | "I do not believe there is enough literature to support obtaining only 5 measurements." | "Focus on 10 measurements until more literature is available."
| Closed Issues - # Measurements | 150 | H | "The IQR/Median of 0.3 is if measurements are being reported in stiffness (kPa). If m/s, then this cut-off is actually 0.15." | none listed
| Closed Issues - Max Liver Depth | 150 | M | "Some recommendations report a distance of 1.5 cm from the capsule, instead of 2.0 cm." | none listed
| 2.1 Proposed Clinical Interpretation | 256-271 | M | "The SRU consensus panel has submitted an update for publication that changes the docus from staging fibrosis (eg. Metavir F0 or F4), to conveying a risk of clinically-significant fibrosis (compensated advanced chronic liver disease)." | none listed
| 3.3 Staff Qualification | 527 | L | "How was 20 chosen? What if a site has either a large number of sonographers or a small number of elastography requests that would make this requirement impractical. Also, what if a site is unable to obtain an elastic phantom?" | none listed
| 3.5 Subject Selection | 574 | M | "For an FDA-approved clinical technique, why would informed consent be needed?" | none listed
| 3.6 Subject Handling | N/A | M | "Would it be prudent for an actor to verify other potential confounders such as presence of markedly elevated AST/ALT, evidence of biliary obstruction, and evidence of right heart failute/passive hepatic congestion?" | none listed
| 3.7.2 Transducer Position | 723 | M | "Why not include hepatic segment 6? In some patterns of hypertrophy/atrophy, 6 is quite prominent, and is easily imaged, esp when the patient is 30 degrees semi-decub." | none listed
(AAPM) Cynthia McCollough, PhD, FAAPM, FACR, AFIMBE; Response replies to Richard J. Martin, JD | General Remark | N/A | N/A | "Essentially, the liver is not as homogenous as we had hoped. Considering the current recommendation of sampling just one area, AAPM asks what do the committee members think about the limited sampling issue and how do the users know which area is the best to measure?"
| General Remark | N/A | N/A | Have any members been struggling with significant underestimation in stiffness with US shear wave elastography (SWE) compared to magnetic resonance elastography (MRE)? Part of the value of imaging methods of assessing liver fibrosis compared to biopsy approaches is that imaging methods allow evaluation of the entire liver. Sampling only one area of the US image discards important information about heterogeneity in the liver.
| General Remark | N/A | N/A | Discussion of pediatric patients is limited. AAPM asks whether pediatric patients have been adequately considered in this Profile.
| Closed Issues | 105,(see also P20 L381 et seq., P21 L415 et seq.) | N/A | The specifications of the phantom are not consistent. Attenuation is given as 0.6 or 0.5 dB/cm/MHz with a standard deviation of 0.1 or 0.2. In addition, speed of sound value and tolerance range are also not consistent.
| Closed Issues | 105 (See also P26 L557, P30 L659, P32 L725) |  | There are several places talking about measurements being 2 cm below the liver capsule and within 6.5 cm from the probe surface. It would be helpful to define this depth: Is it the center/top/bottom of the ROI within the point SWE and 2D SWE image? A definition similar to that given on page 37 would be helpful.
| 2. Clinical Content | 161 |  | The phrase "moderate but significant liver fibrosis" should be changed to “moderate to significant.”
| 2. Clinical Content & Claims | 228-230 and 246-247 |  | Measurements in viscoelastic medium can be quite different for scanners from different vendors. The cut points in the table of page 15 for System A and System B seems to differ by 20% already (also see different cut points given on page 51 and page 56 for two vendors), which seems to suggest that the scannerdifference for in vivo measurement can be as high as 20%. However, some reproducibility coefficients in tables 2-3 and 2-4 are below 20% for the claim of using different ultrasound scanners. Therefore, some clarifications may be needed to address this gap.
| 2.1 Proposed Clinical Interpretation | 258 |  | Replace missing period mark.
| 2.1 Proposed Clinical Interpretation | 279 |  | Closing parenthesis is missing and the line should be re-aligned.
| 2.2 Discussion | 294 |  | Should remove the words 'or not'. The phrase is used two times.
| 3.1.1 Pre-delivery | 395 |  | Do the phantom specifications make sense for a pediatric patient as well (attenuation & backscatter)?
| 3.1.1 Pre-delivery | 434 |  | Are there any calibration specifications for a scale that would weigh the phantom?
| 3.1.1 Pre-delivery | 466 |  | It may be helpful to add discussions of the impact of extreme temperature during shipping of calibration phantoms and any precautions to take during shipping.
| 3.1.2 Phantom Testing | 473 |  | The link is broken; cannot evaluate the AIUM QA manual for Gray Scale US Scanners.
| 3.1.3 Specification | 482 |  | Are the MFR specifications and certification available to the user? Only available upon request?
| 3.5.1 Subject Selection | 557 |  | >=2cm distance from the liver capsule seems to be different compared to most existing literatures which state ~1.5-2cm. Are there any data to support this decision? In addition, how to handle the inconsistency with the vendor recommendations? For example, >=1cm from Canon (page 48).
| 3.5.1 Subject Selection | 557 |  | Is the maximum 6.5cm recommendation corresponding to the lower edge of the acquisition ROI? What about patients with ascites? Because of the lower attenuation, many of these cases appear to have sufficient signal at 6.5cm or slightly deeper, even though they are more susceptible to motion artifacts.
| 3.7.1 SWS Image Acquisition | 611-614 |  | The efficacy of phantom calibration may depend on whether the phantom is viscoelastic or elastic. Elastic phantom may not be able to provide a good calibration among different scanners for in vivo liver measurements. Some comments about phantom properties may be helpful here.
| 3.7.1 Transducer Pressure | 651 |  | Newer data seem to suggest that stronger pressure is fine for intercostal acquisition, which actually decreases the acquisition failure rate without changing the quantification results (Byenfeldt, et al. 2019 Ultrasound in Med & Biol). One manufacturer suggests using this approach. Obviously, it would be different for subcostal acquisition.
| 3.7.2 Specification | 725 |  | Under parameter/number of measurements, "Shall make a minimum of 5 measurements should be made" should be changed to “Shall make a minimum of 5 measurements.”
| 4.1.2 Image Performance | 772-773 |  | The link is broken.
| 4.2.1 Site Assessment Tools & Tests | 809 |  | Are there any pediatric-specific phantoms that are required during this evaluation? It states only that two phantoms will be used.
| 4.2.1.2.B | 945 |  | There needs to be a closing parenthesis to end the sentence.
| Appendix D | 1300 et seq. |  | For GE LE9 users, it is extremely important to maintain the elastography gain at 55. It is inherently tied to a quality threshold. It's an easy change for the operators to make, which could actually alter the quantitative measurements. Similar requirement exists on S8 with a different gain value.
Note on A10: Alton, Karen *: QST Hospital, National Institute of Radiological Sciences 4-9-1, Anagawa, Chiba 262-0033, Japan
Note on A21: Alton, Karen *: QST Hospital, National Institute of Radiological Sciences 4-9-1, Anagawa, Chiba 262-0033, Japan
Note on A31: Alton, Karen *: Cynthia McCollough President, AAPM Brooks-Hollern Professor Professor of Medical Physics and Biomedical Engineering Director, CT Clinical Innovation Center and X-ray Imaging Core Department of Radiology, Mayo Clinic