FDG-PET/CT as an Imaging Biomarker Predicting Response to Cancer Therapy for Patient Management and Oncologic Drug Development

Version 0.2

31 January 2011

Table of Contents

I. Executive Summary	3

II. Clinical Context and Claims	4

III. Profile Details	6

0. Reserved	6

1. Reserved	6

2. Reserved	6

3. Subject Scheduling	6

4. Subject Preparation	7

5. Imaging-related Substance Preparation and Administration	13

6. Individual Subject Imaging-related Quality Control	18

7. Imaging Procedure	19

8. Image Post-processing	30

9. Image Analysis	32

10. Image Interpretation	36

11. Archival and Distribution of Data	40

12. Quality Control	43

13. Imaging-associated Risks and Risk Management	48

IV. Compliance	51

References	52

Appendices	53

Appendix A: Acknowledgements and Attributions	53

Appendix B: Background Information	53

Appendix C: Conventions and Definitions	53

Appendix D: Documents included in the imaging protocol (e.g., CRFs)	54

Appendix E: Associated Documents	54

Appendix F: TBD	55

Appendix G: Model-specific Instructions and Parameters	55

# I. Executive Summary

This Profile documents specifications and requirements for the use of FDG-PET as an imaging biomarker in oncologic indications. It covers both clinical trial usage as well as individual patient management. The document is built in part from extracting relevant information from its sister UPICT protocol as indicated below as well as directly through the efforts of the QIBA FDG-PET Technical Subcommittee.

Summary of Clinical Trial Usage as described in "Consensus Protocol for FDG-PET/CT as an Imaging Biomarker Predicting Response to Cancer Therapy and Oncologic Drug Development "

(draft pending completion of the extraction process) – This UPICT Protocol is intended to guide the performance of whole-body FDG-PET/CT within the context of single- and multi-center clinical trials of oncologic therapies by providing acceptable (minimum), target, and ideal standards for all phases of the imaging examination as defined by the UPICT Template (ref) with the aim of minimizing intra- and inter-subject, intra- and inter-platform, inter-examination, and inter-institutional variability of primary and/or derived data that might be attributable to factors other than the index intervention under investigation. The specific potential utilities for the FDG-PET/CT study(ies) as performed in accordance with this Protocol within any particular clinical trial could be to utilize qualitative, semi-quantitative, and/or quantitative data for single time point assessments (e.g., diagnosis, staging, eligibility assessment, investigation of predictive and/or prognostic biomarker(s)) and/or for multi-time point comparative assessments (e.g., response assessment, investigation of predictive and/or prognostic biomarker(s)). More generally, such standardization of FDG-PET/CT within the conduct of clinical trials should 1) support internal decision-making in drug, biologic, and device development, 2) provide data to support registration and market-label indications, and 3) allow the eventual qualification of one or more imaging biomarkers (perhaps as surrogates for clinical endpoints) by supporting meta-analyses of multiple clinical trials (possibly over different compounds or devices and as contributed by different companies).

This document does include specifications for the performance of CT for the purposes of attenuation correction and/or localization, but does not address the performance of diagnostic CT within the context of FDG-PET/CT; although the integration of diagnostic CT in conjunction with FDG-PET/CT for oncology is acknowledged as potentially useful and appropriate. When the integration of diagnostic CT is desired as part of the imaging protocol within the clinical trial, specifications for the CT portion of the imaging protocol may be derived from other UPICT protocol(s).

# II. Clinical Context and Claims

The clinical context sets out the utilities and endpoints for clinical trial usage and then proceeds to identify targeted levels of accuracy for named measurement read-outs that may be used in the clinical trial setting and/or individual patient management.

Utilities and Endpoints for Clinical Trials

Consolidated Statement – The specific utility(ies) for which the FDG-PET/CT imaging study(ies) in the clinical trial authored using this UPICT Protocol should be clearly stated within the documentation of the clinical trial. This Protocol has been derived from various referenced standards documents and publically listed clinical trials. The specific utilities for the FDG-PET/CT imaging stated in those source materials include:

• diagnosis and staging of tumors (EU, Neth, ACRIN protocol #s 6671, 6685)

• prognostic stratification / biomarker (Neth, Hallett, ACRIN protocol # 6685)

• treatment planning or triage (ACRIN protocol # 6685)

• edge detection of tumors in radiotherapy planning (EU)

• lesion localization and characterization (EU, ACRIN protocol #s 6671, 6685)

• evaluate tumor response / predictive stratification / biomarker (EU, Neth, Hallett, ShankarNCI, ACRIN protocol #s 6665, 6678)

• correlation between imaging and tissue biomarkers and/or pathway activity (ACRIN protocol # 6665)

Claim 1

FDG-PET scans are sensitive and specific for detection of malignant tumors. FDG-PET scans reliably reflect glucose metabolic activity of cancer cells and can be measured with high reproducibility over time. Longitudinal changes in tumor FDG activity during therapy predict clinical outcomes (define examples?) earlier than changes in standard anatomic measurements. Therefore, tumor response or progression as determined by tumor FDG activity will be able to serve as an endpoint in well-controlled Phase II and III efficacy studies of cytotoxic and targeted therapies in FDG-avid tumors. In tumor/drug settings where the preceeding phase II trials have shown a statistically significant relationship between FDG-PET response and an independent measure of outcome, changes in tumor FDG activity can then serve as the primary endpoint for regulatory drug approval in registration trials.

Profile specified for use with: **target subpopulation, for the following indicated biology: **indicated biology, and to serve the following purpose: **biomarker purpose.

Obtain actual maximum  activity concentration or voxel value (units Bq/ml) and maximum SUV (BW, LBM, BSA normalised), which are the activity concentration value and the SUV of a single voxel from the actual original image that has the highest value within a volume of interest drawn around the object of interest. The latter will avoid additional error in activity concentration values and SUVmax due to viewing/display interpolation and/or minimize round-off errors in SUV calculation within defined tolerances.

Compliance Levels for Measurement Read-outs

| Measurement or Categoric Result | Performance Levels Achieved under Bull's Eye Conditions |
| --- | --- |
| SUVmax | If Activities are Performed at Acceptable level If Activities are Performed at Target Level <0.1% If Activities are Performed at Ideal Level 0% |
| **for example, PR | If Activities are Performed at Acceptable level **e.g., coef. of corrlelation 85% If Activities are Performed at Target Level **e.g., coef. of correlation 90% If Activities are Performed at Ideal Level **e.g., coef. of correlation 95% |
| **for example, presence of carcinoma | If Activities are Performed at Acceptable level e.g., AUC of ROC 80% If Activities are Performed at Target Level e.g., AUC of ROC 87% If Activities are Performed at Ideal Level e.g., AUC of ROC 90% |

# III. Profile Details

## 0. Reserved

## 1. Reserved

## 2. Reserved

## 3. Subject Scheduling

Considerations and specifications related to subject scheduling are drawn from the protocol in the following sections.

Consolidated Statement: Prior to scheduling potential and/or already accrued subjects for FDG-PET/CT with its inherent (albeit minimal) risks confirmation of appropriateness for imaging (e.g., history, physical examination, staging, biopsy for diagnosis, etc.) should be performed and documented. Scheduling diabetic subjects may require special attention (please see Section 4.2.2 for additional details) and therefore this should be specifically queried at the time of scheduling. At the time of scheduling, the study team should determine that inclusion of the subject does not violate any of the study-specific inclusion and exclusion criteria pertinent to the FDG-PET/CT study. (SNM GHS) For considerations related to the scheduling of subjects who are known to be diabetic please also see Sections 1.7.2 and 4.2.2.

3.1. Timing Relative to Index Intervention Activity

The following requirements are placed regarding timing relative to index intervention activity.

| Index Intervention Activity | Timing |
| --- | --- |
|  | Acceptable Target Ideal |

3.2. Timing Relative to confounding Activities (to minimize “impact”)

Consolidated Statement – Activities, tests, and interventions that might increase the chances for false positive and/or false negative FDG-PET/CT studies should be avoided prior to scans. The allowable interval between the potentially confounding event and the imaging test will be dependent on the nature of the confounder. For example, a percutaneous or excisional biopsy of a suspicious mass may cause focally increased FDG-PET activity or might lead to the appearance of a non-malignant mass (e.g., hematoma) on the CT portion of the study. A percutaneous ablation procedure of a known malignant focus may cause focally increased FDG-PET activity and/or an immediate post-ablation increase in the apparent volume of the ablation target lesion. The time of onset and the duration of the increased FDG-PET activity and/or the change in lesion volume might be different for these two different confounding factors.

If iodinated contrast is to be used for the CT portion of the PET/CT study, conflict with other tests and treatments should be avoided congruous with community standards of care (e.g., thyroid scan).

| Confounding Activity | Timing |
| --- | --- |
|  | Acceptable Target Ideal |

3.3. Scheduling Ancillary Testing

Consolidated Statement – Avoid scheduling tests that might confound the qualitative or quantitative results of the FDG-PET/CT study within the time period prior to the scan. For example, a glucose tolerance test should not be scheduled during the 24 hours prior to the performance of FDG-PET/CT. Similarly, other tests that might involve increasing plasma glucose, insulin, or corticosteroid levels should also be avoided. Exercise cardiac stress testing should be avoided during the twenty-four (24) hours prior to the performance of FDG-PET/CT. Similarly, other tests that might involve vigorous exercise and thereby increase muscle metabolic function should also be avoided.

| Ancillary Test | Scheduling |
| --- | --- |
|  | Acceptable Target Ideal |

## 4. Subject Preparation

The following sections describe how subjects are prepared.

4.1. Prior to Arrival

Consolidated Statement - The main purpose of subject preparation is to reduce tracer uptake in normal tissue (kidneys, bladder, skeletal muscle, myocardium, brown fat) while maintaining and optimizing tracer uptake in the target structures (tumor tissue). Below is a generally applicable protocol to address (1) Dietary, (2) Fluid Intake and (3) Other activities that may impact the FDG-PET/CT procedure or results. (1) Dietary (for the management of previously known or unknown diabetic subjects please see section 4.2.2):

- According to two sources, subjects should not eat any food for an absolute minimum (acceptable level) of 4 hours prior to start of FDG-PET study (ACRIN), although the target pre-test fasting period is recommend as a 6 hour minimum (Neth, EU). This can be achieved as follows:

o Subjects scheduled to undergo the PET study in the morning should not eat after midnight and preferably have a light meal during the evening prior to the PET study.

o Subjects scheduled for an afternoon PET study may have a light breakfast before 8 am.

o Medication can be taken as prescribed (see Section 4.2.2 for diabetic management)

o Two sources have stated that a low carbohydrate diet should be followed for 24 hours before the study, culminating with fasting for the final six hours. (NCI, ACRIN)

o One large study (ref) has suggested that a high-fat, low-carbohydrate meal is preferred for the last meal prior to commencing the period of fasting; as well one study has suggested that beta blockade may be useful to decrease brown fat uptake. Although there are insufficient data to recommend these strategies as routine at this time (SNM GHS)

- However, on the basis of the SNM harmonization summit the acceptable and target timing for discontinuing enteral nutrition is at least six (6) hours prior to the anticipated time of FDG administration. (SNM GHS)

(2) Fluid Intake:

-All references note that adequate hydration (before and after FDG administration) is important (both to ensure a sufficiently low FDG concentration in urine (less artifacts) and for radiation safety reasons). Whichever hydration strategy is used (how much and when to administer), the protocol should be uniform among sites during a trial. Specific hydration recommendations include: one reference (NCI) suggests oral intake of at least 710-1665 mls of water while fasting, an additional reference (ACRIN) recommends consumption of two to three 8-12 oz water (??? = 473-1065 mls) while fasting, and two references (Neth and EU) suggest 1 liter during 2 hours prior to FDG administration. However, on the basis of the SNM harmonization summit the acceptable and target hydration for studies without iodinated contrast material is at least 500ml PO or IV in the two hours prior to FDG administration and for studies utilizing iodinated contrast 1L PO or IV during the two hour period. The fluid administered should not contain glucose or caffeine. (SNM GHS)

-While two sources (EU, ACRIN) indicates that parenteral nutrition and intravenous fluids containing glucose should be discontinued at least 4 (acceptable) - 6 (target) hours before the PET examination. The infusion used to administer intravenous prehydration must not contain any glucose. However, on the basis of the SNM harmonization summit, the acceptable and target are both set to six (6) hours in the consensus protocol. (SNM GHS)

(3) Other Activities:

- The subject should avoid strenuous or extreme exercise before the PET exam to minimize uptake of the radiotracer in muscle (ACRIN), for a minimum acceptable time period of at least 6 hours (EU) but preferably for the target time period of 24 hours (NCI, EU, ACRIN) prior to the PET exam. Based on the SNM harmonization summit the acceptable and target time for avoidance of strenuous physical activity and cold exposure is 24 hours prior to the study. Other activities that might be avoided are contained in sections 3.2 and 3.

Insert screening for claustrophobia here with ACRIN attribution. Insert screening and pretreatment for iodinated contrast issues here with ACRIN attribution – separate paragraph.

| Preparation Step | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

4.2. Upon Arrival

The following sections describe steps taken upon arrival.

4.2.1. Confirmation of subject compliance with instructions

**Add any needed introduction or replace the following content from the Protocol.

Consolidated Statement – Upon arrival 1) confirmation of subject compliance with pre-procedure instructions and 2) the occurrence of potentially confounding events should be documented on the appropriate case report forms. The documentation might include some or all of the following:

• timing, character, and amount of the most recent previous oral and/or intravenous intake of fluid and nutrients

• timing and dosages of relevant non-prescription and prescription medications taken prior to the PET/CT scan (e.g., the last cycle of chemotherapy or non-cytotoxic pharmacotherapy, administration of growth factors, cytokines, steroids, beta blockers, etc.)

• extent of physical activity and most recent exposure to cold temperature for the preceding 24 hours

• timing and description of medical procedures performed prior to the PET/CT scan (e.g., radiation therapy, biopsy, surgery)

• timing and description of relevant medical tests performed prior to the PET/CT scan (e.g., invasive tests and/or tests that involve the administration of exogenous substances and/or tests that involve vigorous physical activities)

• confirmation that the subject has completed the trial Informed Consent Document.

The FDG-PET/CT procedure should be explained to the subject and exam-specific consent should be obtained if that is the standard of care for the site or the standard established for the specific clinical trial. There should be documentation of subject-specific risk factors including, but not limited to, previous contrast reactions (if iodinated contrast is to be used).

| Instruction | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

4.2.2. Ancillary Testing

*Consolidated Statement: - Subject height and body weight must be measured precisely with standardized measurement devices and with the subject in a gown or light clothing and recorded as the minimum acceptable standard (EU, Neth, NCI, ACRIN). The target standard would add that for serial studies in the same subject, weight should be measured directly prior to each PET study since body weight often changes during the course of the study (EU, SNM GHS).  Blood glucose monitoring, measurement and documentation and the appropriate management/disposition of hyperglycemic/ diabetic subjects are addressed by all references and should be included as a minimum acceptable standard of performance.

- It is important to measure and document subject blood glucose level shortly prior to and target within the 2 hours prior to (ideally within 1 hour for all subjects and target within 1 hour for insulin-requiring diabetic subjects) FDG administration (all, SNM GHS).

- The ideal is blood glucose level < 120 mg/dL (all).

- Subjects with blood glucose level > 200 mg/dL should be rescheduled (all).

- For subjects with blood glucose measurements between 120 mg/dL and 200mg/dL, there are varying actions suggested by the different references.

o Two references (EU, ACRIN) indicate that the subject must be rescheduled, and adjustments to diet and medications made if necessary, so that the fasting blood glucose concentration can be brought down to the acceptable range at the time of FDG injection, or excluded depending on the subject circumstances and the trial being conducted.

o One reference (NCI) indicates that for diabetic subjects, the allowable glucose reference range is up to 150-200mg/dL. The handling of non-diabetic subjects is not specifically addressed.

o One reference (Neth) indirectly indicates that for diabetics (and non-diabetics?) with glucose levels between 120-200 mg/dl, PET scanning can proceed

o There is no consensus from these references for diabetic or non-diabetic subject management in the glucose range of 120-200 mg/dL. The imaging protocol for each individual clinical trial should indicate the glucose cut-off thresholds and the exact management for diabetic and non-diabetic subjects with plasma glucose levels between 120-200 mg/dl, especially if the data from the FDG-PET/CT examination will be used towards a primary or secondary endpoint and/or will be compared in a serial manner over the course of the protocol.

• Secondary to recognized problems with administration of insulin (due to alteration of FDG biodistribution and diminished accuracy of SUV determination-NCI), insulin must not be given to reduce pre-FDG-administration glucose levels (EU, NCI), unless the interval between administration of insulin and FDG is more that 4 hours (EU). One reference (ACRIN) suggests that if a hyperglycemic subject is to be rescheduled, the subject’s physician should be contacted to manage the blood glucose prior to the rescheduled study.

• Additional scheduling recommendations for diabetic subjects are suggested by two references (EU and Neth). These include the following:

o For type I diabetes:

 Ideal to achieve euglycemia prior to PET study

 Schedule study for late morning by eating normal breakfast at 7 am and taking normal amount of insulin; then fast for at least 4 hours till exam

o For type II diabetes:

 Schedule study for late morning

 Comply with at least 4 hour fast till exam  Continue oral medication (hypoglycemic) as usual

- One reference (ACRIN) suggests the following for diabetic management:

o Diabetic subjects should be scanned early in the morning before the first meal, and doses of insulin and/or hypoglycemic medication should be titrated appropriately in consultation with the subject's referring physician.

o Before scheduling an FDG-PET study, diabetics subjects should test their ability to maintain reasonable plasma glucose levels after fasting, while avoiding insulin close to the time that FDG would be administered.

| Ancillary Test | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

4.2.3. Preparation for Exam

Consolidated Statement – In order to avoid artifactual distribution of the FDG, it is critical that subject preparation after arrival and prior to imaging are standardized among all sites and subjects throughout the conduct of the clinical trial. (EU, Neth, NCI, Hallet, ACRIN)

• The waiting and preparation rooms should be relaxing and warm (> 75° F or 22° C) during the entire uptake period (and for as long as reasonably practicable prior to injection, at least 15 minutes is suggested as acceptable). Blankets should be provided if necessary. (SNM GHS)

• The subject should remain recumbent or may be comfortably seated; activity and conversation should be kept to an absolute minimum. For example, the subject should be asked to refrain from speaking, chewing, or reading during the uptake period (SNM GHS). For brain imaging the subject should be in a room that is dimly lit and quiet for FDG administration and subsequent uptake period (ACRIN).

• The subject may use the rest room, but if possible not for the 30 minutes immediately after injection of FDG. The subject should void immediately (5 – 10 minutes) prior to the FDG-PET/CT image acquisition phase of the examination (SNM GHS).

• Bladder catheterization is not routinely necessary; but if necessary the catheter should be placed prior to injection of FDG. Bladder catheterization may be important for the evaluation of pelvic tumors (e.g., cervix or prostate cancer).

• Following the administration of FDG, the subject should drink 500 ml (or 8 – 12 oz 237-354 ml per ACRIN) of water (or receive by intravenous administration 250 - 500 ml of non-glucose containing fluid). Fluid intake may need to be modified for those subjects on fluid restriction.

• For specific areas of anatomic interest (e.g., tumors located in the lower abdomen, pelvis or kidney) intravenous diuretic agents may be used (e.g., 20 – 40 mg of furosemide given nearly contemporaneously (within 10 – 15 minutes) with the administration of FDG). Per the SNM harmonization summit if bladder catheterization is performed IV diuretics should be administered as described herein so as to ensure that the concentration of activity in the renal collecting systems and bladder is relatively dilute.

• Sedation is not routinely required, but is not contraindicated provided that the sedative used does not interfere with the uptake of FDG. If sedation might be used, the subject was be instructed in advance that operation of a motorized vehicle will be prohibited after the FDG-PET/CT test. Sedation may have utility in specific clinical circumstances such as brain or head and neck tumors, claustrophobic subjects, or children.

• The amount of fluid intake and use of all medications (e.g., diuretic, sedative) must be documented on the appropriate case report form.

• Subjects undergoing a CT scan should empty their pockets and remove any clothing containing metal and any metallic jewelry from the body parts to be scanned, changing into a hospital gown if necessary (ACRIN SOP).

• Subjects undergoing a CT scan should be asked about the presence of implanted electronic devices (e.g. pacemakers, neural stimulators, cochlear implants) (ACRIN SOP).

| Preparation Step | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

## 5. Imaging-related Substance Preparation and Administration

Considerations and specifications related to imaging-related substance preparation and administration are drawn from the protocol in the following sections, augmented with DICOM-field capture for recording purposes.

IV and oral iodinated contrast is not discussed as part of this document as its utility is related to the diagnostic CT examination.

Per the SNM harmonization summit, the FDG must meet USP specifications or meet other current specifications as defined by the FDA or other appropriate regulatory agency.

5.1. Substance Description and Purpose

FDG is a glucose analogue. Its use in oncology is based on the fact that most types of tumours need more glucose than most other types of normal tissue (Neth).

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| Subtstance description | Acceptable Record in DICOM Target Ideal |
| Substance purpose | Acceptable Record in DICOM Target Ideal |

5.2. Dose Calculation and/or Schedule

Consolidated Statement - The 18F - FDG dose is usually around 5mCi in Europe and between 10mCi (=370 MBq) (Hallet) and 20 mCi (=740 MBq) (ACRIN) in the United States. Further FDG dose refinement and/or dose reduction can be achieved by taking into account: (1) patient weight, for example by applying a dose of 5 – 8 MBq/kg (NL, EANM, Hallet, NCI, ACRIN); (2) 2D versus 3D scanning mode (NL, EANM, Hallet, NCI); (3) acquisition time per bed position (NL, EANM, HALLET) and; (4) percentage bed overlap of subsequent bed positions (NL, EANM). The exact dose and the time at which dose is calibrated should be recorded (all, SNM GHS). Residual dose remaining in the tubing, syringe or automated administration system or any dose spilled during injection should be recorded (all).

• In the case of using a syringe for administration, residual activity within the syringe and administration system must be measured and the residual dose should be documented (EU, NL).

• In the case of using an automated system and procedures must assure a net administered FDG activity within 3% accuracy (this must be ensured by manufacturer and verified by the user); i.e., the actual administered activity may not deviate more than 3% from that indicated by the reading of that device or used dose calibrator following instructions given by the manufacturer of the automated administration system (EU).

• Residual activity as determined by the above methods should be used to correct the administered dose for any quantitative results reported.

Any upper dose limits recommended by the manufacturer to take dead time/count rate limitations into account should be considered (EANM, ACRIN). Moreover, (upper) dose limits may apply because of national or local legislation (EANM). In case upper dose limits apply, consistent image quality across sites should be accomplished by increasing scanning time (time per bed, NL, EANM, Hallet). For pediatric studies, other (existing) guidelines may apply, such as the EANM pediatric dose card (EANM).

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| Dose calculation | Acceptable Record in DICOM Target Ideal |
| Dose schedule | Acceptable Record in DICOM Target Ideal |

5.3. Timing, Subject Activity Level, and Factors Relevant to Initiation of Image Data Acquisition

Consolidated Statement:  FDG uptake into both tumors and other body tissues is a dynamic process that peaks and plateaus at various time points dependent upon multiple variables. Therefore, it is extremely important that (1) the time interval between FDG administration and the start of emission scan acquisition is consistent and (2) when repeating a scan on the same subject, it is essential to use the same interval between scans performed across different time points.

The suggested consensus time (from all references) between FDG administration and scan acquisition is 60 minutes based on historical use of this test; assuming this is the ‘ideal’, the target window, ofan acceptable windowility is often cited asis +/- 5 minutes (55-65 minutes). Two references (NCI and ACRIN) allow the acceptable window to be +/- 10 minutes (50-70 minutes), which is considered the absolute minimum of acceptability.

However, on the basis of the SNM harmonization summit while the “target” tracer uptake time is 60 minutes, the “acceptable” window is from 55 to 75 minutes so as to ensure that imaging does not begin prematurely so as to allow adequate tumor uptake of FDG and to account for the practicality of work flow which often does not accommodate imaging at exactly 60 minutes after FDG injection. The exact time of injection must be recorded; the time of injection initiation should be used as the time to be recorded per the SNM GHS. Ideally, the injection and flush should be completed within one minute with the rate of injection appropriate to the quality of the vein accessed for FDG administration so as to avoid compromising the integrity of the injection vein.

Insert paragraph suggesting that more recent evidence might justify a target interval of greater than 60 minutes for a particular trial. If a target time greater than 60 minutes is chosen for a specific trial, the imaging protocol should justify the specific time chosen as well as the acceptable window about this target time. Furthermore, as routine clinical practice might not allow the use of pre-recruitment scan for the study, the protocol should include a plan for repeating the baseline scan if necessary to allow appropriate inter-time-point comparisons.

Multiple references (EU, ACRIN, and NCI) imply that the consistency of timing (relative to baseline) for a follow-up time point should be prioritized for a case where the baseline scanning ideal/target was not achieved; rather than defaulting to a 55 – 65 75 minute (SNM harmonization summit) expectation at follow-up if baseline was performed outside this targeted window. This issue should be specifically addressed for each individual protocol.

When repeating a scan on the same subject, especially in the context of therapy response assessment, it is essential to apply the same time interval with target window of +/- 5 10 minutes (with an acceptable window of +/- 15 minutes) provided that the scan must not begin prior to 55 minutes after the injection of FDG (SNM GHS). If a limited or targeted scan is obtained at follow-up after a whole body scan was performed at baseline, one should consider adjusting the timing of the follow up scan to be congruent with the timing for the same anatomic region as achieved during the baseline study.

If, for scientific reasons, an alternate time (between dose administration and scan acquisition) is targeted for a specific protocol, then the rationale for this deviation should be stated; inter-time point consistency must still be followed (NCI, EU, Neth).

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| Subject activity level | Acceptable Record in DICOM Target Ideal |

5.4. Administration Route

Consolidated Statement – FDG should be administered intravenously through a large bore (≥21 gauge) indwelling catheter placed anatomically remote (e.g., contralateral extremity to site of disease if at all possible) to any site(s) of suspected pathology, preferably in an antecubital vein (SNM GHS). Intravenous ports should not be used, unless no other venous access is available. If a port is used addition flush volume should be used (SNM GHS). As reproducible and correct administration of FDG is required for quantitation purposes extravasation or paravenous administration should be avoided (EU, Neth, NCI, ACRIN). If an infiltration is suspected, the fact should be recorded and if the study is quantitative, i.e. SUVs will be measured, then the infiltration site should be imaged and the approximate amount of infiltration should be calculated If the infiltration is greater than 5% of the administered dose and the quantitative result from the FDG-PET/CT study is a primary or secondary endpoint, the data point might be censored from review or the subject might not be included in study (SNM GHS). The injection site should be documented on the appropriate case report form (ACRIN).

Presuming that the IV access site is properly functioning, the same route of administration may be used for iodinated contrast as is used for FDG.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| Administration route | Acceptable Record in DICOM Target Ideal |

5.5. Rate, Delay and Related Parameters / Apparatus

Consolidated Statement – Either manual or automated injection systems may be used to administer the FDG.

• In the case of manual administration, a three-way valve system should be attached to the previously placed intravenous cannula (See Section 5.4) so as to allow at least a 10 cc normal (0.9% NaCl) saline flush following FDG injection. Residual activity within the syringe, and as much of the administration system as is available (inclusive of the needle cap) must be measured and the residual dose should be documented; and if the residual dose is >1%, the remaining activity should be measured in a calibrated dose calibrator (See Section 5.2) (EU, NL, ACRIN)

• In the case of an automated administration system, the manufacturer’s instructions should be followed. However, the automated system and administration procedures must be ensured by the manufacturer and verified by the user to perform within the characteristics specified in Section 5.2) (EU)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| Rate | Acceptable Record in DICOM Target Ideal |
| Delay | Acceptable Record in DICOM Target Ideal |

5.6. Required Visualization / Monitoring, if any

From the Protocol:

Not applicable.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

5.7. Quality Control

See 12.2

## 6. Individual Subject Imaging-related Quality Control

See 12.3

## 7. Imaging Procedure

This section describes needed specifications and requirements related to the actual imaging procedure. It is structured in such a way as to extract relevant content from the Protocol and then extending that with extensions necessary to explicitly meet the claims as set in the Profile.

**When the imaging procedure involves acquisition of multiple series, Section 7 may be repeated (7a, 7b) as necessary to describe each acquisition.

7.1. Required Characteristics of Resulting Data

**Add any needed introduction or replace the following content from the Protocol.

**Describe the characteristics of the imaging data resulting from acquisition and reconstruction that are relevant to its use in the clinical trial. This description is generally independent of the vendor, platform, and version of the imaging equipment.

7.1.1. Data Content

**Add any needed introduction or replace the following content from the Protocol.

Consolidated Statement   For most Oncology indications, anatomic coverage should include from the skull base (external auditory meatus (EU, ACRIN), skull base (Neth), angle of jaw (NCI)) to the proximal to mid-thigh, considered a ‘whole body’ scan. Per the SNM GHS the acceptable / target / ideal field of coverage is from the external auditory meatus to mid-thigh. However, other ranges could be used as appropriate for specific clinical trials (with justification????). Threeo references (EU and, Neth, SNM GHS) specify that scanning direction should be caudiocranial to minimize effect from increasing bladder activity during the scan. Scanning direction should be protocol specified. It is critical that for a given subject, scanning direction on baseline scans be duplicated at follow-up time points (NCI, SNM GHS).

Two references (EU, NCI) note that extended ‘whole body’ anatomic coverage may be performed for tumors that show higher probability of metastasis or direct extension to other levels (e.g. head, skull, brain or extremities).

One reference (EU) recognizes an option of performing a two-step protocol for examination of head and neck tumors (head / neck portion and apex of lung through mid thigh portion).

One reference (EU) suggests that limited-area tumor imaging can be considered for follow-up examinations, if the disease is restricted to a defined region. If this strategy is used, attention to scan timing is critical to provide comparison with earlier timepoint(s) – reference section 5.3.

Various scanning strategies may be employed in a given specific trial. Potential acceptable scanning strategies include:

For FDG-PET/CT in which the CT is used for attenuation correction and localization only:

1. Scout, followed by

2. CT for anatomic localization and attenuation correction (ideally without IV contrast and either 1) without bowel contrast or 2) with negative or dilute positive bowel contrast) with subject breath-hold instructions per section 7.2.2, followed by

3. Emission scan acquisition.

For FDG-PET/CT in which a diagnostic CT is performed in conjunction with FDG-PET/CT there are two main options, (1) to perform a single CT study prior to or after the emission scan for all purposes (i.e., anatomic localization, attenuation correction, and diagnostic information) or (2) to perform an additional diagnostic CT prior to or after the emission scan.

Note there is currently no consensus on all of the details of these options among various clinical trials and authoritative sources. However, note is made of the following consideration. The diagnostic CT is usually performed in maximal inspiration breath-hold which could result in image artifacts due to mis-registration of the lung-liver interface between emission and CT images when the diagnostic CT is being used for attenuation correction (i.e., there is only one CT scan performed for both diagnosis and attenuation correction). Therefore two sources state that the diagnostic CT cannot be used for attenuation correction purposes (EU, Neth).

Three references indicate that the subject should be fully included within the CT and PET FOV in order to avoid attenuation and scatter correction artifacts resulting in non-quantifiable data (NL, EANM, ACRIN)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.1.2. Data Structure

**Add any needed introduction or replace the following content from the Protocol.

**Describe how the data should be organized/sampled. (e.g. Spatial Resolution, Collimation Width, Slice Interval, Temporal Resolution)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable The matrix size, slice thickness, and reconstruction zoom will yield a target reconstructed voxel size of 3 – 4 mm in all three dimensions (i.e., not achieved through post-processing), although not necessarily isotropic. – for QC see section 12.1.1. QIBA and/or SNM Image Reconstruction Committee????? Target Same as Acceptable Ideal Reconstructed voxel size (i.e., not achieved through post-processing) should be as small as possible without introducing artifacts and also so as to be consistent across all trial sites; with current technology 2 – 3 mm in all three dimensions is achievable. |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.1.3. Data Quality

**Add any needed introduction or replace the following content from the Protocol.

Image quality (as defined by SUV calibration, SUV Recovery Coefficient, and SNR) should be such that when applying the same acquisition and reconstruction protocol as used in subject scanning to the protocol specified phantom(s) the output should meet the QC standards as stated in Section 12.1.1.

Treatment response assessment and classification (based on criteria) require several quantitative and qualitative assessments:

First of all, lesion eligibility will be based on tumor tracer uptake above a certain threshold (section 10 ?).

Secondly, response will be determined by measuring changes in SUV over time (either in the same lesion(s), EORTC, or in the ‘hottest per scan’lesion, PERCIST). Apart from a percentage change above or below a threshold to determine progression, stable disease or response, a minimal absolute change is suggested as well (Percist)

Finally, an important factor in the assessment of disease progression is not only an increase in SUV above a certain specified threshold, but also the appearance of new lesions.

For the first two aspects (lesion eligibility and measuring % changes) harmonization of quantitative image quality, e.g. be means of harmonizing recovery coefficients measured in specific dedicated phantoms, will result in more uniform lesion selection and response assessments across institutes. Consequently, harmonizing quantitative performance of PET/CT systems should be aimed at in a multicenter trial

For the assessment of progression when new lesion appears it is important to set a minimal threshold for image quality with respect to lesion detectability. As such scanners need to have a minimal image quality performance/lesion detectability/SNR in order to be suitable to be used in trials.

Insert text per discussion 01052011 – Ronald to author first draft regarding the balance of detectability and uniformity and quantitative harmonization, SNR, confidence of new lesion detection, smoothing, etc. (attribution to SNM GHS)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.2. Imaging Data Acquisition

Consistency issues:

- user should be able to use the same SW version and settings for all longitudinal studies for one patient

- user should be able to use the same SW version and settings for all studies within a trial

- Alternative: vendor has implemented an acquisition/ reconstruction protocol that ensures consistent qualitative and quantitative image characteristics independent of SW version. For example this could be verified by looking at lesion detectability (SNM phantom), uniformity (cylinder) and SUV recoveries (modified NEMA). Also we could specify a minimal COV of pixel values within a VOI measured under specific experimental conditions in a phantom.

For serial scan of the same subject, every attempt should be made to use the same scanner, and the same scanner model throughout the trial. (ACRIN).   However, in some cases a different scanner that has been previously qualified and is the same platform as the scanner used at baseline can be used for a subject’s follow-up scan in the instance of equipment malfunction (ACRIN 6678).

Per the SNM GHS, the ideal level of performance is that all serial scans on a subject should be performed on the same scanner with the same software version; target performance is that all serial scans on a subject should be performed on equivalent scanners (i.e., the same model) but also with the same software version. (SNM GHS)

Additionally, all scan acquisitions for a given subject should include identical transmission and emission scan technique and emission scan duration per bed position (ACRIN). There is no consensus provided on emission scan time range. The number of bed positions and the acquisition time per bed position will be scanner specific. Typical parameters are 6 bed positions and an acquisition of 2 – 5 min per bed position (ACRIN 6678). Per the SNM harmonization summit the minimum acceptable time per bed position should be between 2 and 4 minutes for a 3D acquisition with 2D acquisitions typically requiring at least 1.5 - 2x longer depending on the administered FDG dose; although the absolute impact on image quality by scan time per bed position is currently undefined it is dependant on several pertinent factors including, but not limited to, administered dose, body weight and habitus, bed overlap, and specific model / version of the imaging platform used. In general, increased scan time per bed position will improve the SNR and thus it may be important to increase scan time when quantitative metrics are used towards a primary endpoint. A more detailed set of preliminary recommendations for time per bed position relative to pertinent factors is provided in Appendix XYZ (or is this Table X in Section 13.2??).

One reference (NCI) notes that whole body acquisitions can be in either 2- or 3- dimensional mode with attenuation correction, but a consistent method should be chosen for all serial scanning of an individual subject throughout the trial.

Two references (NL, EU) describe a relationship between applied FDG dose, acquisition time per bed position, percentage bed overlap and scanning mode (2D, 3D) in order to harmonize image quality (and avoid bias in quantification). Using this relationship these parameters are directly linked, e.g. a higher FDG dose can be offset by shorter acquisition times per bed position etc. Please see the equations and Table X in Section 13.2 for examples relating dose, bed positions, scan time per bed position, scanning mode, and bed overlap, etc. (NL, EU).

Acceptable: While there may be variance based on type of scanner, scanning algorithm, model, and software version, the following guidelines are meant to assist each site in achieving the desired data quality as specified in Sections 5.2, 7.1.3, xxxx, and yyyy.

For a dose of 5 MBq/kg or higher (370 MBq or more for a 75 kg patient) the minimal time per bed position using the manufacturers’ recommended bed overlap specifications. The time per bed position should be at least 2 mins for 3D systems showing ≥50% bed overlap and at least 4 min for 3D systems showing less than< 50%. Time per bed position may be modified inversely proportional to alteration in injected dose per body weight within the limits of the scanner performance as determined by the manufacturer or an appropriately qualified independent standard-setting organization or peer-reviewed publication.

For 2D systems these times per bed should be at least 1.5 times longer for the same injected dose based on body weight. Time per bed position may be modified inversely proportional to alteration in injected dose per body weight within the limits of the scanner performance as determined by the manufacturer or an appropriately qualified independent standard-setting organization or peer-reviewed publication.

In general, increased scan time per bed position will improve the SNR and thus it may be important to increase scan time when quantitative metrics are used towards a primary endpoint.

Target (just an idea we are working on at EANM See the additionally attached document); Use an anthropomorphic phantom (SNM, EANM) with a uniform area to assess ‘noise’ by means of the coefficient of variation of the pixel values within a specified region of interest (ROI) positioned in the uniform area. Phantom should be filled such that activity concentration in the uniform area is similar (within 150%) to the expected average normal tissue concentration seen in an average weight (70-80 kg) subject in combination with the intended FDG dosage. Phantom should than be scanned with various times per bed and reconstructed using the acceptable reconstruction methods and settings (e.g. minimal and/or harmonized resolution criteria etc etc) . Next COV will be derived in each image following the NEMA NU-2 protocol. The COV within the above indicated ROI should be plotted as function of time per bed. The minimal allowable time per bed position is then given by the time per bed that provides a COV of 105% or better. Afterwards higher FDG doses may be offset by longer scan durations

Ideal: I guess this is yet unknown.

7.2.1. Subject Positioning

**Add any needed introduction or replace the following content from the Protocol.

Consolidated Statement –   During PET-CT, subjects should be positioned in the center of the field of view (FOV), preferably with the subjects’ arms to be positioned over head (to minimize beam hardening and FOV truncation artifacts) (ACRIN, SNM GHS). Alternatively, the arms can be positioned along the side for head and neck imaging (for two-step procedure – see section 7.1.1) (EU, SNM GHS). One reference notes that subject may be unable to maintain arms above head for the examination, in which case protocol specific handling needs to be defined. Arm positioning in a particular subject should be consistent as possible across all time points. (ACRIN) One reference (EU) notes that if PET-CT data are used to radiation planning, the examination should be carried out in the radiation position using the same dedicated radioopaque positioning devices as used in the radiotherapy department. Support devices, under the back and/or the legs, may be used to enable the subject to comfortably maintain his/her position throughout the exam (ACRIN).

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.2.2. Instructions to Subject During Acquisition

**Add any needed introduction or replace the following content from the Protocol.

Consolidated Statement - The diagnostic CT is usually performed in maximal inspiration breath-hold which could result in image artifacts due to mis-registration of the lung-liver interface between emission and CT images when the diagnostic CT is being used for attenuation correction (i.e., there is only one CT scan performed for both diagnosis and attenuation correction). Therefore, two references (EU, Neth) indicate that during CT-AC, shallow breathing is recommended for multi-slice CT scanners; one reference (EU) notes that for single or two-slice CT scanners, breath hold acquisition should be acquired. Per the SNM GHS, the CT acquisition for attenuation correction should be done with shallow breathing without regard to the CT technology used (acceptable / target / ideal).

One reference (ACRIN) notes that the subject should be monitored and coached (1) in the breathing protocol, if one is used and (2) to remain motionless throughout the scan

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.2.3. Timing/Triggers

**Add any needed introduction or replace the following content from the Protocol.

**(e.g., relative to administration of imaging agents; inter-time point standardization)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.2.4. Model-specific Parameters

Appendix G.1 lists acquisition parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

7.2.5. Archival Requirements for Primary Source Imaging Data

See 11.3

7.2.6. Quality Control

See 12.3

7.3. Imaging Data Reconstruction

From wishlist (see also section …):

Easy and fast procedure to adjust calibration of scanner, eg when dose calibrator is recalibrated then it would be very convenient to simply adjust calibration of scanner with a simple correction factor rather than an extensive procedure. Preferably a correction factor that is automatically generated and applied (upon user request) based on a calibration QC experiment.

Flexibility issues (may also be put in section ....)

- vendors should allow users to modify a number of basic acquisition and reconstruction parameters in order to optimize or harmonize image quality for a specific task. Parameters that need to be available for the user to be specified (and stored in a study specific acquisition/reconstruction protocol) are:

o scan duration per bed position (and variable within one study)

o number of bedpositions

o image matrix size (128-512), zoom (1-2) OR voxel size (2,3,4,5 mm). Should be reconstructed voxel size and not obtained after rebinning of reconstructed image)

o number of iterations and subsets (and relaxation factor)

o ToF kernel size

o Image filtering (type and FWHM)

o ??

o ??

Consolidated Statement (n.b., concepts to capture STILL A WIP, NOT COMPLETED)

- PET emission data must be corrected for geometrical response and detector efficiency (normalization), system dead time, random coincidences, scatter and attenuation (EU, Neth, ACRIN).

- Data acquired in the 3D mode can be reconstructed directly using a 3D reconstruction algorithm or rebinned in 2D data and subsequently be reconstructed with a 2D reconstruction algorithm (EU).

- Iterative reconstruction algorithms are current standard for PET (rather than filtered back projection) (EU).

- (ACRIN, Must) Perform reconstructions with and without attenuation correction (EU, ACRIN).

-Scanners must be properly normalized and calibrated to ensure uniformity and accuracy of SUV measurements within the limits of the spatial resolution (ACRIN)

- Standardization of reconstruction settings is necessary to obtain comparable resolution and SUV recoveries across the same subject and inter-subject across sites < model specific issues to allow multicenter convergence - reference in template to 7.3.1 and Appendix??> (Netherlands, EU)

7.3.1. Input Data to Be Used

**Add any needed introduction or replace the following content from the Protocol.

**Place requirements on the input data to be used.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.3.2. Methods to Be Used

Claim is that data can be reconstructed including all corrections needed for quantification as well as without scatter and attenuation correction. Iterative reconstruction method should be applied.

PET/CT systems are equipped with a harmonized reconstruction protocol (NB not the algorithm or settings are harmonized, but the resulting specified quantitative image characteristics for specified phantoms/experiments) – SNM Image Reconstruction Workgroup.

**Place requirements on the methods to be used.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.3.3. Required Characteristics of Resulting Data

(section 11.4) Store PET image data always in units of Bq/ml in DICOM files, decay reference time should equal (PET) scan acquisition time and integral model should be used for decay correction.

Vendors should clearly specify which time (series, study, acquisition, dose essay time, injection time) is used as reference for decay correction when data are given in units Bq/ml.

All needed information for fully corrected nett injected activity (e.g. residual activity, injection time, calibration time) is required. Scanner performs all decay corrections (not the operator) (see also section 9.2)

The matrix, slice thickness, and reconstruction zoom will yield a target voxel size of 3 – 4 mm in all three dimensions, although not necessarily isotropic and not achieved by rebinning etc of the reconstructed images

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.3.4. Platform-specific Instructions

Appendix G.2 lists reconstruction parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.2.

7.3.5. Archival Requirements for Reconstructed Imaging Data

See 11.4

7.3.6. Quality Control

See 12.4

## 8. Image Post-processing

**Add any needed introduction or replace the following content from the Protocol.

Consolidated Statement: Standard whole-body FDG-PET oncology studies typically include all necessary data corrections and processing within the reconstruction process and do not require additional post-processing other than data de-identification (see section 11.2). More advanced studies such as those including dynamic imaging may require additional post-processing as specified in the individual protocol.

8.1. Input Data to Be Used

**Add any needed introduction or replace the following content from the Protocol.

**Describe required input data and any necessary validation or adjustments which should be performed on it. (e.g. particular image series or views; raw, processed, both)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

8.2. Methods to Be Used

**Add any needed introduction or replace the following content from the Protocol.

**Describe how the analysis should be performed.  (e.g. algorithms to be used; where measurements should be taken; definition of key anatomical points or pathology boundaries; related annotations)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

8.3. Required Characteristics of Resulting Data

**Add any needed introduction or replace the following content from the Protocol.

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

8.4. Platform-specific Instructions

Appendix G.3 lists post-processing parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 8.

8.5. Archival Requirements

See 11.5

8.6. Quality Control

See 12.5

## 9. Image Analysis

**Add any needed introduction or replace the following content from the Protocol.

Consolidated Statement – n.b., WIP certainly not complete, high level concepts only; citations may be incomplete or in error; also need to be parsed among the various subsections within Section 9 – BUYER BEWARE.

Issues to address in Analysis:  What data are to be used and for each type of data for what purposes is it appropriate to use which specific data?  SUV measure to be utilized needs to be specified for each protocol and needs to be used consistently across all subjects and across all sequential lesion measurements. Categories for consideration and definition include:

1. Options for Region-of-Interest Statistics -

- Values to report: SUVmax, SUVmean, SUVpeak, TLG?

All references indicate that SUVmax (maximum voxel value or hottest voxel) is required for each lesion which is reported as specified in the study protocol and/or considered clinically relevant.

Multiple references (NCI, ACRIN, EU) also indicate that SUVmean of the ROI/VOI obtained be reported. The SUVpeak equals the SUVmean only when the VOI is a sphere with a specified diameter which is also indicated as a reportable statistic (EU, ACRIN).

CONSENSUS: (rw3)

- SUV max single pixel (normalized for LBM) - ?acceptable

- SUV peak 3D for 1 cm volume (1.2 cm sphere) - ? ideal

- SUV peak 2D (1.2 cm diameter) - ?target - SUV values based on bw, lbm, bsa

CONSENSUS: (rw2) - SUV normalized to lean body mass (SUL?) - ? ideal <equation to be used needs further review>

- SUV based on body weight should be reported - ?target

- Other values of ROI/VOI to be reported (e.g. number of pixels, counts/pixel, size)

2. Definition of the region for quantitation -

- ROI (or VOI) tool to be utilized to define lesion constraint condition and strategy to define edge detection; region growing versus based on threshold condition (% of max / absolute threshold / gradient sampling/ iterative / fuzzy clustering)

CONSENSUS: (rw5) - Tumor ROI's reflecting the metabolic volume of the tumor is desirable. Volumes based on a 50% and (?) 70% threshold of the peak tumor SUV should be produced. Alternative is to use the PERCIST threshold at 2SD above mean liver activity as a floor value for the ROI's. (?) These can reflect tumor volume and tumor volume x mean SUV lean - viewed as exploratory.

CONSENSUS: (rw4) - Suggested that use of partial pixel values to secure a 1.2 cm diameter (1cc volume) ROI was appropriate and desirable, since standard pixel sizes would not allow selection of a 1cc volume precisely in most cases.

- Automatic vs. Semi-automatic vs. Manual and circumstance qualifiers

- Strategy for analysis of necrotic core lesion

3. Normalization (if required) (e.g. based on lesion size, glycemic status)

- Strategies around the issue of partial volume effects

- Strategies for glucose correction

Analysis to be based on original reconstructed PET image (no additional rebinning, resampling, smoothing by user is allowed Issues relevant to dynamic analysis

9.1. Input Data to Be Used

**Add any needed introduction or replace the following content from the Protocol.

**Describe required input data and any necessary validation or adjustments which should be performed on it.  (e.g. particular image series or views; raw, processed, both)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

Tumor ROI's reflecting the metabolic volume of the tumor is desirable. Volumes based on a 50% and (?) 70% threshold of the peak tumor SUV should be produced. Alternative is to use the PERCIST threshold at 2SD above mean liver activity as a floor value for the ROI's. (?) These can reflect tumor volume and tumor volume x mean SUV lean - viewed as exploratory.

Measurement system (viewing and analysis software) shall provide the actual maximum voxel value in Bq/ml and the actual maximum SUV (SUVmax), which is the value or SUV of a single voxel from the actual original image (digital reference object, DRO) that has the highest value within a volume of interest drawn around the object of interest (maximum voxel value test object in DRO).

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Zero error Target Ideal |
| ROI extraction of activity concentrationjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjjj | Acceptable Zero error Target Ideal |

9.2. Methods to Be Used

**Add any needed introduction or replace the following content from the Protocol.

**Describe how the analysis should be performed.  (e.g. algorithms to be used; where measurements should be taken; definition of key anatomical points or pathology boundaries; scoring scales and criteria, related annotations)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

ROI computations

Report actual max. voxel value

Capablity to lable, save, recall (in a std format, e,g, Advanced Technology Consortium (ATC) for Clinical Trials Quality Assurance)

Report VOI activity as Bq/ml

Report SUVs normalized to Weight, LBM, BSA

Report volume in cc (or ml) and voxels

Describe voxel inclusion policy

Use weighting for partial voxels

Fully included voxels use weight of 1.0

Output statistics inlcude mean, max, std. dev. to 3 sig figs

Describe rounding policy for outputs

VOI dimensions (sphere) should be specifiable to ±1 mm

Automatic centering for peak VOI calculation

Three threshold methods: fixed, % of max, gradient

Produce correct SUVs from QIBA DRO

All needed information for fully corrected nett injected activity (e.g. residual activity, injection time, calibration time) is required. Scanner performs all decay corrections (not the operator)

| Parameter | Compliance Levels |
| --- | --- |
| Activity concetnration | Acceptable Target Ideal |
| glucose correction | Acceptable Target Ideal |

9.3. Required Characteristics of Resulting Data

**Add any needed introduction or replace the following content from the Protocol.

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| SUVmax | Acceptable single pixel (normalized for LBM) Target Ideal |
| SUVmean | Acceptable Target Ideal |
| SUVpeak 3D | Acceptable Target Ideal for 1 cm volume (1.2 cm sphere) |
| SUVpeak 2D | Acceptable Target Ideal |
| SUL | Acceptable Target Ideal SUV normalized to lean body mass |
| TLG | Acceptable Target Ideal |

9.4. Platform-specific Instructions

Appendix G.4 lists image analysis parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 9.

9.5. Archival Requirements

See 11.6

9.6. Quality Control

See 12.6

## 10. Image Interpretation

**Add any needed introduction or replace the following content from the Protocol.

Consolidated Statement: n.b., WIP certainly not complete, high level concepts only; citations may be incomplete or in error; also need to be parsed among the various subsections within Section 9 – BUYER BEWARE.

Issues to address -

1. Quality assessment of exam and actions based on assessment (of what specific parameters??)

Both uncorrected and attenuation corrected images need to be assessed in order to identify any artifacts caused by contrast agents, metal implants and/or subject motion (EU).

CONSENSUS: (rw8) - Image quality should be noted. For example, movement or misregistration can lead to poor quality quantitative data and invalid numbers. Some images may be too poor in quality to quantify.

CONSENSUS: (rw9) - Statistical quality of images is important to report, but not a full substitute for quality. Liver noise per PERCIST was considered a reasonable start.

2. Presence or absence of abnormal FDG accumulation in combination with size and intensity of focal accumulation is evaluated

3. Use of qualitative vs. quantitative measurements

4. Target, Non-target, New lesion binning and metadata requirements

5. Normal appearing tissue reporting requirement (e.g. liver, muscle, blood pool background)

CONSENSUS: (rw7) - Internal normalization should be reported. Considerations included use of normal liver and use of blood pool activity. For liver, mean SUL was suggested using 3 cm diameter VOI (PERCIST standard). These outputs could be used as QC parameter as well as available for normalization (if desired).

6. Baseline target lesion requirements (number, size and metabolic activity thresholds/dependencies)

CONSENSUS: (rw6) - Number of target lesions – up to 5; not more than 2/organ

CONSENSUS: Minimum threshold requirements:

- Baseline size (rw6): GTD > 2 cm

- Baseline metabolic (rw1): Qualification as quantifiable target lesion requires Tumor to Background ratio (TBR) of >2:1 - using liver mean ??or other background tissue?? Recognition of recommendation of PERCIST of 1.5 x liver mean + 2xSD of liver noise. Lesion has to be "hot" enough at baseline so that a reduction in tumor metabolism can be detected at follow up.

7. Response Classification strategies (e.g. EORTC, lymphoma, PERCIST)

- The SUV change metric

- use of % SUV change only or also absolute SUV change?

- Whether/how to assess non-targets

- Whether/how to identify new lesions and incorporate into scheme

- How to deal with mixed response

CONSENSUS: (rw10)Response Interpretation briefly discussed PERCIST criteria a reasonable beginning. Issue of confirmation was raised; agreement that PD with new lesion consistent with tumor also if CT worsened and PET progressed (i.e. no need to confirm). However, if PET worsened without change in CT, conservative approach would be to confirm (but not firm on this point).

Response assessment bins and definitions (this was discussed, but not in consensus notes to date):

- PD = new lesions or incr. of > 25-30%?

- SD = no new; between PR and PD quantitatively

- PR = > 25-30 % reduction

- CR = no visible tumor remaining

<issue of reporting confidence level of interpretations based on QA and/or examination intertimepoint variables not addressed> (esp note)

10.1. Input Data to Be Used

**Add any needed introduction or replace the following content from the Protocol.

**Describe required input data and any necessary validation or adjustments which should be performed on it. May also specify data which should not be used until after the clinical trial interpretation is recorded. (e.g. particular image series or views; before and after processing versions of images to evaluate/validate the effects of processing; analysis results)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| Image quality | Acceptable Record in DICOM Target Ideal |

10.2. Methods to Be Used

**Add any needed introduction or replace the following content from the Protocol.

**Describe how the interpretation should be performed.  (e.g. definition of key anatomical points or pathology boundaries; scoring scales and criteria such as BIRADS, interpretation schema such as RECIST, related annotations)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

10.3. Required Characteristics of Resulting Data

**Add any needed introduction or replace the following content from the Protocol.

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

10.4. Platform-specific Instructions

Appendix G.5 lists image interpretation parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 9.

10.5. Archival Requirements

See 11.7

10.6. Quality Control

See 12.7

## 11. Archival and Distribution of Data

**Add any needed introduction or replace the following content from the Protocol.

The following sections describe the archival and distribution of data.

11.1. Reserved

11.2. Reserved

11.3. Primary Source Imaging Data

**Add any needed introduction or replace the following content from the Protocol.

Not mentioned in either of the sources.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

11.4. Reconstructed Imaging Data

CT: Daily water equivalent phantom values tracked in DICOM header

PET: Daily/weekly/monthly scanner QA values included in DICOM header

PET: Daily (or frequent) uniform cylinder analysis, with link to results in patient DICOM header

Two sources (EU, ACRIN) mention use and storage of DICOM formatted data. One source (EU) indicates that data should be stored in DICOM format Part 10: Media Storage and File Format for Media Interchange. DICOM format should meet the Conformance Statement written by manufacturer of the PET/CT system.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol:

Software Version Tracking: Use enhanced DICOM attributes to follow version number of software for: 1 Acqusition, 2 Reconstruction, 3 post-processing, 4 Display/VOI analysis, 5 Dynamic Analysis. Build list (on console) of dates of all s/w versions

QA/QC Tracking: Dose calibrator is calibrated for F-18 using NIST-traceable source with information included in patient DICOM header   (section 11.4) Store PET image data always in units of Bq/ml in DICOM files, decay reference time should equal (PET) scan acquisition time and integral model should be used for decay correction.

QA/QC Tracking (or section 11.4)

CT: Daily water equivalent phantom values tracked in DICOM header

PET: Daily/weekly/monthly scanner QA values included in DICOM header

PET: Daily (or frequent) uniform cylinder analysis, with link to results in patient DICOM header

List SUV computation (w/ DICOM tags)

Blood glucose (from CLIA compliant device) at time of injection is recorded with DICOM patient information (would require entry in GUI of pet/ct console before or at start of acquisition)

Patient meta-data recorded, e.g. using enhanced DICOM attributes

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

11.5. Post-Processed Data

**Add any needed introduction or replace the following content from the Protocol.

Not mentioned in either of the sources.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

11.6. Analysis Results

**Add any needed introduction or replace the following content from the Protocol.

Not mentioned in either of the sources.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| Quantitative reporting | Acceptable Include reference tissue value (e.g. 3 cm diameter VOI in liver) Populate reports from DICOM header information Mechanism for flagging artifacts (motion, extravasation, etc.) NOTE: should contact RSNA group working on structured reporting (David Clunie or Chris Carr) Target Ideal |

11.7. Interpretation Results

**Add any needed introduction or replace the following content from the Protocol.

Not mentioned in either of the sources.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

## 12. Quality Control

**Add any needed introduction or replace the following content from the Protocol.

The following sections describe all aspects of Quality Control are identfied and described.

12.1. QC Associated with the Site

The following sectins consider various aspects of quality control.

12.1.1. Quality Control Procedures

Add QC check with a test object

Draft Consolidated Statement: - PET and PET-CT systems (and dose calibrator, scales, glucose monitoring equipment, etc.) should routinely be assessed for quantitative integrity and stability by being tested using various imaging protocols on a standard phantom (NL, EANM, Hallet, ACRIN). n.b., no consensus from extractions but for consideration: Currently there is no single “standard phantom” agreed upon by all references or extracted studies. Phantoms might need to be specified for certain types of cancers or anatomic locations and therefore might vary from trial to trial based on diagnosis, treatment, and/or anatomic location. Options that might be considered on a per protocol basis including, but are not limited to 1) each site would use the same phantom for the duration of the trial (but the phantoms might not be exactly the same among all sites), 2) all sites would use the same general type of phantom for the duration of the trial, 3) all sites would use phantoms that are precisely specified for the duration of the trial, 4) all sites would share the exact same phantom for the duration of the trial. For SUV measurements, this assessment should include a comparison against a dose calibrator to ensure accuracy; that is, a comparison of the absolute activity measured, versus the measured injected, should be performed (5% NL, 10% EANM, ACRIN). This comparison is particularly important after software or hardware upgrades (NL, EANM, ACRIN). In case absolute besides relative (longitudinal, response assessment) quantitation is required, it should be considered to include an image quality and/or contrast recovery QC assessment as part of the routine QC procedures and/or scanner validation process, see Appendix E. (NL, EANM).

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Minimally required is that images can be used in a quantitative manner. The latter implies at least correct calibration and/or cross-calibration of the PET-CT systems against the (locally) used dose calibrator within 10%. (NL, EANM, ACRIN, NCI, Hallet) The same scanner with the same acquisition/reconstruction protocol, software and settings should be used for each subject study to ensure consistent image quality characteristics for all scans performed for each individual patient (). As well the QC procedures should utilize the same acquisition/reconstruction protocol, software and settings that are used for the subject scans. In case absolute besides relative (longitudinal, response assessment) quantitation is required, it should be considered to include an image quality and/or contrast recovery QC assessment as part of the routine QC procedures and/or scanner validation process (NL, EANM, SNM-CTN). Target SUV calibration accuracy should be within 5% (NL), voxel size should be 3 – 4 mm in all three dimensions but not necessarily isotropic, SUV recovery should be measured and stored as meta-data for the trial, but there is no minimum performance threshold, uniformity should be XXXX, SNR should be YYYY. These can be verified using various QC methods as described in Appendix E. The same (type and set of) phantoms, metrics and specifications should be used across all sites participating in a specific trial . Vendors implement tools or provide functionality such that users can modify acquisition and image reconstruction settings (iterations, filters, matrix size etc) that change these image characteristics and thereby allow users or CTNs to implement an acq.recon.protocol that collects data which meets the trial specified image characteristics (see Appendix G). The same scanner with the same acquisition/reconstruction protocol and settings should be used for each patient study. Ideal All vendors have implemented one acquisition and reconstruction protocol on each of their machines such that use of this protocol ensures that the specified image characteristics are met (and by implication the SUV Recovery Coefficient is not only measured and saved as meta-data but is standardized by default at some value for any and all PET/CT platforms). This implies that vendors may need to revise this protocol to deal with SW or HW upgrades, i.e. in all cases use of the ‘standard’ protocol should ensure that image characteristic are and remain globally harmonized. Scanner clocks and all clocks used to record dose timing should be synchronized and periodically monitored as part of ongoing QC program (EU, ACRIN, Neth, SNM GHS). |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.1.2. Reserved

12.1.3. Reserved

12.2. QC Associated with Imaging-related Substance Preparation and Administration

**Add any needed introduction or replace the following content from the Protocol.

The following requirements are placed on QC associated with imaging-related substance preparation and administration.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.3. QC Associated with Individual Subject Imaging

**Add any needed introduction or replace the following content from the Protocol.

The following sections describe aspects of QC associated with individual subject imaging.

12.3.1. Phantom Imaging and/or Calibration

**Add any needed introduction or replace the following content from the Protocol.

<Do we want to place reference here to Appendix D?>

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.3.2. Quality Control of the Subject Image and Image Data

**Add any needed introduction or replace the following content from the Protocol.

Consolidated Statement – The integrity of DICOM image headers should be reviewed and confirmed for regulatory compliance (HIPAA), protocol compliance, and consistency with source data such as CRFs. In some cases, internal references such as the liver can be used for quality control to confirm acceptable ranges of SUVs (ACRIN 6678). For details of performance see Appendix D.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.4. QC Associated with Image Reconstruction

**Add any needed introduction or replace the following content from the Protocol.

Consolidated Statement - CT images should be reviewed for potential artifacts such as beam hardening, metal objects, and motion. PET images should be compared to the CT images for proper image registration and potential attenuation correction artifacts. (ACRIN 6678). For details of performance see Appendix D.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.5. QC Associated with Image Processing

**Add any needed introduction or replace the following content from the Protocol.

The following requirements are placed on QC associated with image processing.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.6. QC Associated with Image Analysis

**Add any needed introduction or replace the following content from the Protocol.

The following requirements are placed on QC associated with image analysis.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
| QC | Acceptable QC with Digital reference object ? Target Ideal |

12.7. QC Associated with Interpretation

**Add any needed introduction or replace the following content from the Protocol.

The following requirements are placed on QC associated with interpretation.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

## 13. Imaging-associated Risks and Risk Management

**Add any needed introduction or replace the following content from the Protocol.

The following sections consider various imaging-associated risks and risk management.

13.1. Radiation Dose and Safety Considerations

**Add any needed introduction or replace the following content from the Protocol.

The radiation dose of the PET/CT study results from the radiation exposure from the injection of FDG and the CT study (EANM, ACRIN, HALLET). One source (EANM) indicates that CT scans can be performed as low dose CT to be used for attenuation correction purposes to minimize radiation dose. Two sources (EANM, HALLET) indicate that radiation dose from the CT scans should be estimated specific to the system and imaging protocol used (EANM) or by means of standard estimates (HALLET). These standard estimates can be utilized within the framework of local regulatory requirements for risk analysis (HALLET), which will also depend on patient populations and life expectancy (HALLET) and particular considerations to reduce radiation exposure should be given for paediatric applications (EANM). The protocol should contain language estimating the total radiation dose and its comparative radiation risk. The exact estimates of radiation and risk will be protocol-specific and based on factors such as the number and frequency of studies.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13.2. Imaging Agent Dose and Safety Considerations

**Add any needed introduction or replace the following content from the Protocol.

(ACRIN) Approximately 1 person in 1000 may have an allergic reaction from the iodinated contrast drugs. These reactions are temporary and treatable. Allergic reactions may include: mild itching or hives (small bumps on the skin), and shortness of breath and swelling of the throat or other parts of the body. The subject should be instructed to tell the technologist immediately if s/he experience any of these symptoms so s/he can be treated promptly.

(ACRIN) The placement of intravenous catheters has the associated risk of making the patient temporarily uncomfortable and a small bruise may form. A slight bruise may form where the needle has been in a vessel. There is a slight risk of infection at the site, but sterile technique reduces this risk nearly completely. The patient may also experience claustrophobia from the imaging ring apparatus or discomfort from lying on the scanner table for 60-120 minutes

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13.3. Imaging Hardware-specific Safety Considerations

**Add any needed introduction or replace the following content from the Protocol.

Not mentioned in either source

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

The following set of requirements extends what has been stated in the protocol.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13.4. Reserved

13.5. Reserved

# IV. Compliance

Acqusitisiton System

What we do to test for compliance by acquisitions systems

Ananlysis System

Viewing System

**the compliance method for that actor.

# References

**a reference

# Appendices

## Appendix A: Acknowledgements and Attributions

NEED TO ACKNOWLEDGE ALL TEAMS and INDIVIDUALS

don't forget the source documents from which the extractions were made

don't forget the SNM worldwide summit

**Add any additional or content, or replace content from the Protocol.

## Appendix B: Background Information

**

**Add any additional or content, or replace content from the Protocol.

## Appendix C: Conventions and Definitions

Definitions: Review this document and define relevant terms and notations here.

Acquisition vs. Analysis vs. Interpretation: This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be computational steps that transform the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge.  (The transformation of knowledge into wisdom is beyond the scope of this document.)

Bulls-eye Compliance Levels Acquisition parameter values and some other requirements in this protocol are specified using a “bullseye” approach. Three rings are considered from widest to narrowest with the following semantics:

ACCEPTABLE: failing to meet this specification will result in data that is likely unacceptable for the intended use of this protocol.

TARGET: meeting this specification is considered to be achievable with reasonable effort and equipment and is expected to provide better results than meeting the ACCEPTABLE specification.

IDEAL: meeting this specification may require unusual effort or equipment, but is expected to provide better results than meeting the TARGET.

An ACCEPTABLE value will always be provided for each parameter. When there is no reason to expect better results (e.g. in terms of higher image quality, greater consistency, lower dose, etc.), TARGET and IDEAL values are not provided.

Some protocols may need sites that perform at higher compliance levels do so consistently, so sites may be requested to declare their “level of compliance”. If a site declares they will operate at the TARGET level, they must achieve the TARGET specification whenever it is provided and the ACCEPTABLE specification when a TARGET specification is not provided. Similarly, if they declare IDEAL, they must achieve the IDEAL specification whenever it is provided, the TARGET specification where no IDEAL level is specified, and the ACCEPTABLE level for the rest.

**Add any additional or content, or replace content from the Protocol.

## Appendix D: Documents included in the imaging protocol (e.g., CRFs)

**(Material the site needs to submit)

**Subject preparation

**Imaging agent dose calculation

**Imaging agent

**Image data acquisition

**Inherent image data reconstruction / processing

**Image analysis

**Interpretation

**Site selection and Quality Control

**Phantom Imaging and Calibration

**Add any additional or content, or replace content from the Protocol.

## Appendix E: Associated Documents

**e.g. the Imaging Charter, Site Manual, Standard Operating Procedures, etc.

**Add any additional or content, or replace content from the Protocol.

## Appendix F: TBD

## Appendix G: Model-specific Instructions and Parameters

The presence of specific product models/versions in the following tables should not be taken to imply that those products are fully compliant with the QIBA Profile. Compliance with a profile involves meeting a variety of requirements of which operating by these parameters is just one. To determine if a product (and a specific model/version of that product) is compliant, please refer to the QIBA Conformance Document for that product.   G.1. Image Acquisition Parameters  The following technique tables list acquisition parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

These technique tables may have been prepared by the submitter of this imaging protocol document, the clinical trial organizer, the vendor of the equipment, and/or some other source. (Consequently, a given model/version may appear in more than one table.) The source is listed at the top of each table.

Sites using models listed here are encouraged to consider using these parameters for both simplicity and consistency. Sites using models not listed here may be able to devise their own acquisition parameters that result in data meeting the requirements of Section 7.1 and conform to the considerations in Section 13.

In some cases, parameter sets may be available as an electronic file for direct implementation on the imaging platform.

**Add any additional or content, or replace content from the Protocol.