X.  Imaging Protocol – <<Insert Catchy Protocol Name>>

Instructions to Clinical Trialists who are adapting this imaging protocol for inclusion in their Clinical Trial Protocol are shown in italics.  All italic text should generally be removed as part of preparing the final protocol text.

Executive Summary

Provide a brief (less than 250 words) synopsis to let readers quickly determine if this imaging protocol is relevant to them.  Sketch key details such as the primary utility, imaging study design, specific aims, context, methods, expected results, risks, and deliverables.

This protocol describes imaging, measurements and interpretation for quantitatively evaluating the progression/regression of lung tumors greater than 1cm in Longest Diameter.  It is intended to provide “twice the sensitivity of RECIST”.

<Need more complete summary.>

Context of the Imaging Protocol within the Clinical Trial

Refine the following sub-sections to accurately and specifically describe how this imaging protocol interfaces with the rest of your clinical trial. E.g. what are the specific utilities of the imaging protocol in your trial.

Utilities and Endpoints of the Imaging Protocol

This image acquisition and processing protocol is appropriate for quantifying the volume of a solid tumor of the lung, and longitudinal changes in volume within subjects.

<<It has been suggested that the UPICT Protocol should state that it’s useful for quantifying, but not make any statements about expected accuracy and repeatability; such statements can be made in the QIBA Profile document.  Should the following paragraph stay or go then?>>
Following this protocol is expected to provide volume repeatability of at least 18% (in order to be “twice as sensitive as RECIST”, based on the idea that for uniformly expanding cubes and solid spheres, an increase in the RECIST defined uni-dimensional Longest Diameter of a Measurable Lesion corresponds to an increase in volume of about 72% and to diagnose Progressive Disease at a change of about one half that volume, 36%, the noise needs to be less than about 18%).

This protocol is otherwise agnostic to the clinical settings in which the measurements are made and the way the measurements will be used to make decisions about individual patients with cancer or new treatments for patients with cancer.  Typical uses might include assessing response to treatment, establishing the presence of progression for determining TTP, PFS, etc, or determining eligibility of potential subjects in a clinical trial.

Timing of Imaging within the Clinical Trial Calendar

Describe for each discrete imaging acquisition the timing that will be considered “on-schedule” preferably as a “window” of acceptable timing relative to other events in the clinical trial calendar.  Consider presenting the information as a grid which could be incorporated into the clinical trial calendar.

This protocol does not presume a specific timing.  Generally, per RECIST 1.1, "all baseline evaluations should be performed as close as possible to the treatment start".

Management of Pre-enrollment Imaging

Describe the evaluation, handling and usage of imaging performed prior to enrollment.

Clearly identify purposes for which such imaging may be used: eligibility determination, sample enrichment, stratification, setting the measurement base-line, etc.

(e.g. What characteristics or timing will make the imaging acceptable for the purpose?

Will digitized films be accepted?
        Will low-grade images be annotated and/or excluded from parts of the trial?
        Is there normalization that should be done to improve low-grade priors?

How should such imaging be obtained, archived, transferred, etc.)

To quantify volumes and volume changes with the precision claimed in this protocol, the pre-enrollment image acquisition and processing must meet or exceed the minimum specifications described in this protocol in order to serve as the “baseline” scan.

Management of pre-enrollment imaging, including decisions on whether to accept lower precision or to require a new baseline scan, are left to the Clinical Trial Protocol author.

Management of Protocol Imaging Performed Off-schedule

Describe the evaluation, handling and usage of imaging performed according to the Procedure below but not within the “on-schedule” timing window described in Section 1.2.

(e.g. For what purpose(s) may such imaging be used (for clinical decision-making; for data

analysis; for primary endpoints; for secondary endpoints; for continued subject eligibility

evaluation; to supplement but not replace on-schedule imaging, etc.)?

What characteristics or timing will make the imaging acceptable for the purpose?

Is there normalization that should be done to account for the schedule deviation?

What is the expected statistical impact of such imaging on data analysis?

How should such imaging be recorded, archived, etc.)

This protocol does not presume a specific imaging schedule.  It is intended to measure tumor volume change between two arbitrary time points.

Management of the clinical trial calendar, deviations from the calendar, and potential impacts of deviations or non-uniformity of interval timing on derived outcomes such are Time-To-Progression (TTP) or Progression-Free-Survival (PFS) time are left to the Clinical Trial Protocol author.

Management of Protocol Imaging Performed Off-specification

Describe the evaluation, handling and usage of imaging described below but not performed completely according to the specified Procedure.  This may include deviations or errors in subject preparation, the acquisition protocol, data reconstruction, analysis, interpretation, and/or adequate recording and archiving of necessary data.

(e.g. For what purpose(s) may such imaging be used (for clinical decision-making; for data

analysis; for primary endpoints; for secondary endpoints; for continued subject eligibility

evaluation; to supplement but not replace on-schedule imaging, etc.)?

What characteristics or timing will make the imaging acceptable for the purpose?

Is there normalization that should be done to account for the schedule deviation?

What is the expected statistical impact of such imaging on data analysis?

How should such imaging be recorded, archived, etc.)

Deviation from this specification will likely degrade the quality of measurements.

Management of off-specification imaging, including decisions on whether to accept lower precision or to require repeat scans, are left to the Clinical Trial Protocol author.

Management of Off-protocol Imaging

Describe the evaluation, handling and usage of additional imaging not described below.  This may include imaging obtained in the course of clinical care or potentially for research purposes unrelated to the clinical trial at the local site.

(e.g. For what purpose(s) may such imaging be used (for clinical decision-making; for data

analysis; for primary endpoints; for secondary endpoints; for continued subject eligibility

evaluation; to supplement but not replace on-schedule imaging, etc.)?

What characteristics or timing will make the imaging acceptable for the purpose?

Is there normalization that should be done to account for the schedule deviation?

What is the expected statistical impact of such imaging on data analysis?

How should such imaging be recorded, archived, etc.)

Management of Off-protocol imaging is left to the Clinical Trial Protocol author.

Subject Selection Criteria Related to Imaging

Relative Contraindications and Mitigations

Describe criteria that may require modification of the imaging protocol.

This protocol involves ionizing radiation.  Risk and Safety considerations, e.g. for young children or pregnant women, are referenced in section 13.1.  Local standards for good clinical practice (cGCP) should be followed.

This protocol involves the use of intravenous contrast.  Risk and Safety considerations, e.g. for subjects with chronic renal failure, are referenced in section 13.2.  Local standards for good clinical practice (cGCP) should be followed.  The use of contrast in section 5 assumes there are no known contra-indications in a particular subject.

Absolute Contraindications and Alternatives

There are few, if any, absolute contra-indications to the CT image acquisition and processing procedures described in this protocol.  Local standards for good clinical practice (cGCP) should be followed.

No alternative imaging protocols are currently available to reference.

Site Selection, Qualification and Training

Personnel Qualifications

This protocol does not presume specific personnel or qualifications beyond those normally required for the performance and interpretation of CT exams with contrast.

Technical

Physics

Physician

Other (Radiochemist, Radiobiologist, Pharmacist, etc.)

Imaging Equipment

This protocol requires the following equipment:

- CT scanner with the following characteristics:

  - while multi-slice is not required, it will produce better results.  
Acceptable: Single slice, Target: 16 or greater, Ideal: 64 or greater

  - See section 7 for required acquisition capabilities

  - conforms to the Medical Device Directive Quality System and the Essential Requirements of the Medical Device Directive

  - designed and tested for safety in accordance with IEC 601-1, as well as for ElectroMagnetic Compatibility (EMC) in accordance with the European Union’s EMC Directive, 89/336/EEC

  - Labelled for these requirements, as well as ISO 9001 and Class II Laser Product, at appropriate locations on the product and in its literature

  - CSA compliant

- Measurement Software

  - See section 9 for required capabilities

Participating sites may be required to qualify for, and consistently perform at, a specific level of compliance (See discussion of Bulls-eye Compliance Levels in Appendix C).   Documentation of Acceptable/Target/Ideal Levels of Compliance will appear in relevant sections throughout this document.

Infrastructure

List required infrastructure, such as subject management capabilities, internet capability, image de-identification and transmission capability.

Update this section to reflect your data archival and distribution requirements as described in section 11.

No particular infrastructure is specified.

Quality Control

Procedures

See 12.1.1 for procedures the site must document/implement.

Baseline Metrics Submitted Prior to Subject Accrual

See 12.1.2 for metric submission requirements.

Metrics Submitted Periodically During the Trial

See 12.1.3 for metric submission requirements.

Additional task-specific Quality Control is described in sections below.

Protocol-specific Training

No protocol-specific training is specified beyond familiarity with the relevant sections of this document.

Physician

Physics

Technician

Subject Scheduling

Describe requirements and considerations for the physician when scheduling imaging and other activities, which may include things both related and unrelated to the trial.

Timing Relative to Index Intervention Activity

Define the timing window for imaging relative to the index intervention activity.  This parameter is significantly influenced by the specifics of the index intervention (e.g. the specific pharmaceutical under investigation).

Timing Relative to confounding Activities (to minimize “impact”)

This protocol does not presume any timing relative to other activities.  Fasting prior to a contemporaneous FDG PET scan or the administration of oral contrast for abdominal CT are not expected to have any adverse impact on this protocol.

Scheduling Ancillary Testing

This protocol does not depend on any ancillary testing.

Subject Preparation

Prior to Arrival

No preparation is specified beyond the local standard of care for CT with contrast.

Upon Arrival

Confirmation of subject compliance with instructions

No preparation is specified beyond the local standard of care for CT with contrast.

Ancillary Testing

No ancillary testing is specified beyond the local standard of care for CT with contrast.

Preparation for Exam

No exam preparation is specified beyond the local standard of care for CT with contrast.

Imaging-related Substance Preparation and Administration

Substance Description and Purpose

The use of contrast is not an absolute requirement for this protocol.  However, the use of intravenous contrast material is often medically indicated for the diagnosis and staging of lung cancer in many clinical settings.

Contrast characteristics influence the appearance and quantification of the tumors, therefore a given subject must be scanned with the same contrast agent and administration procedures for each scan, even if that means no contrast is given due to it not being given in previous exams of this subject in this trial.

A subject should be scanned with the same brand of contrast agent for each scan (Target).  Another brand or type of contrast may be used if necessary (Acceptable).

Dose Calculation and/or Schedule

For a given subject, the same contrast dose should be used for each scan (Target).  If a different brand or type of contrast is used, the dose may be adjusted to ensure comparability if appropriate and as documented by peer-reviewed literature and/or the contrast manufacturers’ package inserts (Acceptable).

Site-specific sliding scales that have been approved by local medical staffs and regulatory authorities should be used for patients with impaired renal function (e.g. Contrast Dose Reduction Based On Creatinine Clearance).

Timing, Subject Activity Level, and Factors Relevant to Initiation of Image Data Acquisition

For a given subject, image acquisition should start at the same time after contrast administration for each scan (Target).

Scan delay after contrast administration is dependent upon the both the dose and rate of administration, as well as the type of scanner being used. Contrast administration should be tailored for both the vascular tree as well as optimization of lesion conspicuity in the solid organs. (These guidelines do not refer to perfusion imaging of single tumors.) Generally, since there are multiple concentrations of contrast as well as administration rates and scanning speeds, it is difficult to mandate a specific value. Generally institutional guidelines should be followed so as to optimize reproducibility of the scan technique.

<Should we discuss adjustment of imaging delay and/or timing on the basis of cardiac output as determined by some sort of pre-imaging bolus protocol?>>

Administration Route

Intravenous.

Rate, Delay and Related Parameters / Apparatus

Contrast may be administered manually (Acceptable), preferably at the same rate for each scan (Target), which is most easily achieved by using a power injector (Ideal).

If a different brand or type of contrast is used, the rate may be adjusted to ensure comparability if appropriate and as documented by peer-reviewed literature and/or the contrast manufacturers’ package inserts (Acceptable).

Required Visualization / Monitoring, if any

No particular visualization or monitoring is specified beyond the local standard of care for CT with contrast.

Quality Control

See 12.2.

Individual Subject Imaging-related Quality Control

See 12.3.

Imaging Procedure

Required Characteristics of Resulting Data

This section describes characteristics of the acquired images that are important to this protocol.  Characteristics not covered here are left to the discretion of the participating site.

Additional details about the method for acquiring these images are provided in section 7.2.

Data Content

These parameters describe what the acquired images should contain/cover.

| Parameter | ComplianceLevel * |  |
| --- | --- | --- |
| Anatomic Coverage | Acceptable | Entire Lung Fields, Bilaterally (Lung apices through bases) |
|  | Target | Entire Lung Fields, Bilaterally (Lung apices through adrenal glands) |
| Field of View : Pixel Size | Acceptable | Complete Thorax : 0.8 to 1.0mm |
|  | Target | Outer Thorax : 0.7 to 0.8mm |
|  | Ideal | Rib-to-rib : 0.55 to 0.75mm |

* See Appendix C for a discussion of Bulls-eye Compliance Levels

Field of View affects Pixel Size due to the fixed image matrix size used by most CT scanners.  If it is clinically necessary to expand the Field of View to encompass more anatomy, the resulting larger pixels are acceptable.

Data Structure

These parameters describe how the data should be organized/sampled.

| Parameter | ComplianceLevel * |  |
| --- | --- | --- |
| Collimation Width | Acceptable | 5 to 160mm |
|  | Target | 10 to 80mm |
|  | Ideal | 20 to 40mm |
| Slice Interval | Acceptable | Contiguous or up to 20% overlap |
| Slice Width | Acceptable | <= 5.0mm |
|  | Target | 1.0 to 2.5mm |
|  | Ideal | <= 1.0mm |
| Pixel Size |  | See 7.1.1 |
| Isotropic Voxels | Acceptable | (5:1) Slice width <= 5 x Pixel Size |
|  | Target | (1:1) Slice width = Pixel Size |
| Scan Plane | Acceptable | Same for each scan of subject |
|  | Target | 0 azimuth |
| Rotation Speed | Acceptable | Manufacturer’s default |

* See Appendix C for a discussion of Compliance Level

Collimation Width (defined as the total nominal beam width) is often not directly visible in the scanner interface.  Wider collimation widths can increase coverage and shorten acquisition, but can introduce cone beam artifacts which may degrade image quality.

Slice intervals that result in discontiguous data are unacceptable as they may “truncate” the spatial extent of the tumor, degrade the identification of tumor boundaries, etc.

<Pitch?> impacts dose since the area of overlap results in additional dose to the tissue in that area.  Overlaps of greater than 20% have insufficient benefit to justify the increased exposure.

Slice Width directly affects voxel size along the subject z-axis.  Smaller voxels are preferable to reduce partial volume effects and (likely) provide higher precision due to higher spatial resolution.

Pixel Size directly affects voxel size along the subject x-axis and y-axis.  Smaller voxels are preferable to reduce partial volume effects and (likely) provide higher measurement precision.

Isotropic Voxels are expected to improve the reproducibility of tumor volume measurements, since the impact of tumor orientation (which is difficult to control) is reduced by more isotropic voxels.

Scan Plane may differ for some subjects due to the need to position for physical deformities or external hardware, but should be constant for each scan of a given subject.

Faster Rotation Speed reduces the breath hold requirements and reduces the likelihood of motion artifacts.

Data Quality

These parameters describe the quality of the images.

| Parameter | ComplianceLevel * |  |
| --- | --- | --- |
| Motion Artifact | Acceptable | Minimal (see below) |
|  | Target | No Artifact |
| Noise Metric | Acceptable | std. dev. in 20cm water phantom < 40 HU |
|  | Target |  |
|  | Ideal |  |
| Spatial Resolution Metric | Acceptable | >= 6 lp/cm |
|  | Target | >= 7 lp/cm |
|  | Ideal | >= 8 lp/cm |

* See Appendix C for a discussion of Bulls-eye Compliance Levels

Motion Artifacts may produce false targets and distort the size of existing targets.  “Minimal” artifacts are such that motion does not degrade the ability of image analysts to detect the boundaries of target lesions.

Proposal: Remove Noise Metric and Spatial Resolution Metric until we can properly document the procedure for generating these values on site systems in a reliable fashion.  Work with 1C groundwork activities to test the concept and prepare such procedure specifications.  < When time comes to first publish this protocol, resolve this based on the then current status of 1C>

Noise Metric quantifies the level of noise in the image pixel values.  The procedure for obtaining the noise metric for a given acquisition protocol on a given piece of equipment is described in section XX.   Greater levels of noise may degrade segmentation by image analysis operators or automatic edge detection algorithms.

Noise can be reduced by using thicker slices for a given mAs.  A constant value for the noise metric might be achieved by increasing mAs for thinner slices and reducing mAs for thicker slices.

Spatial Resolution Metric quantifies the ability to resolve spatial details.  It is stated in terms of the number of line-pairs per cm that can be resolved in a scan of an ACR resolution phantom (or equivalent). The procedure for obtaining the spatial resolution metric for a given acquisition protocol on a given piece of equipment is described in section XX.  Lower spatial resolution can make it difficult to accurately determine the borders of tumors.

Spatial resolution is mostly determined by the scanner geometry (not under user control) and the reconstruction algorithm (which is under user control).

Imaging Data Acquisition

Subject Positioning

For a given subject, they may be placed in a different position if medically unavoidable due to a change in clinical status (Acceptable), but otherwise the same positioning should be used for each scan (Target) and if possible, that should be Supine/Arms Up/Feet First (Ideal).

If the previous positioning is unknown, the subject should be positioned Supine/Arms Up/Feet First if possible.  This has the advantage of promoting consistency, and reducing cases where intravenous lines, which could introduce artifacts, go through gantry.

Subject positioning shall be recorded, manually by the staff (Acceptable) or in the image dataset (Target).

Consistent positioning is required to avoid unnecessary variance in attenuation, changes in gravity induced shape, or changes in anatomical shape due to posture, contortion, etc.  Careful attention should be paid to details such as the position of their upper extremities, the anterior-to-posterior curvature of their spines as determined by pillows under their backs or knees, the lateral straightness of their spines, and, if prone, the direction the head is turned.

Factors that adversely influence patient positioning or limit their ability to cooperate (breath hold, remaining motionless, etc.) should be recorded in the corresponding DICOM tags and case report forms, e.g., agitation in patients with decreased levels of consciousness, patients with chronic pain syndromes, etc.

Instructions to Subject During Acquisition

Breath Hold

Subjects should be instructed to hold a single breath at full inspiration (Target) or at least near high % of end inspiration (Acceptable) for the duration of the acquisition.

Breath holding reduces motion which might degrade the image.  Full inspiration inflates the lungs which is necessary to separate structures and make lesions more conspicuous.

Timing/Triggers

(e.g., relative to administration of imaging agents; inter-time point standardization)

For each subject, the time-interval between the administration of intravenous contrast and the start of the image acquisition should be determined in advance, and then maintained as precisely as possible during all subsequent examinations.

(Describe a pre-bolus time to target, account for different circulation,

Acceptable: use a standard time; Target: evaluate “manually”

Ideal: “smart-prep”™ features,

Model-Specific Parameters

Appendix G.1 lists acquisition parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

Archival Requirements for Primary Source Imaging Data

See 11.3.

Imaging Data Reconstruction

These parameters describe general characteristics of the reconstruction.

| Parameter | ComplianceLevel * |  |
| --- | --- | --- |
| Reconstruction Kernel Characteristics | Acceptable | soft to overenhancing |
|  | Target | standard to enhancing |
|  | Ideal | slightly enhancing |
| Reconstruction Interval | Acceptable | <= 5mm |
|  | Target | <= 3mm |
|  | Ideal | <= 1mm |
| Reconstruction Overlap | Acceptable | Contiguous (e.g., 5mm thick slices, spaced 5mm apart or 1.25mm spaced1.25 mm apart) |
|  | Target | 20% Overlap (e.g. 5mm thick slices, spaced 4mm apart or 1.25mm spaced 1mm apart) |

* See Appendix C for a discussion of Bulls-eye Compliance Levels

Reconstruction Kernel Characteristics should be the same for each scan of a given subject.  A softer kernel can reduce noise at the expense of spatial resolution.  An enhancing kernel can improve resolving power at the expense of increased noise.  Moderation on both fronts is recommended with a slight bias towards enhancement.

Reconstruction Interval should be the same for each scan of a given subject.

Reconstruction Overlap should be the same for each scan of a given subject.

• Decisions about overlap should consider the technical requirements of the clinical trial, including effects on measurement, throughput, image analysis time, and storage requirements.

• Reconstructing datasets with overlap will increase the number of images and may slow down throughput, increase reading time and increase storage requirements.

It should be noted that for multidetector row CT (MDCT) scanners, creating overlapping image data sets has NO effect on radiation exposure; this is true because multiple reconstructions having different kernel, slice thickness and intervals can be reconstructed from the same acquisition (raw projection data) and therefore no additional radiation exposure is needed.

<Gary: Based on this discussion, should the UPICT protocol provide advice to those sites with multiple types of scanners as to the characteristics of the preferred scanner – multidetector, multirow, vs. single slice, etc.?>

Model-Specific Parameters

Appendix G.2 lists reconstruction parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

Archival Requirements for Reconstructed Imaging Data

See 11.4.

Quality Control

See 12.4.

Image Post-processing

(e.g. spatial registration, spatial re-orientation, re-slicing, feature enhancement, 3D view generation)

No post-processing shall be performed on the reconstructed images sent for image analysis.  Such processing, if performed, has the potential to disrupt the consistency of the results.

Image Analysis

Each lung lesion shall be characterized as described in this section.

Lesions of interest include: a) small pulmonary nodules surrounded by air; b) small to medium pulmonary nodules surrounded by air and/or with adjacent normal and abnormal (non-neoplastic) anatomic structures; c) large pulmonary masses surrounded by air and/or with adjacent normal and abnormal (non-neoplastic) anatomic structures and/or confluent with mediastinum, chest wall, and diaphragm.

Neoplastic tissue volume? necrotic tissue volume?

<Do we want to say anything about our target being tumors greater than 10mm in Longest Diameter to achieve the repeatability numbers?  Can it be used for smaller tumors?> Depends on response criteria, perhaps leave this to the PI to define when they incorporate this protocol in their study.

Many separate things which should be noted in the analysis but when we will not be defining here.

(Pulmonary hemorage, hemorage into a cystic lesion, etc.)

Input Data to Be Used

The reconstructed images may be used directly since no post-processing is specified.

No other data is required for this Analysis step.

Methods to Be Used

Each lesion shall be characterized by determining the boundary of the lesion (referred to as segmentation) and taking certain measurements of the segmented lesion.

Segmentation may be performed automatically by a software algorithm, manually by a human observer, or semi-automatically by an algorithm working with human guidance/intervention.

Measurements may be performed automatically by a software algorithm, manually by a human observer with “e-calipers”, or semi-automatically by an algorithm working with human guidance/intervention.

It is expected that automated boundary detection algorithms will place segmentation edges with greater precision, accuracy and speed than an operator can draw by hand with a pointing device.  It is also expected that automated algorithms for finding the Longest Diameter (LD) and Longest Perpendicular (LP) within each ROI will have greater speed and precision of measurement than an operator using electronic calipers.  The performance of the algorithms will, however, depend on the characteristics of the lesions may be challenged by complex lung tumors.

For each method of segmentation and measurement a site chooses to use, the baseline intra- and inter-rater reliability for segmentation and for linear measurement shall be measured using the methods described in section 9.6 (and provided with the resulting data?).

<Mention adjudication – but needs to be in the context of specific response criteria> <Repeatability of a measurement vs repeatability of a “cognitive” process> <Detectability of new lesions is quite different from the repeatable measurement of existing “known” lesions>

The segmentation reliability (reproducibility? accuracy? precision?) shall be greater than 80% (Acceptable) and preferably greater than 90% (Target).

The linear measurement reliability (reproducibility? accuracy? precision?) shall be greater than 80% (Acceptable) and preferably greater than 90% (Target).

(do we want to define where tumor measurements should be taken; algorithms to be used; definition of key anatomical points or pathology boundaries; scoring scales and criteria, related annotations)

<Gary: Should we define that whatever segmentation and/or measurement stipulations are used for baseline should be used consistently for all subsequent studies and therefore need to be archived along with the studies?  For example, measuring leading edge to leading edge, outer edge to outer edge, etc.?>

Proposal: Save the following speed performance details for the QIBA Profile.   Speed mostly affects the workflow and different clinical trials will have different workflows/needs so speed is mostly beyond the scope of a UPICT protocol.

(Ask products to state their speed performance/style (batch vs realtime) on a declared CPU specification)(Remind PI to state speed requirements if they have any)(Encourage differentiated offerings that exceed a minimum baseline requirement)

The software can process an algorithm in a specified period of time with a CPU with the following specifications. (Note that the time performance depends on the workflow of the site and slower may be acceptable, for other workflows, batch processing would be unworkable, e.g. if the operator has to validate the result immediately or take it into consideration in the subsequent step).  Software allows segmentation algorithms to be corrected and adjust the segmentation results. The maximum number of mouse clicks to perform segmentation on a lesion will be specified,

Required Characteristics of Resulting Data

While all measurement metrics are surrogates for tumor burden, it is still uncertain which measurement metric is optimal to assess for change.  Accordingly, multiple overlapping measurements are specified here.

For each lesion analyzed, the data shall include:

- lesion volume, in mm3 or mL (volumetric metric)

- (RECIST 1.0, 1.1 instead of “re-interpreting?” but might be nice to clarify ambiguities for more consistent results)(note no difference between 1.0 and 1.1 for lung lesions)

- greatest maximal? diameter, in mm (uni-dimensional metric) (in the volume? in the “best” axial plane)

- greatest maximal diameter and longest perpendicular (LP), in mm (bi-dimensional metric)<in the same plane>

<Need better definitions to clarify whether it’s in the axial plane, etc.><RECIST says in the plane, but other software is trying others><We need to define a baseline

Additional lesion measurements which would be desirable? to obtain are:

- WHO measurement (reference to the paper)

- maximum 3D diameter, in mm

- shape parameters like roundishness or others (should explain how to score/characterized this in the methods section 9.2)

Provide error margins for each measurement?

For each lesion segmentation, the following data shall be provided for the voxels within the segmentation:
(is this for evaluating the quality of the segmentation? If so, should it stay here or go into QC 9.6 and explain how it should be used; or is this something that gets recorded and archived but not sent to the Clinical Trail Center)

- histogram of the HU (Hounsfield Unit) values

- the min, max, mean, and standard deviation of the HU values

- the markup of the lesion boundary (and calculated unidimensional and volumetric measurement? or is the data up above sufficient)

- <Gary: If contrast is used, are we suggesting contrast densitometry or DCE-CT with AUC, upslope, persistence of contrast, etc.?>

Platform-specific Instructions

Appendix G.4 lists parameter values and/or instructions for specific models/versions that can be expected to produce data meeting the requirements of Section 9.3.

(Can we ask Definiens, Kitware, Siemens, …)

Archival and Distribution Requirements

See 11.6.

Quality Control

See 12.6.

Software to provide accuracy of Z% (Petrick phantom data) using a predefined phantom test data set and Y% using a predefined clinical data set (MSK Coffee break data)(Fenimore RIDER data may be good for testing repeatability but not for validating accuracy)(LIDC Consortium [NCI-Sponsored] database may also be useful – 4-6 radiologists marking up each study)

(Also need to calibrate accuracy and repeatability scores against the difficulty of the dataset; if using a static dataset we can use a single score)

(this is done in some studies today.  Issues include size of test dataset; tweaking algorithms to game the test set but not perform as well on others; need general comment about the algorithm working on independent dataset;

Image Interpretation

Describe the diagnostic conclusions of interest to be drawn from the images.

(e.g. progression of disease, presence/absence/degree of pathology, viable tumor vs. necrotic)

While Analysis is primarily about computation; Interpretation is primarily about judgment.  Interpretation may be performed at both the lesional / target level and in the aggregate at the subject level (e.g., in an oncology study each index lesion may be measured in longest diameter during the analysis phase, but in this phase a judgment may be made as to whether there is a new “non-index” lesion; the aggregation of the measured lesions with comparison to previous studies coupled with the judgment as to the presence or absence of a new lesion will result in the RECIST classification at the subject level).

Input Data to Be Used

Describe required input data and any necessary validation or adjustments which should be performed on it. May also specify data which should not be used until after the clinical trial interpretation is recorded.

(e.g. particular image series or views; before and after processing versions of images to evaluate/validate the effects of processing; analysis results)

Methods to Be Used

Describe how the interpretation should be performed.

(e.g. definition of key anatomical points or pathology boundaries; scoring scales and criteria such as BIRADS, interpretation schema such as RECIST, related annotations)

Required Characteristics of Resulting Data

Platform-specific Instructions

Appendix G.5 provides instructions for specific models/versions that can be expected to produce data meeting the requirements of Section 10.3.

Reader Training

Archival Requirements

See 11.7.

Quality Control

See 12.7.

Archival and Distribution of Data

Describe the required data formats, transmission methods, acceptable media, retention periods, …

(e.g. Is the site required to keep local copies in addition to transmitting to the trial repository?  Must all intermediate data be archived, or just final results? At what point may various data be discarded?)

Central Management of Imaging Data

Ideal: electronic transmission of encrypted data over a secure network

Target: electronic transmission with a secure file transfer protocol

Acceptable: courier shipment of physical media containing electronic copies of the data

Note: The submission of films for digitization is not acceptable

Imaging data for analysis at central laboratories should be de-identified according to 11.2 prior to transfer.

Get Merck Clinical Computer Validation and Quality Assurance to propose a passage that can be vetted by other pharma companies

De-identification / Anonymization Schema(s) to Be Used

The de-identification software should be certified as fit-for-purpose by regulatory authorities at both the site of origin and site of receipt.

All personal patient information that is not needed for achieving the specific aims of the trial should be removed.

Pre-specified data, such as height, weight, and in some cases, sex, race, or age, may be retained if it has been approved for use by regulatory authorities. Quality assurance procedures must be performed by the recipient to verify that the images that will be submitted for analysis have been properly de-identified.

Acceptable: Data should be transferred to the "quarantine area" of a "safe harbor" for de-identification by professional research organizations or trained operators using procedures that have been certified by regulatory authorities at both the site of origin and the site of receipt. Quality assurance procedures performed by the recipient should verify that the images that will be submitted for analysis have been properly de-identified. Images that were not properly de-identified prior to receipt by the central archiving facility should be obliterated after assuring that copies conform to quality standards for patient privacy.

Primary Source Imaging Data

This protocol presumes no archiving the pre-reconstruction image data.

Reconstructed Imaging Data

Reconstructed images shall be archived locally, formatted as either DICOM CT image objects or DICOM Enhanced CT image objects.

Retention period and policy is left to the Clinical Trial Protocol author.

Post-Processed Data

No post processing is specified, however if post-processing is performed, the images shall be archived the same as 11.4.

Analysis Results

Segmentation results may be recorded as DICOM Segmentation Objects, or STL Model Files.

Measurement results may be recorded as …

The data described in 9.3 may be provided in any of the following formats:
(we should probably tighten this up)

- DICOM SR

- DICOM RTSS?

- DICOM secondary capture

- XLS, CSV, XML

Interpretation Results

Quality Control

QC Associated with the Site

Quality Control Procedures

Describe required procedures and documentation for routine and periodic QC for the site and various pieces of equipment.

Baseline Metrics Submitted Prior to Subject Accrual

List required baseline metrics and submission details.

Metrics Submitted Periodically During the Trial

List required periodic metrics and submission details.

QC Associated with Imaging-related Substance Preparation and Administration

QC Associated with Individual Subject Imaging

Acquisition System Calibration

Ideal: A protocol specific calibration and QA program shall be designed consistent with the goals of the clinical trial.

This program shall include (a) elements to verify that sites are performing the specified protocol correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address issues relating to both radiation dose and image quality (which may include issues relating to water calibration, uniformity, noise, spatial resolution -in the axial plane-, reconstructed slice thickness z-axis resolution, contrast scale, CT number calibration and others). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.

Target: A protocol specific calibration and QA program shall be designed consistent with the goals of the clinical trial.

This program may include (a) elements to verify that sites are performing the specified protocol correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address a limited set of issues primarily relating dose and image quality (such as water calibration and uniformity). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.

Acceptable: Site staff shall conform to the QA program defined by the device manufacturer.

Phantom Imaging and/or Calibration

[Document the procedure for acquiring images and measuring the image quality metrics in the acquisition protocol description, e.g. uniformity, noise, effective resolution]

Quality Control of the Subject Image and Image Data

QC Associated with Image Reconstruction

QC Associated with Image Processing

QC Associated with Image Analysis

QC Associated with Interpretation

Imaging-associated Risks and Risk Management

Radiation Dose and Safety Considerations

It is recognized that X-ray CT uses ionizing radiation and this poses some small, but non-zero risk to the patients in any clinical trial. The radiation dose to the subjects in any trial should consider the age and disease status (e.g. known disease or screening populations) of these subjects as well as the goals of the clinical trial. These should inform the tradeoffs between desired image quality and radiation dose necessary to achieve the goals of the clinical trial.

Imaging Agent Dose and Safety Considerations

Imaging Hardware-specific Safety Considerations

Management and Reporting of Adverse Events Associated with Imaging Agent and Enhancer Administration

Management and Reporting of Adverse Events Associated with Image Data Acquisition

Acknowledgements and Attributions

This imaging protocol is proffered by the Radiological Society of North America (RSNA) Quantitative Imaging Biomarker Alliance (QIBA) Volumetric Computed Tomography (v-CT) Technical Committee.

The v-CT technical committee is composed of scientists representing the imaging device manufacturers, image analysis software developers, image analysis laboratories, biopharmaceutical industry, academia, government research organizations, professional societies, and regulatory agencies, among others. All work is classified as pre-competitive. A more detailed description of the v-CT group and its work can be found at the following web link: http://qibawiki.rsna.org/index.php?title=Volumetric_CT

The Volumetric CT Technical Committee (in alphabetical order):

- Avila, R	 	Kitware, Inc.

- Buckler, A 	(Chair) Buckler Biomedical LLC

- Dorfman, G 	(UPICT liaison) Cornell

- Fenimore, C 	(WG 1C leader) Nat Inst Standards & Technology

- Ford, R	 	RadPharm, Inc.

- Gottlieb, R 	Roswell Park Cancer Center

- Hayes, W 	Bristol Myers Squibb

- Hillman, B 	Metrix, Inc.

- McNitt-Gray, M University California Los Angeles

- Mozley, PD 	(pharma industry co-chair) Merck & Co Inc/PhRMA

- Mulshine, JL 	Rush

- Nicholson, D 	Definiens, Inc.

- O'Donnell, K 	(IHE liaison) Toshiba

- Petrick, N 	(WG 1A leader) US Food and Drug Administration

- Schwartz, LH 	(academic co-chair)

- Sullivan, DC 	(RSNA Executive Sponsor) Duke University

- Zhao, B., 	Memorial Sloan Kettering Cancer Center

The v-CT Committee is deeply grateful for the remarkable support and technical assistance provided by the staff of the Radiological Society of North America, including Susan Anderson, Linda Bresolin, Joseph Koudelik, and Fiona Miller.

Background Information

The long-term goal of the v-CT committee is to qualify the quantification of anatomical structures with x-ray computed tomography (CT) as biomarkers. The v-CT group selected solid tumors of the chest in patients with lung cancer as its first case-in-point. The rationale for selecting lung cancer as a prototype is that the systems engineering analysis, the groundwork, profile claims documents, and roadmaps for biomarker qualification in this specific setting can serve as a general paradigm for eventually quantifying volumes in other structures and other diseases.

The specific aim of this image acquisition and processing protocol is to describe procedures that seem sufficient for quantifying the volumes of neoplastic masses in the chest that have relatively simple geometric shapes and are adequately demarcated from surrounding non-neoplastic tissues. This particular image acquisition and processing protocol is limited to masses that have measurable diameters of 10 mm or more. The profile claims document on which this protocol is based asserts that following these image acquisition and processing procedures will produce volume measures with less than 18% test-retest variability.

The protocol describes, in predominantly chronological order, procedures that are required to achieve this level of precision.

The protocol describes procedures that should be universally followed in this setting, regardless of the instrument that is used to acquire the data. It also provides links to tables that list specific settings on various makes-and-models of CT scanners.

This protocol should be considered for use in the care of individual patients in conventional medical settings, as well as in clinical trials of new therapies for lung cancer. Table 1 summarizes how staging relates to lung cancer drug therapy approaches, the imaging approaches used in those stages and issues relative to the image requirements.

Table 1: Summary of Image Processing Issues Relative to Stage of Lung Cancer

| Stage | % of Cases | 5-year Survival % | Imaging Focus / Therapy Focus | Imaging Tool | Issues | Thoracic Segment. | Hi-Res |
| --- | --- | --- | --- | --- | --- | --- | --- |
| I | 16% | 49% | Primary tumor / Neo and adjuvant RX | sCT | Small cancers surrounded by air | Can be straight forward | Need |
| II / III | 35% | 15.2% | Primary, hilar and mediastinal lymph nodes / Combined modality | sCT, PET | Larger tumors and nodes abut other structures | Often challenging | Opt. |
| IV | 41% | 3% | Primary/regional nodes and metastatic sites / Chemotherapy | sCT, PET, Bone, Brain Scan | Tumor response often determined outside the chest | Often challenging | Opt. |

Conventions and Definitions

Bulls-eye Compliance Levels

Acquisition parameter values and some other requirements in this protocol are specified using a “bullseye” approach.  Three rings are considered from widest to narrowest with the following semantics:

ACCEPTABLE: failing to meet this specification will result in data that is likely unacceptable for the intended use of this protocol.

TARGET: meeting this specification is considered to be achievable with reasonable effort and equipment and is expected to provide better results than meeting the ACCEPTABLE specification.

IDEAL: meeting this specification may require unusual effort or equipment, but is expected to provide better results than meeting the TARGET.

An ACCEPTABLE value will always be provided for a specified parameter.  When there is no reason to expect better results (e.g. in terms of higher image quality, greater consistency, lower dose, etc.), TARGET and IDEAL values are not provided.

Some protocols may need sites that perform at higher compliance levels do so consistently, so sites may be requested to declare their “level of compliance”.  If a site declares they will operate at the TARGET level, they must achieve the TARGET specification whenever it is provided and the ACCEPTABLE specification when a TARGET specification is not provided.  Similarly, if they declare IDEAL, they must achieve the IDEAL specification whenever it is provided, the TARGET specification where no IDEAL level is specified, and the ACCEPTABLE level for the rest.

<Gary: Maintaining this performance standard would be an obligation for all subjects in the entire clinical trial and should not be a per subject or per test variable.  In some trials, it will be necessary for all sites to perform at a single performance level even if certain sites could perform at a higher level of compliance.  For those sites that achieve a higher level of compliance for SOC imaging, it may be necessary to provide image data at both the mandated level of compliance as well as the higher SOC level of compliance presuming there is no increased radiation risk and the additional time and effort can be appropriately recognized.>

Acquisition vs. Analysis vs. Interpretation

This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge.

Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject.  Analysis is primarily considered to be computational steps that transform the data into information, extracting important values.  Interpretation is primarily considered to be judgment that transforms the information into knowledge.

(The transformation of knowledge into wisdom is beyond the scope of this document.)

Definitions

Review this document and define relevant terms and notations here.

Documents included in the imaging protocol (e.g., CRFs)

(Material the site needs to submit)

Subject preparation

Imaging agent dose calculation

Imaging agent

Image data acquisition

Inherent image data reconstruction / processing

Image analysis

Interpretation

Site selection and Quality Control

Phantom Imaging and Calibration

Associated Documents (derived from the imaging protocol or supportive of the imaging protocol)

e.g. the Imaging Charter, Site Manual, Standard Operating Procedures, etc.

TBD

Model-specific Instructions and Parameters

The following sections provide instructions for various equipment models/versions that are expected to produce data meeting the requirements of the relevant activity.

The presence of specific product models/versions in the following tables should not be taken to imply that those products are fully compliant with the QIBA Profile.  Compliance with a profile involves meeting a variety of requirements of which operating by these parameters is just one.  To determine if a product (and a specific model/version of that product) is compliant, please refer to the QIBA Conformance Document for that product.

Image Acquisition Parameters

The following technique tables list acquisition parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

These technique tables may have been prepared by the submitter of this imaging protocol document, the clinical trial organizer, the vendor of the equipment, and/or some other source.  (Consequently, a given model/version may appear in more than one table.)  The source is listed at the top of each table.

Sites using models listed here are encouraged to consider using these parameters for both simplicity and consistency.  Sites using models not listed here may be able to devise their own acquisition parameters that result in data meeting the requirements of Section 7.1 and conform to the considerations in Section 13.

In some cases, parameter sets may be available as an electronic file for direct implementation on the imaging platform.

Table G.1a

Generic: This represents parameters for a generic CT.  The v-CT committee has not yet completed the process of vetting these parameters as fit for purpose.

Model Y: <description of manufacturer, model, version, etc.>

Model Z: <description of manufacturer, model, version, etc.>

Source: QIBA v-CT Cmte                                                  Date: 2009-mm-dd

| Parameter | Compliance Level* | Generic | Model Y | Model Z |
| --- | --- | --- | --- | --- |
| kVp | Acceptable | 110 to 140 |  |  |
|  | Target | 110 to 130 |  |  |
|  | Ideal | 120 |  |  |
| mAs (medium patient) | Acceptable | 40 to 350 |  |  |
|  | Target | 60 to 200 |  |  |
|  | Ideal | 80 to 160 |  |  |
| Scan Duration | Acceptable | < 30 sec. |  |  |
|  | Target | < 15 sec. |  |  |
|  | Ideal | < 10 sec. |  |  |
| Table Speed | Acceptable |  |  |  |
|  | Target |  |  |  |

* See Appendix C for a discussion of the Levels of Compliance

kVp and mAs should be adjusted as necessary, depending on the body habitus of individual patients.  The values should be consistent for all scans of the same patient.

Scan Duration values are intended to allow completion of the scan in a single breath hold for most/a majority/nearly all subjects respectively.

Table Speed values are intended to yield an IEC Pitch Value of approximately 1 while achieving the corresponding Scan Duration.

Table G.1b

The following table provides sample parameters sets that meet the “Target” Level of Compliance for specific models.

Model A1: <description of manufacturer, model, version, etc.>

Model A2: <description of manufacturer, model, version, etc.>

Source: <submitted by who>                                     Date: <submitted when>

| Parameters | Philips | GE | ACRIN |
| --- | --- | --- | --- |
|  | MxIDT 8000 (Thin) | MxIDT 8000 (Thick) | Ultra | VCT-64 | 6678 |
| Data Content |  |  |  |  |  |
| Anatomic Coverage |  |  |  |  |  |
| Field of View : Pixel Size |  |  |  |  | Rib-to-rib: 0.55-.75mm |
| Data Structure |  |  |  |  |  |
| Collimation Width | 16x0.75 mm | 16x1.5 mm |  |  | (TBA) |
| Slice Interval |  |  |  |  |  |
| Slice Width | 0.8 mm | 5.0 mm |  |  | 1.0 mm |
| Pixel Size |  |  |  |  | 0.55 mm |
| Isotropic Voxels |  |  |  |  | (2:1) |
| Scan Plane |  |  |  |  |  |
| Rotation Speed |  |  |  |  | 0.5 sec |
| Data Quality |  |  |  |  |  |
| Motion Artifact |  |  |  |  |  |
| Noise Metric |  |  |  |  |  |
| Spatial Resolution Metric |  |  |  |  |  |
| Acquisition |  |  |  |  |  |
| Tube Voltage | 120 kVp | 120 kVp |  |  | 120 kVp |
| Exposure | 100 mAs | 100 mAs |  |  | 100 mAs |
| Pitch | 1.2 | 1.2 |  |  |  |
| Reconstruction |  |  |  |  |  |
| Recon. Kernel | Detailed filter | Detailed filter |  |  | Standard |
| Recon. Interval |  |  |  |  |  |
| Recon. Overlap | 50% | 50% |  |  | 20% |

* See Appendix C for a discussion of the Levels of Compliance

Image Reconstruction Parameters

See above.

Post-Processing Instructions

None provided.

Analysis Instructions

None provided.

Interpretation Instructions

None provided.