QIBA DWI Profile v1.06b:

- Clinical context

  - Gain insight into microstructure and composition in tumors using precise measurements of ADC for robust tissue characterization and longitudinal tumor monitoring.

- Claims

  - Using in vivo water tissue mobility can be characterized by measurement of the apparent diffusion coefficient (ADC).  ADC is determined with MRI by applying different b-values to a subject and fitting the resulting signal intensities to an exponential decay.

  - At isocenter, ADC measurements of an ice water phantom should exhibit minimum bias, within 5% of the gold standard value of 1.1 x 10-9 m2/s, regardless of coil type and field strength.

  - When acquiring ADC values in solid tumors greater than 1 cm in diameter or twice the slice thickness (whichever is greater), one can characterize in vivo diffusion with at least a 15% test/retest coefficient of variation, intrascanner and intrareader.

- Profile detail/protocol

  - Executive Summary

    - Word about what is the state of the art in research and clinical trials

    - Why would standardization help

    - Few sentences what this profile is for.

4. Clinical Context

Tumor tissues normally demonstrate an abnormal microstructure and physiology, which might be related to their specific tumor microenvironment and biologic aggressiveness.

Cytotoxic agents and novel molecular tumor therapies early affect the tumor microstructure and physiology, and might result under effective treatment in a tumor necrosis and shrinkage. However, early changes of the tumor microstructure and physiology will not necessarily reflected by classical measurements of size changes (e.g. RECIST), and response classification by these conventional criteria will need several weeks (routinely first follow-up acquired 6-8 weeks after treatment initiation). Since most tumor therapies also cause side effects, and novel molecular drugs are expensive in the preclinical development and daily clinical use, robust non-invasive biomarkers are strongly needed for early assessment of treatment response for patient care, drug discovery, and economic reasons.

Role of DWI in a response to therapy assessment

Diffusion- weighted imaging (DWI) provides qualitative and quantitative information of the tumor microstructure, cellularity, and integrity of the cellular membrane.

Cancer could be detected due to an increased cell density (e.g. lymphoma or prostate cancer), and the calculated "apparent diffusion coefficient" (ADC) might predict tumor aggressiveness and therapy response at baseline.  DWI can also detect relatively small changes in tumor microstructure at the cellular level allowing for quantification of early treatment-induced changes. Very soon, hours to days after therapy initiation, cellular edema could occur, resulting in a transient decrease of the ADC. A few days to weeks after effective therapy, tumor necrosis with a loss of cell membrane integrity and an increase of the extracellular space typically result in an increasing ADC measurement. During the following weeks and months, the tumor may show a shrinkage with a resorption of the free extracellular fluid and fibrotic conversion leading to a decrease of the ADC. However, tumor relapse and regrowth could also result in an ADC reduction, but are typically associated with unchanged or increasing tumor size.

  - Challenges to profile use (biology only)

    - Necrotic components

    - Hemorrhages

    - Lipid-rich tumors

    - Mucin-rich tumors

    - Susceptibility effects

5. Subject scheduling

Baseline examinations should be ideally within 14 days, but at least within 30 days prior to treatment start. DWI should not be performed within 14 days after biopsy, and there should be no other tumor treatment at the meantime. Otherwise measured tumor tissue cellularity may not reflect the status of the tumor prior to initiation of therapy.

Intervals between follow-up examinations should be generally for early treatment monitoring more 24- 48 hours after therapy initiation and for severe therapy related changes more than 2- 4 weeks, but as defined by the clinical trial of the new treatment and determined by current standards for GCP.

6. Subject preparation

For DWI patients should prepared according to the local standard of care (e.g. removal of all metal objects and electronic devices), but no specific patient preparation procedures are required. Patients should be comfortably positioned, in appropriate attire to minimize patient motion and stress, which might affect the imaging results.

7. Imaging Procedure

This section describes the imaging protocols and procedure for conducting a DW-MRI exam.  Suitable localizer (scout) images must be collected at the start of the exam and be used to confirm correct coil placement as well as selection of the appropriate region to image.  This will be followed by routine T2-weighted sequences to delineate the number, location, and limits of tumor extension.

7.1 Required Characteristics of Resulting Data

The DWI portion of the exam will consist of a single-shot echo planar imaging sequence (SSEPI) performed at several b-values.  The details of the protocol and imaging parameters (b, TE, TR, etc.) are body region and organ-specific, and are described in the sections below. In general, in tissues with a substantial perfusion component, the inclusion of b=0 data in the analysis to determine ADC should be avoided, as it biases true ADC.

7.1.1 Region-specific imaging protocol- Abdomen

The abdomen represents imaging challenges due to subject motion from breathing, as well as local fat content.  For these reasons, imaging within the abdomen necessitates the use of fat suppression techniques, as well as motion compensation.  Details vary across organs within the abdomen, and specific details are provided for liver and kidney in the subsections below. General hardware requirements and imaging protocol for the abdomen are listed directly below.

- Pulse sequence: 3D single shot echo planar imaging

- Coils: Transmit- body coil; Receive- Phased array receive coil

- Frequency-encoding direction: The frequency-encoding direction should be adjusted so as to minimize motion artifact.  This decision will be based on the location of the tumor being interrogated and its relationship to moving structures.

- Digitized bit depth: The maximum dynamic range should be utilized, e.g., “extended dynamic range”, or equivalent.

- Receiver Bandwidth: Target: maximum possible; Acceptable: >1000 Hz/voxel

7.1.1.1 Organ-specific imaging protocol- Liver

Imaging parameters specific to the liver are listed in this subsection.

- Motion compensation: Breath-hold or respiratory-triggered motion compensation are ideal practice. Details can be found in section Y.

- Lipid suppression: Spectrally selective methods, such as SPAIR and SPIR are preferred due to the higher SNR as compared to strictly IR-based methods, with the former being ideal at 3 T, due to increased B1 inhomogeneity.

- Number of b-values: Ideal: >3; Target: 3; Acceptable: 2

- Minimum highest b-value strength: Ideal: 800 s mm-2; Target: 500-800 s mm-2; Acceptable: two b-values >100 s mm-2.  b=0 should be avoided, as lower b-values will cause measured ADCs to be weighted by perfusive effects.

- Slice thickness: Ideal: <5 mm; Target: 5-7 mm; Acceptable: 7-9 mm

- Gap thickness: Ideal: 0 mm; Target: 1 mm; Acceptable: >1-2 mm

- Field of view: 300-450 mm FOV along both axes

- Matrix: Set by FOV and desired resolution, typically 160 x 160 or higher

- Number of averages: Ideal: >4; Target: 4; Acceptable: 2-3

- Parallel imaging: >=2

- Plane of view:

- TE: Ideal: <60 ms; Target: 60-75 ms; Acceptable: 75-85 ms

TR: Ideal: >2000 s (depends on anatomic coverage, i.e. # of slices)

7.1.1.2 Organ-specific imaging protocol- Kidney

7.1.1.3 Organ-specific imaging protocol- Lung?

7.1.2 Region-specific imaging protocol- Brain

- Field Strength: 1.5T or 3T.

- Acquisition Sequence: Ideal: 3-orthogonal DW axes, SSEPI, Target: TRSE; Acceptable: single-echo spin-echo

- Coil Type: Ideal>=32; Target: 8-31; Acceptable: 8-channel head coil

- Lipid suppression: On.

- Number of b-values: 2

- Highest b-value strength: Ideal/Target:1000 s mm-2; Acceptable: 700-1200 s mm-2

- Slice thickness: Ideal: <5 mm; Target: 5 mm; Acceptable: 5-6 mm

- Gap thickness: Ideal: 0 mm; Target: 1 mm; Acceptable: >1-2 mm

- Field of view: 220-240 mm FOV along both axes

- Resolution/Acquired Matrix: Ideal: 128x128+ Target in-plane acquired resolution ~2 mm or better, Acquired matrix typically 128 x 128 or higher;

- Number of averages: >2

- Parallel imaging Factor: Target =2

- TE: Ideal/Target: minimum TE; Acceptable: <110 ms

- TR: Ideal: >2000 s (depends on anatomic coverage, i.e. # of slices)

- Receiver Bandwidth: Target: maximum possible; Acceptable: >1000 Hz/voxel

7.1.3 Region-specific imaging protocol- Breast

- Field Strength: 1.5T or 3T.

- Acquisition Sequence: 3-orthogonal DW axes, SSEPI, Target: single-echo spin-echo; Acceptable: TRSE

- Coil Type: Ideal>=16; Target: 8-15; Acceptable: 7-channel breast coil

- Lipid suppression: Spectrally selective methods, such as SPAIR and SPIR are preferred due to the higher SNR as compared to strictly IR-based methods. SPAIR is preferred at 3 T due to B0 inhomogeneity effects.

- Number of b-values: Ideal: >3; Target: 3; Acceptable: 2

- Minimum highest b-value strength: Ideal: 800 s mm-2; Target: 500-800 s mm-2; Acceptable: two b-values >100 s mm-2.

- Slice thickness: Ideal: <5 mm; Target: 4 mm; Acceptable: 3-5 mm

- Gap thickness: Ideal: 0 mm; Target: 1 mm; Acceptable: >1-2 mm

- Field of view: bilateral axial 300-380 mm FOV along both axes

- Resolution/Acquired Matrix: Target in-plane acquired resolution ~2mm, Acquired matrix typically 160 x 160 or higher; Acceptable: 128x128 – 192x192

- Number of averages: Ideal: >4; Target: 4; Acceptable: 2-3

- Parallel imaging Factor: Target >2, Acceptable: 1.5T (1-3); 3T (2-3)

- TE: Ideal: <60 ms; Target: minimum TE; Acceptable: 75-85 ms

- TR: Ideal: >3000 ms; Target: 2000-3000 ms; Acceptable: 2000-8000 ms

- Receiver Bandwidth: Target: maximum possible; Acceptable: >1000 Hz/voxel

7.1.4 Region-specific imaging protocol- Pelvis

7.1.4.1 Organ-specific imaging protocol- Prostate

Generally : at 1.5 T and 3.0T using a 8- to 16-channel heart or pelvic phased array coil, usage of endorectal coils due to their signal inhomogeneties is not recommended for quantitative imaging. Anti-peristaltic drugs (Buscopan®, Glucagon®) should be given. No motion compensation necessary.

•	Number of b-values: Ideal: >3; Target: 3; Acceptable: 2

•	Minimum highest b-value strength: Ideal: 1000-800 s mm-2; Target: 500-800 s mm-2; Acceptable: two b-values >100 s mm-2.

•	Slice thickness: Ideal: <5 mm; Target: 5-7 mm; Acceptable: 7-9 mm

•	Gap thickness: Ideal: 0 mm; Target: 1 mm; Acceptable: >1-2 mm

•	Field of view: 300-400 mm FOV along both axes

•	Matrix: Set by FOV and desired resolution, typically 128 to 192, or higher

•	Reconstructed Voxel size: in-plane resolution: 1.5×1.5 mm to 2.0×2.0 mm at 1.5 T and 1.0×1.0 mm to 1.5×1.5 mm at 3 T

•	Number of averages: Ideal: >4; Target: 4; Acceptable: 2-3

•	Parallel imaging: 2x

•	TE: Ideal: <60 ms; Target: 60-75 ms; Acceptable: 75-90 ms

•	TR: Ideal: > 4000 ms; Target: 3000-4000 ms; Acceptable: 1500-3000 ms

- Receiver Bandwidth: Target: maximum possible; Acceptable: >1000 Hz/voxel

7.1.4.2 Organ-specific imaging protocol- Cervical

7.1.4.1 Organ-specific imaging protocol- Bladder

7.1.4.1 Organ-specific imaging protocol- Rectal

7.1.4.1 Organ-specific imaging protocol- Ovarian

7.1.5 Region-specific imaging protocol- Head and Neck (non-brain)

7.1.6 Region-specific imaging protocol- Whole Body

7.2 Data Content and Structure

All imaging data should be stored in the DICOM format.  All DWI data should be contained in a single series.

7.3 Data quality

A quality review, confirming that all imaging parameters are compliant with the protocol, that the data structure is correct, should be performed before the data are submitted for analysis.

# 8. Respiratory motion compensation in DWI

Three approaches in motion compensated acquisition strategies in body (abdomen and whole body) were reported in the literature review: breath hold, free breathing, respiratory-triggered and navigated.

### 8.1 Breath-hold single shot EPI

The key advantage of breath-hold acquisition is short acquisition time. The entire liver can be covered in one or two breath-holds of up to 20 seconds. Parallel imaging with the EPI sequence allows for short TE (~40-70 ms), thus preserving SNR (1). Theoretically breath hold scans are more effective for evaluating lesion heterogeneity and small lesion ADC. However, single-shot sequences are inherently noisy. Motion artifacts are reduced, but pulsatile flow & motion artifacts remain. Some authors advise combining with cardiac pulse triggering (1), but triggering prolongs scan time. Cardiac pulsations are reported to increase ADC in left lobe of the liver (1).  For good SNR thicker slices are needed (6-8 mm). Breath-hold scans are limited in resolution and in number of b values per breath hold, which may impact ADC accuracy, or limit multi-exponential analysis.

### 8.2 Free breathing with multiple averaging

Free breathing allows multiple b values and thinner slices (4-5 mm), with 3 to 6 minutes scan time for whole liver evaluation (2). Free breathing scans are typically acquired with a higher number of averages (4 to 6) resulting in higher SNR. Cyclical breathing is a coherent motion that doesn’t attenuate signal in liver (2). It is possible to perform MPR and MIP for qualitative evaluation and fusion with anatomical images to combine functional and anatomical information (1).

However, multiple averaging causes slight image blurring. Small lesion ADC and heterogeneity are less accurate because of motion averaging.  The shortcomings of free breathing with multiple averaging raises interest in respiratory (1) and cardiac triggering to improve image registration for ADC measurement.

Free breathing DWI can be extended to multiple stations for whole body DWI, also known as DWIBS (diffusion weighted whole body imaging with background body signal suppression). DWIBS is easier to perform with dedicated whole-body coils (commercially available TIM, dStream for example). Otherwise images can be acquired with the quadrature body coil (with no parallel imaging) or using coil/table sliding solutions (X-Tend Table™ for example).

### 8.3 Respiratory triggering and navigation

Respiratory-triggered scans are acquired using respiratory bellow controls or respiratory navigation with a 2D navigator excitation.

High quality images are acquired with good anatomical detail (2). Liver detection is improved compared to breath-hold DWI (4).  Image quality, SNR, and ADC quantification are improved. Better CNR and decreased scattering of ADC is reported (1).

The penalty of respiratory-triggered acquisition is increased scan time (-> 5-6 minutes), and thus increased chance of patient motion. Risk of pseudo-anisotropy artifact can lead to errors in ADC (5). Cardiac motion causes spin dephasing artifacts in left liver lobe (2).  Cardiac triggering can reduce the cardiac pulsation artifacts (1, 7).

In addition to respiratory triggering using respiratory belts, a navigator echo technique can be used for motion compensation. A pencil-beam excitation pre-pulse is placed at the interface of liver and lung. The diaphragm position is determined from the navigator signal. The diaphragm position can be used to trigger the acquisition in end-expiration, but also to adjust the acquired slice displacement according to the diaphragm position.

In order to circumvent the increased scan due respiratory triggering, Takahara et al (8) introduced a modified, “tracking-only” (TRON) navigated DWI acquisition. With TRON the navigator echo is used only to track and correct for tissue displacement, and not for gating. Thus slices are acquired during the entire breathing cycle. This technique was implement at 1.5T (8) and 3T (9) field strengths.

### References

- Koh DM , Takahara T , Imai Y , Collins DJ. Practical aspects of assessing tumors using clinical diffusion-weighted imaging in the body . Magn Reson Med Sci 2007; 6 : 211 –224.

- Taouli & Koh, Radiology: Volume 254: Number 1—January 2010

- Kwee TC , Takahara T , Ochiai R , Nievelstein RA , Luijten PR . Diffusion-weighted whole body imaging with background body signal suppression (DWIBS): features and potential applications in oncology . Eur Radiol 2008; 18 : 1937 – 1952.

- Parikh T , Drew SJ , Lee VS , et al . Focal liver lesion detection and characterization with diffusion-weighted MR imaging: comparison with standard breath-hold T2-weighted imaging. Radiology 2008; 246 : 812 – 822 .

- Nasu K , Kuroki Y , Fuji H , Minami M . Hepatic pseudo-anisotropy: a specific c artifact in hepatic diffusion-weighted images obtained with respiratory triggering . MAGMA 2007; 20 : 205 – 211.

- Influence of cardiac motion on diffusion-weighted magnetic resonance imaging of the liver.

- Kwee TC, Takahara T, Niwa T, Ivancevic MK, Herigault G, Van Cauteren M, Luijten PR. MAGMA. 2009 Oct;22(5):319-25.

- Takahara T, Kwee TC, Van Leeuwen MS, Ogino T, Horie T, Van Cauteren M, Herigault G, Imai Y, Mali WP, Luijten PR, Diffusion-weighted magnetic resonance imaging of the liver using tracking only navigator echo: feasibility study. Invest Radiol. 2010 Feb; 45(2):57-63.

- Ivancevic MK, Kwee TC, Takahara T, Ogino T, Hussain HK, Liu PS, Chenevert TL. Diffusion-weighted MR imaging of the liver at 3.0 Tesla using TRacking Only Navigator echo (TRON): a feasibility study. J Magn Reson Imaging. 2009 Nov;30(5):1027-33.

    - Imaging post-processing (Mango, Hendrik)

      - Image distortion correction: Distortion correction is desirable, and should be implemented when available. Please refer to the appendix for vendor-specific recommendations.

9. Image Analysis (Alan Jackson)

    - Image formation (obtaining an ADC value) (Ona)

      - Fit

    - Monoexponential fit, pixel-by-pixel

      - Pixelwise, whole Tumor Mean/Median, histogram

    - ROI protocol (Thorsten P)

      - Low b (0-100)/T2W ROI

      - high b for ROI

      - Challenges (ROI vs VOI)

  - Statistical analysis of resulting ADC maps

    - Mean/Median/Histogram

    - Bad pixels

    - Exclusion of ADC=0/NaN in mean/median

  - Archival and distribution of data (Michael)

    - Archiving segmentations

    - Saving segmentation masks (numeric)

10. Quality Control

The following section deals with all aspects of quality control in DWI-MRI studies.  Primary objectives of a DWI QA/QC program are: (a) to confirm DWI acquisition protocol compatibility and compliance across participating centers; (b) assess performance of each MRI system in measuring key DWI/ADC quantities; (c) certification of systems/sites to meet quantitative performance thresholds or identify source of performance deficiency; and (d) establish ongoing quality control.  This includes selection of imaging centers and specific scanners.  In addition, the use of DWI phantom imaging and analysis of phantom data are discussed.  Finally, post DWI acquisition quality assessment is described.  Details of these procedures will necessarily vary for the specifics of each trial thus need adjustment, although the common framework is shared.

Guidelines for appropriate patient selection, tumor selection, and post processing are also discussed below.

10.1	Selection of appropriate imaging centers for DWI studies

Typically sites are selected based on a record of competence in clinical oncology and access to a sufficiently large patient population under consideration in the clinical trial.  Sites should also be competent in standard MRI procedures, DWI methodology applied to the relevant anatomical area(s), other advanced MR procedures that may be employed in the trial (eg. MRS, DCE-MRI), as well as access to quality-maintained clinical MRI systems.  In order to ensure high quality DWI results, it is essential to implement procedures that ensure quality assurance of the scanning equipment and reliable image acquisition methodology.  These processes must be established at study outset and maintained for the duration of the study.  A site “imaging capability assessment” is required and should include evaluation of:

- Appropriate MR equipment and standard QC processes

- Experienced MR technologists

- Experienced MR radiologists

- Experienced MR physicists or MR imaging scientists

- Procedures to assure protocol compliance during the trial

10.1.1	DWI acquisition scanner

DWI studies targeted by this profile require a 1.5T or 3T scanner.  The scanner software/hardware versions should be identified and tracked with time over the course of a clinical trial.  Sites often have multiple scanners at the same or variable software/hardware platforms.  It is beneficial to identify and qualify multiple scanners at a given site if such are available in the event a study-eligible scanner is temporarily unavailable.  However, adherence by the site to a use of a specific scanner or pool of scanners for trial subjects must be established by study design.  Likewise, rules for serial scanning a given trial subject on one or multiple systems must be clearly established.  Means to confirm adherence to study design, in terms of eligible scanner for each patient and time point, should utilize specific scanner identifiers available in the DICOM header.

The MRI scanner must undergo routine QA/QC processes and have a service plan that includes a preventative maintenance schedule appropriate for standard clinical MR applications.  In addition, to assure adequate quantitative MR imaging results study-specific quality control measures are required as detailed below.

10.1.2	Site personnel performing DWI studies

(Analogous to DCE profile)

10.1.3	MR Radiologist or other anatomic experts

(Analogous to DCE profile)

10.1.4	Site protocol compliance

(Analogous to DCE profile)

10.2	Site qualification process

10.2.1	Site readiness

(Analogous to DCE profile)

10.2.2	Scanner qualification

(Analogous to DCE profile)

10.2.3	Phantom imaging

To qualify the MRI scanner a DWI phantom imaging procedure is required.  The DWI phantom must contain one or multiple media having known properties of: (a) diffusion coefficient(s), (b) b-value dependence, and (c) isotropy/anisotropy.  Molecular mobility is a function of temperature (eg. water mobility varies ~2.4%/Co), therefore quantitative diffusion coefficient values require knowledge or control of internal phantom temperature.  DWI phantoms at room temperature are convenient for scanning, although the range in room temperature (~10 Co) requires calibrated internal temperature readouts recorded with phantom scans for look-up-table conversion to known diffusion coefficient values.  Alternatively, phantoms designed with an ice-water bath surrounding diffusion media provide an economical means to establish and maintain temperature control at 0Co for several hours.  A test compartment of water at 0Co has a precisely known diffusion coefficient = 1.1x10-3mm2/s, which is comparable to the ADC value of tissue.  However, ice-water phantoms are less convenient since they require on-site preparation.

10.2.4	Phantom imaging data analysis

Phantom data should be analyzed in a uniform manner and preferably by a central analysis site.  Assurance should be made by the analysis center that the phantom scan orientation is correct, and appropriate phantom positioning was performed.

Clinical DWI protocols may require controlled ranges in geometry values (eg. FOV, slice thickness, quantity of slices) to accommodate a range in patient body habitus.  The DWI phantom physical characteristics and imaging protocol can be designed for similarity with the clinical study protocol, but range in all acquisition settings must be minimized.  A small parameter range for DWI phantom scanning may still be required for protocol compatibility across scanner platforms.

The following performance metrics should be measured via DWI phantom images acquired on each candidate MRI system.  Quantitative performance thresholds of these metrics must be established beforehand, and sites need to meet/exceed these thresholds as an essential step for qualification.  Assuming DWI phantom images are acquired across multiple platforms (i.e. manufacturers and software/hardware versions), the central analysis site must be able to derive performance measures regardless of imaging platform source.  While DICOM offers some uniformity in image format, the stored order of DW images is variable and complicates derivation of ADC values from subsets of images extracted from DWI series.  The phantom QC processing center must be able to import and fully process DWI series from all sources, regardless of image order.  One solution is to customize the image import software module for each platform-specific/order-specific condition for conversion to a common internal structure format.  Once converted, all subsequent analysis routines are independent of image source.

10.2.4.1	ADC bias error

In tissue the “apparent” diffusion coefficient (ADC) represents the distillation of complex biophysical processes so that the concept of a “true” ADC is overly simplistic.  In addition, the relative influence of various biophysical processes depends on data acquisition conditions.  An essential first step to assess ADC bias error of an MRI system is measurement of a medium of precisely known diffusion coefficient.  In addition, the functional dependence of mobility on DWI sequence b-value and diffusion time must be known.  In this regard, simple self-diffusion media having no b-value or diffusion time dependence are preferred.  For such media, the standard formula, ADC0,b = { [ln(DWI0/DWIb)] / b } can be used to generate ADC maps over the b-value range 0 to b.  It is appropriate to first apply a noise threshold filter to mask low SNR pixels that otherwise lead to unreliable ADC results.  Phantom scan b-value(s) are set per protocol, but must at least encompass the range used in the associated clinical trial.  Mean and standard deviation from standard-shaped (round, square, rectangle), fixed-sized ROIs defined in test sample compartments are recorded for each ADC map.  Multiple ROIs over slices or regions on the maps for a given test compartment can be combined to create a volume of interest (VOI) to more fully sample the compartment.  ADC Bias Error is derived from the mean ADC measured over the VOI compared to the known diffusion coefficient (DCtrue) of the medium as,

.

10.2.4.2	ADC random error

ADC Random Error is the standard deviation of ADC pixel values measured over each VOI expressed as a percentage of the VOI mean ADC.  A systematic difference between ADC measured in widely separated VOIs would inflate this metric, therefore ADC Random Error should only be derived from single-region VOIs as,

.

10.2.4.3	ADC b-value dependence

If by first-principles, molecular mobility of the diffusing medium is known to not have b-value or diffusion time dependence, any significant difference in DWI phantom ADC value with b-value is artifactual.  Assuming the DWI phantom protocol is designed to measure ADC over multiple b-value intervals, say b=0(b1 and b=0(b2, the level of artifactual b-value dependence is quantified as,

,

Where b2>b1 and the ADC values represent select VOI means.

10.2.4.4	ADC spatial dependence

MRI systems may have a spatial dependence in measured ADC due to systematic imperfections such as gradient nonlinearity.  These errors should be relatively small and nearly symmetric relative to distance from magnet isocenter.  Susceptibility of a study to error due to ADC spatial dependence varies with the expected range of locations for target tissues, as well as reliability in patient positioning.  It is therefore useful to sample the ADC spatial dependence of MRI systems over a relevant range for the clinical trial.  This can be expressed as the percent range in ADC values for widely separated VOIs, say VOI1 and VOI2, that are within the spatial range anticipated for the clinical study as,

.

10.2.4.5	SNR of DWI

Lastly, overall signal-to-noise of the source DW images is an important performance parameter.  While “signal” is relatively straightforward to measure by the mean pixel intensity within an ROI or VOI, “noise” is more difficult.  Often the standard deviation of pixel intensity within an ROI or VOI drawn in the background (ie. air) is used as an estimate of noise.  Unfortunately, the image background regions are often heavily modulated by MRI reconstruction and filter routines, especially when parallel imaging and intensity normalization techniques are employed.  The nature and degree of background modulation is variable across MRI platforms and it is difficult to enforce standardization.  An alternative to deriving noise from background is to acquire immediately sequential serial images acquired under identical conditions then estimate noise of each pixel by the square root of temporal variance measured over the multiple serial passes.  Note, this measure of variance will also include short-term system instability.  Full 3D noise images are created for each b-value image set.  Identical VOIs are then applied to the signal and noise images to estimate SNR at each DWI b-value as,

.

If only two serial passes are available, the noise image is derived from the pixel-by-pixel subtraction of the two passes, and the SNR statistic is defined as,

For comparative purposes, noise can also be estimated by the pixel standard deviation in a large ROI defined in a ghost-free air background region; although as mentioned this metric is potentially compromised by image intensity modulation routines employed on some systems.

10.2.4.6	Spatial distortions due to B0 inhomogeneity and eddy currents

10.2.4.7	Fat suppression effectiveness and uniformity

10.2.4.8	DWI phantom protocol compliance

The established DWI phantom protocol, including allowed parameter ranges, must be tabulated and compared to acquisition settings used for the MRI system being evaluated.  A DICOM parameter “compare” software module is preferred to automate compliance assessment. Protocol compliance or points of violation must be documented in a report.  A Sample QC report is provided in Figure XXX and YYY.

10.2.5	Ongoing MRI scanner quality control

The initial set of DWI phantom images from each site/system are used for certification that the system met or exceeded performance standards set for the clinical trial.  Sites that desire to use multiple MRI scanners for the study, must have each scanner certified.  Once certified, each MRI system must be re-evaluated by the same DWI phantom test procedure at established intervals.

10.3	Quality control of DWI studies

10.3.1	Determination of suitable tumor lesions

(Analogous to DCE profile)

10.3.2	Selection of target lesion

(Analogous to DCE profile)

10.3.3	Determination of subjects unsuitable for DWI analysis

(Analogous to DCE profile)

10.3.4	Determination of DWI exams unsuitable for DWI analysis

(Analogous to DCE profile)

10.3.5	DWI exam protocol compliance

The established clinical trial DWI protocol, including allowed parameter ranges, must be tabulated and compared to acquisition settings used for each trial DWI dataset submitted for evaluation.  A DICOM parameter “compare” software module is preferred to automate compliance assessment.  Protocol compliance or points of violation must be documented in a report.  A Sample QC report is provided in Figure XXX and YYY.

10.3.6	Editing DWI exams prior to DWI analysis

(Analogous to DCE profile)

Material below are sample tables only …………..

Target performance levels for site certification are summarized as follows:

Sample Table 1: DWI Phantom Target Performance Standards §

| ADC Bias Error (ADC vs True Value) | ADC b-value Dependence (ADC0-600 vs ADC0-800) | ADC Spatial Dependence (Right- vs Left- VOIs) | ADC Random Error (single-side VOI) | SNR of High b-value DWI |
| --- | --- | --- | --- | --- |
| < 10% | < 2% | < 5% | < 5% | > 75:1 |

§ See section X for metric definitions

- Test Procedure – DWI Acquisition

The core DWI test sequence was designed for commonality with the ACRIN 6698 DWI sequence used on patients.  Greater details of the QC protocol are provided in Appendix I, although main acquisition elements of the DWI scan are summarized here.

Sample Table 2: DWI QC test sequence:

| Sequence | TR (ms) | TE (ms) | No. of Averages | Parallel Imaging |
| --- | --- | --- | --- | --- |
| 3-axes DW Single-Shot EPI | >8000 | 80 ( 100 | 2 | Factor 2 @ 1.5T Factor 3 @ 3T |
| FOV (mm) | Matrix | Slice Geometry | Encoding | b-values (s/mm2) |
| 320 x 320 | 160 x 160 acq. 256 x 256 recon. | Bilateral axial 30 slices, 4mm thick, 0 gap | Freq axis = A/P Phase axis = R/L | 0, 100, 600, 800 |

An important aspect of the QC protocol involves collection of four sequential DWI “passes”, where each pass is approximately 3minutes.  This design serves two purposes.  Multiple measurements spanning 12minutes are used to confirm the phantom was at thermal equilibrium.  A clear trend of decreasing ADC with each pass suggests the phantom was not at thermal equilibrium.  Secondly, repeated DWI scans provide an estimate of noise in each DWI pixel by the temporal variance of signal.

- Submission of DWI Phantom Images

Do not de-identify DWI phantom images.  Submit all DWI phantom images, including non-DWI scans and screen-shots to the ACRIN Imaging Core Laboratory via TRIAD.  It is important to note that all images must be in DICOM format.

QC Report Generation:Figure 1:  (a) Schematic of DWI phantom; (b) MRI through central plane of phantom; (c) ADC map on quantitative color scale.  Diffusion coefficient of water at 0 oC ( 1.1 x10-3 mm2/s.

11. Compliance

  - Site

  - Scanner

  - Software

12. Artifacts

Le Bihan et al (1986) introduced the concept of a global statistical parameter, the apparent diffusion coefficient (ADC) replacing the physical diffusion coefficient D. ADC not only depends on the diffusion coefficient, but is also influenced by the actual used technical equipment and parameter settings. This creates a source of various potential artifacts and variations in measuring diffusion effects which are covered in the following section. A more detailed view on the topic and an extensive list of further literature will be found in this excellent review article (Le Bihan2006).

12.1 Potential artifacts and error sources in ADC measurements

12.1.1 Gradient system and eddy current effects

Gradient nonlinearity leads to distortion effects (Bammer R, Markl M, Barnett A, et al. Analysis and generalized correction of the effect of spatial gradient field distortions in diffusion-weighted imaging. Magn Reson Med 2003;50:560–9.) while gradient instabilities may cause spiking from sparks in the gradient system and ghosting due to random phase variations.

Systematic imperfections of MRI systems may lead to spatial dependency of measured ADC values as already described above in III. K .2.4.4..

The main impact of gradient system performance on diffusion weighted MRI are eddy current effects. Generally gradient switching produces changes in the static magnetic field introducing eddy currents. Slowly decaying eddy currents lead to a residual magnetic field at read out time superimposing with the read out gradient. This results in geometrical distortions which can be described as contraction or dilation, overall shift and shear in the final image (Le Bihan2006). Intensities of those distortions increase with stronger gradient pulses. Since diffusion images are calculated from images acquired with different b-factors and therefore with different diffusion gradient strengths they will be affected by differing degree. This leads to blurring of the calculated diffusion images.

There are additional effects of eddy currents on the actual voxel size and the actually b-factor ‘experienced’ by water molecules which are described in more detail here (Le Bihan2006).

To a certain degree eddy current effects may prospectively avoided by so-called ‘self-shielded’ gradients and by a technique called pre-emphasis. ‘Self-shielded’ gradients are meanwhile standard. The concept behind pre-emphasis is to adapted gradient waveforms in such a way that the effective gradient including eddy current effects correspond to a stable gradient. (Schmithorst VJ, Dardzinski BJ. Automatic gradient preemphasis adjustment: a 15-minute journey to improved diffusion-weighted echo-planar imaging. Magn Reson Med 2002;47:208–212.

Terpstra M, Andersen PM, Gruetter R. Localized eddy current compensation using quantitative field mapping. J Magn Reson 1998; 131:139–143

Papadakis NG, Martin KM, Pickard JD, Hall LD, Carpenter TA, Huang CLH. Gradient preemphasis calibration in diffusion-weighted echo-planar imaging. Magn Reson Med 2000;44:616–624.)

Despite any measures to minimize eddy current effects experience taught so far that there will always remain an impact to a certain level. Therefore image distortion may additionally approached using post processing software to correct for image distortion and ADC miscalculation as described in section measured ADC values as already described above in III K i.3a .

12.1.2 Motion artifacts

DWI is based on detection of phase shifts induced by microscopic motion. 
Patient motion creates additional phase shift which are 10 to 100 times larger resulting in ghost artifacts along the phase encoding direction. Source of patient motion include eye movement, pulsation of cerebrospinal fluid, cardiac and respiratory motion as well as intended and unintended movement of extremities, face muscles or other body parts.
While especially for the latter ones precaution and correction measures are rapidly at their limits, there are various strategies (see ‘prospective and retrospective measures’) to approach the others.

12.1.3 FAT suppression artifacts

12.1.4 Dielectric Shielding Artifacts

Especially at 3T a standing wave effect or dielectric artifact may occur. The wavelength of the radiofrequency field at 128MHz is 234cm in free space. The dielectric constant of most body tissue reduces both speed and wavelength of RF waves thus resulting for the later in values of about 30cm. This is in the order of typical field of view settings in many body applications. Visible effects are signal brightening or signal loss due to constructive or destructive interference from standing waves. (Collins CM, Liu W, Schreiber W, Yang QX, Smith MB. Central brightening due to constructive interference with, without, and despite dielectric resonance. J Magn Reson Imaging2005 ; 21:192 -196) This is in the order of typical field of view settings in many body applications. Visible effects are signal brightening or signal loss due to constructive or destructive interference from standing waves. The effect depends on the shape of the body anatomy and is particularly pronounced in pregnant patients and patients with ascites.

12.1.5 EPI artifacts, Ghosting

12.1.5.1 Low bandwidth

Any frequency shift larger than the MR frequency of adjacent voxels will lead to a voxel shift. This effect may be minimized by use of high bandwidth, which also allows reducing effective TE, which helps to compensate S/N loss, due to a higher noise coming along with higher frequencies. One main effect of higher bandwidths is a faster readout leaving less time for signal phase changes due to magnetic field inhomogeneities.

While for EPI sequences with high readout bandwidths this is a minor issue in readout direction, it becomes effective along the phase encoding direction. The consequence may be strong distortions.

A naturally given frequency shift is the one between fat and water which is about 220Hz at 1.5T. Depending on the readout bandwidth this leads to a spatial displacement of their signals in the order of a few pixel. Since the effect depends on B0 the displacement is larger on high field systems leading to severe chemical shift artifacts. However for brain diffusion imaging this problem is relatively easy to address by applying some sort of fat suppression since only the water component is of interest. Outside the brain, especially in abdominal DWI fat suppression is a more challenging task which in combination with diffusion measurements is discussed in more detail in (A. Luna et al., p 22 Diffusion MRI Outside the Brain, 17 DOI 10.1007/978-3-642-21052-5_2, Springer-Verlag Berlin Heidelberg 2012).

12.1.5.2 Eddy currents

Applying long lasting diffusion encoding gradients of large amplitude are predestined to induce eddy currents. Such current loops are potentially induced in any conducting materials including elements of the magnet cryostat and its thermal shields. Since induced currents produce their own magnetic field opposed to the applied one these modify the actual shape of gradient pulses. Therefore echo signals are refocused with a slight delay. Since EPI k-space trajectory goes back and forth this result in misalignment of even and odd echoes in k-space. Reconstructed images of such data display a ghost image by half the FOV, also called N/2 ghost. Besides ghosting eddy currents may also produce image distortion (Poupon C, Clark CA, Frouin V, et al. Regularization of diffusion based direction maps for the tracking of brain white matter fascicles. Neuroimage 2000;195.12:184–195.).

12.1.6 Vibration artifacts

DWI with partial Fourier coverage may be affected by localized signal-loss caused by local phase ramps in the image domain which shift the apparent k-space center for a particular voxel outside the covered region (Gallichan D, Scholz J, Bartsch A, Behrens TE, Robson MD, Miller KL. Hum Brain Mapp. 2010 Feb;31(2):193-202. doi: 10.1002/hbm.20856.). The artifact may be avoided by full k-space acquisition. Subsequent increases in TE can be addressed by employing parallel acceleration.

  - Diffusion analysis and interpretation pitfalls

12.2 Prospective measures

  - Perfusion at low b-value

  - Subject motion

One approach to address macroscopic movement is synchronizing the acquisition with the source of motion which may be done with ECG gating and respiration gating. Another way to deal with motion is monitoring it with navigator echoes followed by phase correction of the DWI data. Both approaches may improve data quality significantly. However they have their issues and it’s not always easy to apply them effectively.

The most effective measure is to avoid additional motion effects during data acquisition. This may be achieved by avoiding any motion at all or by measuring faster than effects of motion may effectively compromise the acquired data. The first strategy may only be suitable in animal studies using anesthesia. In human studies EPI sequences with single shot acquisitions in the order of 100msec are the sequence of choice to freeze macroscopic motion usually present in such examinations. However this is only the solution for the single EPI image. Regarding a complete DWI data set, problems may arise at the time of ADC calculation since voxel of subsequent images supposed to origin from the same location are in fact slightly displaced. Especially in the area of tissue interfaces this may lead to signal variations due to partial volume effects.

In abdominal DWI respiratory and cardiac motion is even more critical due to their proximity to the area investigated. Motion control strategies including free breathing, breath-hold acquisition, ECG gating, navigator echoes and combinations of them are discussed in (A. Luna et al., p 22 Diffusion MRI Outside the Brain, 17 DOI 10.1007/978-3-642-21052-5_2, Springer-Verlag Berlin Heidelberg 2012) and in ‘imaging procedure’ section in III. g.ii.1

        - Acq. Plane

        - Image artifacts (wrap, metal, etc…)

      - Retrospective measures

  - Registration methods

A general approach to correct for image distortions is to warp the individual images to a common template (Haselgrove JC, Moore JR. Correction for distortion of echoplanar images used to calculate the apparent diffusion coefficient. Magn. Reson. Med. 1996; 36(6): 960–964.  Bastin ME. Correction of eddy current-induced artefacts in diffusion tensor imaging using iterative cross-correlation. Magn. Reson. Imag. 1999; 17(7): 1011–1024.).

An extension to this is using a mutual information criterion (Poupon C, Mangin J-F, Frouin V, Regis F, Poupon C, Pachot- Clouard M, Bihan DL, Bloch I. Regularization of MR diffusion tensor maps for tracking brain white matter bundles. Proceedings of MICCAI’98. Springer: Berlin, 1998; 489–498.)

Employing additional available information is the basis of directly mapping eddy current induced fields (Horsfield MA. Mapping eddy current induced fields for the correction of diffusion-weighted echo planar images. Magn. Reson. Imag. 1999; 17: 1335–1345.  Bastin ME, Armitage PA. On the use of water phantom images to calibrate and correct eddy current induced artefacts in MR diffusion tensor imaging. Magn. Reson. Imag. 2000; 18: 681–687.) and of methods aiming to correct the phase map by modeling eddy current effects on it (Horsfield MA. Mapping eddy current induced fields for the correction of diffusion-weighted echo planar images. Magn. Reson. Imag. 1999; 17: 1335–1345.  Bastin ME, Armitage PA. On the use of water phantom images to calibrate and correct eddy current induced artefacts in MR diffusion tensor imaging. Magn. Reson. Imag. 2000; 18: 681–687.).

  - Adherence to imaging protocols

### References

- Denis Le Bihan, et al. Artifacts and Pitfalls in Diffusion MRI, JMRI 24:478–488 (2006)

- Bammer R, Markl M, Barnett A, et al. Analysis and generalized correction of the effect of spatial gradient field distortions in diffusion-weighted imaging. Magn Reson Med 2003;50:560–9.

- Le Bihan D, Breton E, Lallemand D, Grenier P, Cabanis E, Laval Jeantet M. MR Imaging of intravoxel incoherent motions: application to diffusion and perfusion in neurologic disorders. Radiology 1986;161:401–407.

- Schmithorst VJ, Dardzinski BJ. Automatic gradient preemphasis adjustment: a 15-minute journey to improved diffusion-weighted echo-planar imaging. Magn Reson Med 2002;47:208–212.

- Terpstra M, Andersen PM, Gruetter R. Localized eddy current compensation using quantitative field mapping. J Magn Reson 1998; 131:139–143

- Papadakis NG, Martin KM, Pickard JD, Hall LD, Carpenter TA, Huang CLH. Gradient preemphasis calibration in diffusion-weighted echo-planar imaging. Magn Reson Med 2000;44:616–624.)

  - Imaging-associated risks and risk management (MIchael)

- Appendices

  - Perfusion effects in various tissues

  - Vendor-specific protocol details

    - Sequence type (Dave [Siemens], Greg [Philips], Ed? [GE])

    - FOV, matrix size, slice thickness

    - 3 orthogonal gradients for each b-value>0

    - Single vs double echo

    - Single vs multishot

    - Parallel imaging

    - Fat-suppression techniques, per region, 1.5 vs 3 T

  - Conventions and Definitions

    - Glossary

- ADC: Apparent Diffusion Coefficient - a quantity (in mm2/s units) representative of water mobility in a medium, including tissue

- DW / DWI: Diffusion-Weighted / Imaging – incorporation of additional gradient pulses to elicit image contrast dependent on water mobility

- SSEPI: Single-Shot Echo-Planar-Imaging – a rapid sequence that encodes data for a given slice in ~100ms to mitigate tissue motion artifact

- b-value: a quantity (in s/mm2 units) representative of the degree to which an MRI sequence is sensitive to water mobility contrast

- SPIR: SPectral Inversion Recovery – incorporation of additional spectrally-selective RF pulse to suppress signal from fat.

- SPAIR: SPectral Attenuated Inversion Recovery – incorporation of additional spectrally-selective RF pulse designed for greater fat suppression uniformity

- STIR: Short inversion Time Inversion Recovery - incorporation of additional non spectrally-selective RF pulse and a specific inversion time suppress signal from fat

- Single-Echo DWI (aka, mono-polar DWI; alternative to TRSE): a DW sequence using single 180o refocusing pulse where pre-180o and post-180o diffusion gradient pulse polarities are the same

- Twice Refocussed Spin Echo (TRSE) DWI (aka, bi-polar DWI, double-echo DWI; alternative to Single-Echo DWI): a DW sequence using two 180o refocusing pulses where diffusion gradient pulses have before, between, and after the 180o pulses have opposite polarities

- Parallel Imaging Factor:

- TR:

- TE:

- FOV:

- Receiver Bandwidth:

    - Sequences

# Comments

Pancreas?

Combine FOV/Res/Matrix into one bullet point

>2 is a placeholder. What are vendor specs?

PE direction?  Perhaps place into Appendix regarding vendor-specific protocol.

Combine FOV/Res/Matrix into one bullet point

Literature may be lacking?

May break up into prostate and non-prostate organ-specific

Literature?

May merge together, also include mets

Will review test/restest literature for ROI protocol

Split abbreviations and defintiions into separate subsections