## Scan FOV1
QIBA 1C Performance Based Protocol Worksheet
TEST SITE | FDA
Step 1 - scanner description
Scanner Manufacturer | Philips
Scanner Model | 16-row IDT mx8000
Site id for scanner | FDA
Software version | Host V3.2.5.1746, Console_1.1.0.5366 18-May-2005
Date of Testing | 1027, 2010
Steps 2-4: Scanning Parameters that Meet Performance Metrics
Step 2: kVp selected | 120
Step 3: Scan Field of View (SFOV) or Patient Size | 350
Step 4: Reconstructed Slice width Selected (in mm) | 1
Step 5: Collimation Rotation Time and Pitch
Number of Data Channels (N) | 16
Width of Each Data Channel (T, in mm) | 0.75
Gantry Rotation Time in seconds | 0.75
Pitch (close to 1.0 is preferred) | 1
Time to cover a 35 cm long thorax (desired < 15 seconds) | 23.3
Step 6: Iterating on Recon Kernels to obtain 6 lp/cm
If using mA, then mA setting
(if using mA, then effective mAs equals this value):
if using mAs/rotation (philips) or effective mAs (Siemens), the record setting here | 120
Recontruction Filter setting (name) - Iteration 1 | C
Size of line pairs seen with this filter (Level =1100, Win=100) # | 7
Recontruction Filter setting (name) - Iteration 2 | B
Size of line pairs seen with this filter (Level =1100, Win=100) | 6
Recontruction Filter setting (name) - Iteration 3
Size of line pairs seen with this filter (Level =1100, Win=100)
Name of Filter that meets 6 lp/cm performance criteria | B
Provide reconstructed images (both DICOM and JPEG) of ACR phantom section 3 (water-equivalent section) with this recon kernel.
Step 7: Iterating on mA (or mAs) settings to obtain std dev of 17 +/1-1 HU
mA or mAs or Eff. mAs setting - Iteration 1 | 120
Standard deviation measured on image of ACR phantom | 11.7
mA or mAs or Eff. mAs setting - Iteration 2 | 50
Standard deviation measured on image of ACR phantom | 17.5
mA or mAs or Eff. mAs setting - Iteration 3
Standard deviation measured on image of ACR phantom
mA or eff. mAs setting that meets Standard Deviation Performance criteria | 50
Scanner reported CTDIvol (in mGy) using this mA or eff. mAs setting | 3.7
Provide reconstructed images (both DICOM and JPEG) of ACR phantom section 4 (spatial resolution) with this recon filter
GUIDANCE ON ESTIMATING RESOLUTION FROM GRILLE PATTERNS
#Figure: Guidance in determining the resolution in a scan of the ACR accreditation phantom. This image has a reported resolution of 6 lp/cm. The blurred lines of the lower right grille indicate the resolution is < 7 lp/cm. The lines are adequately resolved in the 6 lp/cm grille pattern.

## Scan FOV2 (GE & Toshiba)
QIBA 1C Performance Based Protocol Worksheet - FOV2
TEST SITE | Duke / Johns Hopkins
Step 1 - scanner description
Scanner Manufacturer | GE / Toshiba
Scanner Model
Site id for scanner
Software version
Date of Testing | YYYY XX 2010
Steps 2-4: Scanning Parameters that Meet Performance Metrics
Step 2: kVp selected | 120
Step 3: Scan Field of View (SFOV) or Patient Size
Step 4: Reconstructed Slice width Selected (in mm) | 1
Step 5: Collimation Rotation Time and Pitch
Number of Data Channels (N)
Width of Each Data Channel (T, in mm)
Gantry Rotation Time in seconds
Pitch (close to 1.0 is preferred)
Time to cover a 35 cm long thorax (desired < 15 seconds) | #DIV/0!
Step 6: Iterating on Recon Kernels to obtain 6 lp/cm
If using mA, then mA setting
(if using mA, then effective mAs equals this value): | #DIV/0!
if using mAs/rotation (philips) or effective mAs (Siemens), record the setting here
Recontruction Filter setting (name) - Iteration 1
Size of line pairs seen with this filter (Level =1100, Win=100)
Recontruction Filter setting (name) - Iteration 2
Size of line pairs seen with this filter (Level =1100, Win=100)
Recontruction Filter setting (name) - Iteration 3
Size of line pairs seen with this filter (Level =1100, Win=100)
Name of Filter that meets 6 lp/cm performance criteria
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter
Step 7: Iterating on mA (or mAs) settings to obtain std dev of 17 +/1-1 HU
mA or mAs or Eff. mAs setting - Iteration 1
Standard deviation measured on image of ACR phantom
mA or mAs or Eff. mAs setting - Iteration 2
Standard deviation measured on image of ACR phantom
mA or mAs or Eff. mAs setting - Iteration 3
Standard deviation measured on image of ACR phantom
mA or eff. mAs setting that meets Standard Deviation Performance criteria
Scanner reported CTDIvol (in mGy) using this mA or eff. mAs setting
Provide reconstructed image of ACR phantom section 4 (spatial resolution)
with this recon filter