QIBA Profile:

CT Tumor Volume Change (CTV-1)

Version 2.3b

11 April 2014
Status: Reviewed Draft (Public Comments Addressed)

Table of Contents

Closed Issues:	4

1. Executive Summary	7

2. Clinical Context and Claims	8

Utilities and Endpoints for Clinical Trials	8

Claim:  Measure Change in Tumor Volume	8

3. Profile Details	8

3.1. Subject Handling	10

3.1.3.1 Discussion	11

3.1.3.2 Specification	11

3.1.4.1 Discussion	12

3.1.4.2 Specification	13

3.1.5.1 Discussion	13

3.1.5.2 Specification	13

3.2. Image Data Acquisition	14

3.3. Image Data Reconstruction	17

3.X. (Image) QA	20

3.4. Image Analysis	21

4. Compliance Procedures	25

4.1. Assessment Procedure: In-plane Spatial Resolution	25

4.2. Assessment Procedure: Voxel Noise	25

4.3. Assessment Procedure: 	26

4.4. Assessment Procedure: Tumor Volume Change Variability (Image Analysis Tool)	27

4.4.1. Test Image Set	28

4.4.2 Calculate Descriptive Statistics	29

4.2. Performance Assessment: Image Acquisition Site	32

References	35

Appendices	40

Appendix A: Acknowledgements and Attributions	40

Appendix B: Background Information	41

B.1 QIBA	41

B.2 CT Volumetry for Cancer Response Assessment: Overview and Summary	42

B.3 Detailed Literature Review by Indication	46

Appendix C: Conventions and Definitions	57

Appendix D: Model-specific Instructions and Parameters	58

# Closed Issues:

The following issues have been considered closed by the technical committee.  They are provided here to forestall discussion of issues that have already been raised and resolved, and to provide a record of the rationale behind the resolution.

| 1 | Q. Is the claim appropriate/supported by the profile details, published literature, and QIBA groundwork? Is it stated in clear and statistically appropriate terms? A. Basically, yes. Claim reworded to be clear and statistically appropriate. The concept of “levels of confidence” has been introduced (See separate documents and process). Claim seems to be appropriate for the “Reviewed” level of confidence. In terms of anatomy, it is recognized that the acquisition protocols and processing will not be appropriate for all types of tumors in all parts of the body, however it is felt that the conspicuity requirements will make it clear to users of the profile which anatomy is not included. E.g. brain tumors will clearly not have sufficient conspicuity. Despite the selection of the acquisition parameters, it is expected that the segmentation algorithms will be able to handle the breadth. |
| --- | --- |
| 2 | Q. What kind of additional study (if any is needed) would best prove the profile claim? A. Additional study (as described in the evolving Levels of Confidence document) would provide increased confidence. With this stabilized specification QIBA CT can proceed to such testing. |
| 3 | Q. How do we balance specifying what to accomplish vs how to accomplish it? E.g. if the requirement is that the scan be performed the same way, do we need to specify that the system or the Technologist record how each scan is performed? If we don’t, how will the requirement to “do it the same” be met? A: Have made revisions to text to try to achieve an appropriate balance. The details of compliance testing are still not complete and will require further work in future drafts of the profile. |
| 4 | Q. Should there be a “patient appropriateness” or “subject selection” section? A. The claim is conditioned upon the lesion being measurable (and criteria are listed) and a section describes characteristics of appropriate (and/or inappropriate) subjects. |
| 5 | Q. Does 4cm/sec “scan speed” preclude too many sites? A. No. Most 16-slice (and greater) scanners would be able to achieve this (although due to an idiosyncracy of the available scan modes, the total collimation needs to be dropped to 16mm rather than 20mm) Some examples that would meet this include: (a) 16 x 1mm collimation with 0.5 second rotation time and pitch ³ 1.25 OR (b) 16 x 1mm collimation with 0.4 second rotation time and pitch ³ 1 OR (c) 16 x 1.25 mm collimation with 0.5 second rotation time and pitch ³ 1 OR (d) 16 x 1.5mm collimation with 0.5 second rotation time and pitch ³ .833 Keep in mind that 16 x 0.75 mm collimation would require (i) pitch > 1.67 at 0.5 second rotation time (which breaks the Pitch< 1.5 requirement OR (ii) pitch > 1.33 at 0.4 second rotation time (which is fine) A 4cm/sec threshold is needed since it would likely alleviate potential breath hold issues. Because the reconstructed image thickness allowed here was > 2 mm, all of the above collimation settings would be able to meet both the breath hold requirements as well as the reconstructed image thickness requirements. |
| 6 | Q. What do we mean by noise and how do we measure it? A. Noise means standard deviation of a region of interest as measured in a homogeneous water phantom. FDA has starting looking at Noise Power Spectrum in light of recent developments in iterative reconstruction and an interest in evaluating what that does to the image quality/characteristics. QIBA should follow what comes out of those discussions, but since FDA is not mandating it and since few systems or sites toda are in a position to measure or make effective use of it, this profile will not mandate it either. It has promise though and would be worth considering for future profile work. |
| 7 | Q. Is 5HU StdDev a reasonable noise value for all organs? A. No. Will change to 18HU. Not sure where the 5 HU standard deviation came from. The 1C project used a standard deviation of 18HU. At UCLA, our Siemens Sensation 64 will yield a standard deviation of 17 HU for: a. 120kVp, 50 eff. mAs, 1 mm thickness, B30F filter To get this down to 5 HU would require: a. Increasing the eff. mAs to 550, OR b. Increasing the slice thickness to 2 mm AND increasing eff. mAs to 275 |
| 8 | Q. Are there sufficient DICOM fields for all of what we need to record in the image header, and what are they specifically? A. For those that exist, we need to name them explicitly. For those that may not currently exist, we need to work with the appropriate committees to have them added. |
| 9 | Q. Have we worked out the details for how we establish compliance to these specifications? A. Not completely. We are continuing to work on how this is to be accomplished but felt that it was helpful to start the review process for the specifications in parallel with working on the compliance process. |
| 10 | Q. What is the basis of the specification of 15% for the variability in lesion volume assessment within the Image Analysis section, and is it inclusive or exclusive of reader performance? A. For the basis, see the paragraph below the table in Section B.2. It includes reader performance. Allocation of variability across the pipeline (shown in Figure 1) is fraught with difficulty and accounting for reader performance is difficult in the presence of different levels of training and competence among readers. Input on these points to help with this is appreciated (as is also the case for all aspects of this Profile). |
| 11 | Q. Should we specify all three levels (Acceptable, Target, Ideal) for each parameter? A. No. As much as possible, provide just the Acceptable value. The Acceptable values should be selected such that the profile claim will be satisfied. |
| 12 | Q. What is the basis for our claim, and is it only aspirational? A. Our claim is informed by an extensive literature review of results achieved under a variety of conditions. From this perspective it may be said to be well founded; however, we acknowledge that the various studies have all used differing approaches and conditions that may be closer or farther from the specification outlined in this document. In fact the purpose of this document is to fill this community need. Until field tested, the claim may be said to be “consensus.” Commentary to this effect has been added in the Claims section, and the Background Information appendix has been augmented with the table summarizing our literature sources. |
| 13 | Q. What about dose? A. A discussion has been added in Section 2 to address dose issues. |
| 14 | Q. Are there any IRB questions that should be addressed? A. The UPICT protocol that will be derived from this Profile will flush out any IRB issues if they exist. |
| 15 | Q. What mechanisms are suggested to achieve consistency with baseline parameters? A. Basically manual for now. In the future we can consider requiring the parameters be stored in the DICOM image headers or (future) DICOM Protocol Objects, and require systems be able to query/retrieve/import such objects to read prior parameters. |

# 1. Executive Summary

X-ray computed tomography provides an effective imaging technique for assessing treatment response in subjects with cancer. Size quantification is helpful to evaluate tumor changes over the course of illness. Currently most size measurements are uni-dimensional estimates of longest diameters (LDs) on axial slices, as specified by RECIST (Response Evaluation Criteria In Solid Tumors).  Since its introduction, limitations of RECIST have been reported. Investigators have suggested that quantifying whole tumor volumes could solve some of the limitations of diameter measures [1-2] and many studies have explored the value of volumetry [3-12].  This document proposes standardized methods for performing repeatable volume measurements.

This QIBA Profile makes claims about the confidence with which changes in tumor volumes can be measured under a set of defined image acquisition, processing, and analysis conditions, and provides specifications that may be adopted by users and equipment developers to meet targeted levels of clinical performance in identified settings.

The claims are based on several studies of varying scope now underway to provide comparison between the effectiveness of volumetry and uni-dimensional longest diameters as the basis for RECIST in multi-site, multi-scanner-vendor settings.

The intended audiences of this document include:

- Technical staff of software and device manufacturers who create products for this purpose

- Biopharmaceutical companies, oncologists, and clinical trial scientists designing trials with imaging endpoints

- Clinical trialists

- Radiologists, technologists, and administrators at healthcare institutions considering specifications for procuring new CT equipment

- Radiologists, technologists, and physicists designing CT acquisition protocols

- Radiologists and other physicians making quantitative measurements on CT images

- Regulators, oncologists, and others making decisions based on quantitative image measurements

Note that specifications stated as “requirements” in this document are only requirements to achieve the claim, not “requirements on standard of care.”   Specifically, meeting the goals of this Profile is secondary to properly caring for the patient.

# 2. Clinical Context and Claims

## Clinical Context

Quantifying the volumes of tumors and measuring tumor longitudinal changes within subjects (i.e. evaluating growth or regression with serially acquired CT scans and image processing).

Compliance with this Profile by relevant staff and equipment supports the following claim(s):

## Claim:  Measure Change in Tumor Volume

Claim for cross-sectional measurements (i.e., at a single timepoint):

Measurand: Tumor Volume when the given tumor is measurable (i.e., tumor margins are sufficiently conspicuous and geometrically simple enough to be recognized on all images; the tumor is separable from other structures of equal density).

- Bias Profile: Measured tumor volume will be linear with known tumor volume for volumes in the range of 0.5 cm3 to 500 cm3 (equivalent diameter of 10 to 100 mm) as evidenced by 95% confidence limits for slope of 1.0±.1, intercept 0.0±.1, and quadratic term  magnitude 0.0±.1;

- Precision Profile:

  - Repeatability coefficient (RC) will be no more than <fill in>  (log units);

      - Two scans performed and evaluated by same compliant scanner, same compliant analysis software, and same compliant reader.

This corresponds to a within-subject coefficient of variation (wCV) of <fill in>%

  - Reproducibility:

    - Reproducibility coefficient (RDC) when two different but compliant readers, on scans acquired by the same scanner model, and evaluated using a single compliant analysis software application will be no more than <fill in> (log units); corresponding to a between-subject coefficient of variation (wCV) of <fill in>% for tumors with equivalent diameter of 20mm.[I forgot why 20 mm was selected].

    - Reproducibility coefficient (RDC) when the same compliant reader evaluates scans acquired from two different but compliant scanner models are used in combination with a single compliant analysis software application will be no more than <fill in> (log units); corresponding to a within-subject coefficient of variation (wCV) of <fill in>% for tumors with equivalent diameter of 20mm.

    - Reproducibility coefficient (RDC) when the same compliant reader evaluates scans from a single compliant scanner model but with two different but compliant analysis software applications will be no more than <fill in> (log units); corresponding to a within-subject coefficient of variation (wCV) of <fill in>% for tumors with equivalent diameter of 20mm.

Claim for Longitudinal change measurements (i.e., across two distinct timepoints):

Measurand: Change in Tumor Volume when the given meets the measurability criteria indicated for cross-sectional measurements at both timepoints, and where the reader uses the locked sequential read paradigm.  Volume change refers to proportional change, where the percentage change is the difference in the two volume measurements divided by the average of the two measurements.  By using the average instead of one of the measurements as the denominator, asymmetries in percentage change values are avoided.[the nodule screening group moved away from this concept to the more standard and practical (t2-t1)/t1]

- No claim for Bias Profile of change measurements is promulgated by this Profile, however, it is the belief of the Committee that having stated the Bias claim for cross-sectional measures as we have makes it sufficiently likely that change estimation would meet formal definition of an unbiased estimator within reasonable scenarios addressed by the Profile;

- Precision Profile:

  - Repeatability coefficient (RC) will be no more than <fill in>  (log units);

      - Two scans performed and evaluated by same compliant scanner, same compliant analysis software, and same compliant reader.

This corresponds to a within-subject coefficient of variation (wCV) of <fill in> %

  - Reproducibility:

    - Reproducibility coefficient (RDC) when two different but compliant readers, on scans acquired by the same scanner model, and evaluated using a single compliant analysis software application will be no more than <fill in> (log units); corresponding to a between-subject coefficient of variation (wCV) of <fill in>% for tumors with equivalent diameter of 20mm average across timepoints.

    - Reproducibility coefficient (RDC) when the same compliant reader evaluates scans acquired from two different but compliant scanner models are used in combination with a single compliant analysis software application will be no more than <fill in> (log units); corresponding to a within-subject coefficient of variation (wCV) of <fill in>% for tumors with equivalent diameter of 20mm average across timepoints.

    - Reproducibility coefficient (RDC) when the same compliant reader evaluates scans from a single compliant scanner model but with two different but compliant analysis software applications will be no more than <fill in> (log units); corresponding to a within-subject coefficient of variation (wCV) of <fill in>% for tumors with equivalent diameter of 20mm average across timepoints.

While the claim has been informed by an extensive review of the literature, it is currently a consensus claim that has not yet been fully substantiated by studies that strictly conform to the specifications given here.  A standard utilized by a sufficient number of studies does not exist to date.  The expectation is that during field test, data on the actual field performance will be collected and changes made to the claim or the details accordingly.  At that point, this caveat may be removed or re-stated.

# 3. Profile Details

The Profile is documented in terms of “Actors” performing “Activities”.

Equipment, software, staff or sites may claim conformance to this Profile as one or more of the “Actors” in the following table.  Compliant Actors shall support the listed Activities by meeting all requirements in the referenced Section.  Failing to comply with a “shall” is a protocol deviation.  Although deviations invalidate the Profile Claim, such deviations may be reasonable and unavoidable as discussed below.

Table 1: Actors and Required Activities

| Actor | Activity | Section |
| --- | --- | --- |
| Acquisition Device | Subject Handling | 3.1. |
|  | Image Data Acquisition | 3.2. |
| Technologist | Subject Handling | 3.1. |
|  | Image Data Acquisition | 3.2. |
|  | Image Data Reconstruction | 3.3. |
| Radiologist | Subject Handling | 3.1. |
|  | Image Analysis | 3.4. |
| Reconstruction Software | Image Data Reconstruction | 3.3. |
| Image Analysis Tool | Image Analysis | 3.4. |

For the Acquisition Device, Reconstruction Software and Image Analysis Tool actors, while it will typically be the vendor who claims the actor is conformant, it is certainly possible for a site to run the necessary tests/checks to confirm compliance and make a corresponding claim.  This might happen in the case of an older model device which the vendor is no longer promoting, but which a site needs a compliance claim to participate in a clinical trial.

The sequencing of the Activities specified in this Profile are shown in Figure 1:

Figure 1: CT Tumor Volumetry - Activity Sequence

The method for measuring change in tumor volume may be described as a pipeline.  Subjects are prepared for scanning, raw image data is acquired, images are reconstructed and possibly post-processed.  Such images are obtained at two (or more) time points.  Image analysis assesses the degree of change between two time points for each evaluable target lesion by calculating absolute volume at each time point and subtracting.  Volume change is expressed as a percentage (delta volume between the two time points divided by the average of the volume at time point 1 and time point t).

The change may be interpreted according to a variety of different response criteria.  These response criteria are beyond the scope of this document.  Detection and classification of lesions as target is also beyond the scope of this document.

The Profile does not intend to discourage innovation.  The above pipeline provides a reference model.  Algorithms which achieve the same result as the reference model but use different methods are permitted, for example by directly measuring the change between two image sets rather than measuring the absolute volumes separately.

The requirements included herein are intended to establish a baseline level of capabilities. Providing higher performance or advanced capabilities is both allowed and encouraged.  The Profile does not intend to limit how equipment suppliers meet these requirements.

This Profile is “lesion-oriented”.  The Profile requires that images of a given tumor be acquired and processed the same way each time.  It does not require that images of tumor A be acquired and processed the same way as images of tumor B; for example, tumors in different anatomic regions may be imaged or processed differently, or some tumors might be examined at one contrast phase and other tumors at another phase.

The requirements in this Profile do not codify a Standard of Care; they only provide guidance intended to achieve the stated Claim.  Although deviating from the specifications in this Profile may invalidate the Profile Claims, the radiologist or supervising physician is expected to do so when required by the best interest of the patient or research subject.  How study sponsors and others decide to handle deviations for their own purposes is entirely up to them.

Since much of this Profile emphasizes performing subsequent scans consistent with the baseline scan of the subject, the parameter values chosen for the baseline scan are particularly significant and should be considered carefully.

In some scenarios, the “baseline” might be defined as a reference point that is not necessarily the first scan of the patient.

## 3.1. Subject Handling

This Profile will refer primarily to “subjects”, keeping in mind that the requirements and recommendations apply to patients in general, and subjects are often patients too.

3.1.1 Timing Relative to Index Intervention Activity

When the Profile is being used in the context of a clinical trial, refer to relevant clinical trial protocol for further guidance or requirements on timing relative to index intervention activity.

3.1.2 Timing Relative to Confounding Activities

This document does not presume any timing relative to other activities.

Fasting prior to a contemporaneous FDG PET scan or the administration of oral contrast for abdominal CT is not expected to have any adverse impact on this Profile.

3.1.3 Contrast Preparation and Administration

### 3.1.3.1 Discussion

Contrast characteristics influence the appearance, conspicuity, and quantification of tumor volumes. 
Non-contrast CT may not permit an accurate characterization of the malignant visceral/nodal/soft-tissue lesions and assessment of tumor boundaries. Therefore, consistent use of intravenous contrast is required to meet the claims of this Profile.

However, the use of contrast material (intravenous or oral) may be not be medically indicated in defined clinical settings or may be contra-indicated for some subjects.  Radiologists and supervising physicians may omit intravenous contrast or vary administration parameters when required by the best interest of patients or research subjects, in which case lesions may still be measured but the measurements will not be subject to the Profile claims.

The following specifications are minimum requirements to meet Profile claims. Ideally, intravenous contrast type, volume, injection rate, use or lack of a "saline chase," and time between contrast administration and image acquisition should be identical for all time points, and the use of oral contrast material should be consistent for all abdominal imaging timepoints.

Recording the use and type of contrast, actual dose administered, injection rate, and delay in the image header by the Acquisition Device is recommended.  This may be by automatic interface with contrast administration devices in combination with text entry fields filled in by the Technologist.  Alternatively, the technologist may enter this information manually on a form that is scanned and included with the image data as a DICOM Secondary Capture image.

### 3.1.3.2 Specification

| Parameter | Actor | Specification |
| --- | --- | --- |
| Use of intravenous contrast | Radiologist | Shall determine if the contrast protocol is appropriate for the subject. |
|  | Technologist | Shall use intravenous contrast parameters consistent with baseline. Shall document the total volume of contrast administered, the concentration, the injection rate, and whether a saline flush was used. |
| IV Contrast Volume | Technologist | Shall not vary the total volume of contrast administered (grams of iodine) by more than 25% between scans for portal venous phase imaging, |
| IV Contrast Injection Rate | Technologist | Shall achieve a contrast injection rate of at least 2ml/sec for arterial phase imaging and shall not allow it to vary by more than 1ml/sec over the course of imaging. |
| Use of oral contrast | Radiologist | Shall determine if the contrast protocol is appropriate for the subject. |
|  | Technologist | Shall use oral contrast parameters consistent with baseline. Shall document the total volume of contrast administered and the type of contrast. |

3.1.4 Subject Positioning

### 3.1.4.1 Discussion

Consistent positioning avoids unnecessary changes in attenuation, changes in gravity induced shape and fluid distribution, or changes in anatomical shape due to posture, contortion, etc.  Significant details of subject positioning include the position of their arms, the anterior-to-posterior curvature of their spines as determined by pillows under their backs or knees, the lateral straightness of their spines. Prone positioning is not recommended.  Positioning the subject Supine/Arms Up/Feet First has the advantage of promoting consistency, and reducing cases where intravenous lines go through the gantry, which could introduce artifacts.

Artifact sources, in particular metal and other high density materials, can degrade the reconstructed volume data such that it is difficult to determine the true boundary of a tumor.  Due to the various scan geometries, artifacts can be induced some distance from the artifact source.  The simplest way to ensure no degradation of the volume data is to remove the artifact sources completely from the patient during the scan, if feasible.  Although artifacts from residual oral contrast in the esophagus could affect the measurement of small tumors near the esophagus, this is not addressed here.

<<Probably need a reference here to a proper guide on how to/not to position such things.  It is NOT simple or immediately clear to technologists so an informative guide would be very useful if we want this to be done effectively.>>

When the patient is supine, the use of positioning wedges under the knees and head is recommended so that the lumbar lordosis is straightened and the scapulae are both in contact with the table. However, the exact size, shape, etc. of the pillows is not expected to significantly impact the Profile Claim.  It is expected that clinical trial documentation or local clinical practice will specify their preferred patient positioning.

Recording the Subject Positioning and Table Heights in the image header is helpful for auditing and repeating baseline characteristics.

Consistent centering of the patient avoids unnecessary variation in the behavior of dose modulation algorithms during scan.

### 3.1.4.2 Specification

| Parameter | Actor | Specification |
| --- | --- | --- |
| Subject Positioning | Technologist | Shall position the subject consistent with baseline. If baseline positioning is unknown, position the subject Supine if possible, with devices such as positioning wedges placed as described above. |
| Artifact Sources | Technologist | Shall remove or position potential sources of artifacts (specifically including breast shields, metal-containing clothing, EKG leads and other metal equipment) such that they will not degrade the reconstructed CT volumes. |
| Table Height & Centering | Technologist | Shall adjust the table height for the mid-axillary plane to pass through the isocenter. Shall position the patient such that the “sagittal laser line” lies along the sternum (e.g. from the suprasternal notch to the xiphoid process). |

3.1.5 Instructions to Subject During Acquisition

### 3.1.5.1 Discussion

Breath holding reduces motion that might degrade the image. Full inspiration inflates the lungs, which separates structures and makes tumors more conspicuous.

Since some motion may occur due to diaphragmatic relaxation in the first few seconds following full inspiration, a proper breath hold will include instructions like "Lie still, breathe in fully, hold your breath, and relax”, allowing 5 seconds after achieving full inspiration before initiating the acquisition.

Although performing the acquisition in several segments (each of which has an appropriate breath hold state) is possible, performing the acquisition in a single breath hold is likely to be more easily repeatable and does not depend on the Technologist knowing where the tumors are located.

### 3.1.5.2 Specification

| Parameter | Actor | Specification |
| --- | --- | --- |
| Breath hold | Technologist | Shall instruct the subject in proper breath-hold and start image acquisition shortly after full inspiration, taking into account the lag time between full inspiration and diaphragmatic relaxation. Shall ensure that for each tumor the breath hold state is consistent with baseline. |
| Image Header | Technologist | Shall record factors that adversely influence subject positioning or limit their ability to cooperate (e.g., breath hold, remaining motionless, agitation in subjects with decreased levels of consciousness, subjects with chronic pain syndromes, etc.). |
|  | Acquisition Device | Shall provide corresponding data entry fields. |

3.1.6 Timing/Triggers

3.1.6.1 Discussion

The amount and distribution of contrast at the time of acquisition can affect the appearance and conspicuity of tumors.

3.1.6.2 Specification

| Parameter | Actor | Specification |
| --- | --- | --- |
| Timing / Triggers | Technologist | Shall ensure that the time-interval between the administration of intravenous contrast (or the detection of bolus arrival) and the start of the image acquisition is consistent with baseline (i.e. obtained in the same phase; arterial, venous, or delayed). Shall … |
| Image Header | Acquisition Device | Shall record actual Timing and Triggers in the image header. |

## 3.2. Image Data Acquisition

3.2.1 Discussion

CT scans for tumor volumetric analysis can be performed on any equipment that complies with the specifications set out in this Profile.  However, we strongly encourage performing all CT scans for an individual subject on the same platform (manufacturer, model and version) which we expect will further reduce variation.

Many scan parameters can have direct or indirect effects on identifying, segmenting and measuring lesions.  To reduce this potential source of variance, all efforts should be made to have as many of the scan parameters as possible consistent with the baseline.

Consistency with the baseline implies a need for a method to record and communicate the baseline settings and make that information available at the time and place that subsequent scans are performed. Although it is conceivable that the scanner could retrieve prior/baseline images and extract acquisition parameters to encourage consistency, such interoperability mechanisms are not defined or mandated here and cannot be depended on to be present or used.  Similarly, managing and forwarding the data files when multiple sites are involved may exceed the practical capabilities of the participating sites.  Sites should be prepared to use manual methods instead.

The goal of parameter consistency is to achieve consistent performance.  Parameter consistency when using the same scanner make/model generally means using the same values.  Parameter consistency when the baseline was acquired on a different make/model may require some “interpretation” to achieve consistent performance since the same values may produce different behavior on different models.  The parameter sets in Appendix D may be helpful in this task.

The approach of the specifications here, and in the reconstruction section, is to focus as much as possible on the characteristics of the resulting dataset, rather than one particular technique for achieving those characteristics.  This is intended to allow as much flexibility as possible for product innovation and reasonable adjustments for patient size (such as increasing acquisition mAs and reconstruction DFOV for larger patients), while reaching the performance targets.  Again, the technique parameter sets in Appendix D may be helpful for those looking for more guidance.

The purpose of the minimum scan speed requirement is to permit acquisition of an anatomic region in a single breath-hold, thereby preventing respiratory motion artifacts or anatomic gaps between breath-holds. This requirement is applicable to scanning of the chest and upper abdomen, the regions subject to these artifacts, and is not required for imaging of the head, neck, pelvis, spine, or extremities.

Coverage of additional required anatomic regions (e.g. to monitor for metastases in areas of likely disease) depends on the requirements of the clinical trial or local clinical practice.  In baseline scans, the tumor locations are unknown and may result in a tumor not being fully within a single breath-hold, making it “unmeasurable” in the sense of this Profile.

Pitch is chosen so as to allow completion of the scan in a single breath hold.

For subjects needing two or more breath-holds to fully cover an anatomic region, different tumors may be acquired on different breath-holds.  It is still necessary that each tumor be fully included in images acquired within a single breath-hold to avoid discontinuities or gaps that would affect the measurement.

Scan Plane (transaxial is preferred) may differ between subjects due to the need to position for physical deformities or external hardware.  For an individual subject, a consistent scan plane will reduce unnecessary differences in the appearance of the tumor.

Total Collimation Width (defined as the total nominal beam width, NxT, for example 64x1.25mm) is often not directly visible in the scanner interface.  Manufacturer reference materials typically explain how to determine this for a particular scanner make, model and operating mode.  Wider collimation widths can increase coverage and shorten acquisition, but can introduce cone beam artifacts which may degrade image quality.  Imaging protocols will seek to strike a balance to preserve image quality while providing sufficient coverage to keep acquisition times short.

Nominal Tomographic Section Thickness (T), the term preferred by the IEC, is sometimes also called the Single Collimation Width.  It affects the spatial resolution along the subject z-axis.

Smaller voxels are preferable to reduce partial volume effects and provide higher accuracy due to higher spatial resolution. The resolution/voxel size that reaches the analysis software is affected by both acquisition parameters and reconstruction parameters.

X-ray CT uses ionizing radiation.  Exposure to radiation can pose risks; however as the radiation dose is reduced, image quality can be degraded.  It is expected that health care professionals will balance the need for good image quality with the risks of radiation exposure on a case-by-case basis.  It is not within the scope of this document to describe how these trade-offs should be resolved.

Anatomic Coverage recording by the Acquisition Device may or may not require the attention of the Technologist.

The acquisition parameter constraints here have been selected with scans of the chest, abdomen and pelvis in mind.

3.2.2 Specification

| Parameter | Actor | Specification | DICOM Tag |
| --- | --- | --- | --- |
| Scan Duration for Thorax | Technologist | Shall achieve a table speed of at least 4cm per second, if table motion is necessary to cover the required anatomy. | Table Speed (0018,9309) |
| Scanogram | Technologist | Shall confirm on the scanogram the absence of artifact sources that could affect the planned volume acquisitions. |  |
| Anatomic Coverage | Technologist | Shall ensure the tumors to be measured and additional required anatomic regions are fully covered. Shall, if multiple breath-holds are required, obtain image sets with sufficient overlap to avoid gaps within the required anatomic region(s), and shall ensure that each tumor lies wholly within a single breath-hold. | Anatomic Region Sequence (0008,2218) |
| Scan Plane (Image Orientation) |  | Consistent with baseline. | Gantry/Detector Tilt (0018,1120) |
| Total Collimation Width |  | Greater than or equal to 16mm. | Total Collimation Width (0018,9307) |
| IEC Pitch |  | Less than 1.5. | Spiral Pitch Factor (0018,9311) |
| Tube Potential (kVp) |  | Consistent with baseline (i.e. the same kVp setting if available, otherwise as similar as possible). | KVP (0018,0060) |
| Nominal Tomographic Section Thickness (T) |  | Less than or equal to 1.5mm. | Single Collimation Width (0018,9306) |
| Acquisition Field of View (FOV) |  | Consistent with baseline. |  |
|  | Acquisition Device | Shall be capable of performing scans with all the parameters set as described above in this table. |  |
| Image Header | Acquisition Device | Shall record actual Field of View, Scan Duration, Scan Plane, Total Collimation Width, Single Collimation Width, Scan Pitch, Tube Potential, Tube Current, Rotation Time, Exposure and Slice Width in the DICOM image header. |  |

## 3.3. Image Data Reconstruction

3.3.1 Discussion

Image reconstruction is modeled as a separate Activity in the QIBA Profile.  Although it is closely related to image acquisition, and is usually performed on the Acquisition Device, reconstruction may be performed, or re-performed, separate from the acquisition.  Many reconstruction parameters will be influenced or constrained by related acquisition parameters.  This specification is the result of discussions to allow a degree of separation in their consideration without suggesting they are totally independent.

Many reconstruction parameters can have direct or indirect effects on identifying, segmenting and measuring lesions.  To reduce this potential source of variance, all efforts should be made to have as many of the parameters as possible consistent with the baseline.

Consistency with the baseline implies a need for a method to record and communicate the baseline settings and make that information available at the time and place that subsequent reconstructions are performed. Although it is conceivable that the scanner could retrieve prior/baseline images and extract reconstruction parameters to encourage consistency, such interoperability mechanisms are not defined or mandated here and cannot be depended on to be present or used.  Similarly, managing and forwarding the data files when multiple sites are involved may exceed the practical capabilities of the participating sites.  Sites should be prepared to use manual methods instead.

Spatial Resolution quantifies the ability to resolve spatial details. Lower spatial resolution can make it difficult to accurately determine the borders of tumors, and as a consequence, decreases the precision of volume measurements.  Increased spatial resolution typically comes with an increase in noise which may degrade segmentation and quantification of tumors. Therefore, the choice of factors that affect spatial resolution typically represent a balance between the need to accurately represent fine spatial details of objects (such as the boundaries of tumors) and the noise within the image. Maximum spatial resolution is mostly determined by the scanner geometry (which is not usually under user control) and the reconstruction kernel (over which the user has some choice).  Resolution is stated in terms of “the number of line-pairs per cm that can be resolved in a scan of a resolution phantom (such as the synthetic model provided by the American College of Radiology and other professional organizations)”. If a followup scan has a significantly different resolution than the baseline, it is likely that the exposure characteristics will change which can affect repeatability.  The impact of partial volume effects can also change, so reasonable consistency of resolution within a given patient is desirable.

Noise Metrics quantify the magnitude of the random variation in reconstructed CT numbers.  Increased levels of noise can make it difficult to identify the boundary of tumors by humans and automated algorithms.  <<Explain why there is a lower bound on noise>>

Some properties of the noise can be characterized by the standard deviation of reconstructed CT numbers over a uniform region in phantom.  Voxel Noise (pixel standard deviation in a region of interest) can be reduced by reconstructing images with greater thickness for a given mAs. A constant value for the noise metric might be achieved by increasing mAs for thinner reconstructed images and reducing mAs for thicker reconstructed images.  The use of a standard deviation metric has limitations since it can vary with different reconstruction kernels, which will also impact the spatial resolution.  While the Noise-Power Spectrum would be a more comprehensive metric, it is not practical to calculate (and interpret) at this time.  Therefore, the standard deviation metric is the preferred measure for Voxel Noise.  It is not expected that the Voxel Noise be measured for each subject scan, but rather the Acquisition Device and Reconstruction Software be qualified for the expected acquisition and reconstruction parameters.

Reconstruction Field of View affects reconstructed pixel size because the fixed image matrix size of most reconstruction algorithms is 512x512.  If it is necessary to expand the field of view to encompass more anatomy, the resulting larger pixels may be insufficient to achieve the claim. A targeted reconstruction with a smaller field of view may be necessary, but a reconstruction with that field of view would need to be performed for every time point. Pixel Size directly affects voxel size along the subject x-axis and y-axis. Smaller voxels are preferable to reduce partial volume effects and provide higher measurement precision.  Pixel size in each dimension is not the same as spatial resolution in each dimension. The spatial resolution of the reconstructed image depends on a number of additional factors including a strong dependence on the reconstruction kernel.

Reconstruction Interval (a.k.a. Slice spacing) that results in discontiguous data is unacceptable as it may “truncate” the spatial extent of the tumor, degrade the identification of tumor boundaries, confound the precision of measurement for total tumor volumes, etc.  Decisions about overlap (having an interval that is less than the nominal reconstructed slice thickness) need to consider the technical requirements of the clinical trial, including effects on measurement, throughput, image analysis time, and storage requirements.

Reconstructing datasets with overlap will increase the number of images and may slow down throughput, increase reading time and increase storage requirements.  For multi-detector row CT (MDCT) scanners, creating overlapping image data sets has NO effect on radiation exposure; this is true because multiple reconstructions having different kernel, slice thickness and intervals can be reconstructed from the same acquisition (raw projection data) and therefore no additional radiation exposure is needed.

Slice thickness is “nominal” since the thickness is not technically the same at the middle and at the edges.

Reconstruction Kernel Characteristics influence the texture and the appearance of tumors in the reconstructed images, which may influence measurements.  A softer kernel can reduce noise at the expense of spatial resolution. An enhancing kernel can improve resolving power at the expense of increased noise.  The characteristics of different tissues (e.g. lung) may call for the use of different kernels, and implementers are encouraged to use kernels suitable for the anatomic region and tissue imaged.  The use of multiple kernels in a single study is not prohibited by the specification below, but any given tumor must be measured on images reconstructed using consistent kernels at each time point.

The use of iterative reconstruction also may influence the texture and the appearance of tumors in the reconstructed images, which may influence measurements. Therefore the effects of iterative reconstruction on quantitative accuracy and reproducibility are not fully understood as of this writing of this Profile version so it is not currently allowed within the Profile Claim.

The stability of HU between time points and its effect on volume measurements is not fully understood as of the writing of this version of the Profile.

3.3.2 Specification

The Reconstruction Software shall be capable of producing images that meet the following specifications. The Technologist shall set up or configure the reconstruction to achieve the requirements in the following table.

| Parameter | Actor | Specification |
| --- | --- | --- |
| In-plane Spatial Resolution |  | Equivalent to 6-8 lp/cm on the ACR phantom and consistent with baseline (i.e. within 1 lp/cm). See 4.1. Assessment Procedure: In-plane Spatial Resolution |
| Voxel Noise |  | Standard deviation of 10-50HU < 18HU measured near the center of a 20cm water phantom. <<And consistent with baseline within ??HU>> See 4.2. Assessment Procedure: Voxel Noise |
| Reconstruction Field of View | Technologist | Shall ensure the Field of View apans at least the full extent of the thoracic and abdominal cavity, but not significantly greater than required to show the entire body and consistent with baseline. |
| Slice Thickness |  | Less than or equal to 1.0-2.5 mm and consistent with baseline (i.e. within 0.5mm). |
| Reconstruction Interval |  | Less than or equal to 2.5 mm and consistent with baseline. |
| Reconstruction Overlap |  | Greater than or equal to 0 (i.e. no gap, and may have some overlap) and consistent with baseline. |
| Reconstruction Algorithm Type |  | Filtered Back-Projection (orrrrrrrrr Iterative) |
| Reconstruction Kernel Characteristics |  | Consistent with baseline (i.e. the same kernel if available, otherwise the kernel most closely matching the kernel response of the baseline). |
| Image Header |  | The Reconstruction Software shall record actual Spatial Resolution, Noise, Pixel Spacing, Reconstruction Interval, Reconstruction Overlap, Reconstruction Kernel Characteristics, as well as the model-specific Reconstruction Software parameters utilized to achieve compliance with these metrics in the image header. |

## 3.X. (Image) QA

<<This is just a thought experiment to see what a QA Activity section might look like.  If it seems promising it will be proposed to the committee for inclusion in the profile.  Right now it’s just an experiment>>

<<If we leave the text in this location, compliant sites are obliged to confirm/check/notice these on every patient exam.  If we move the text down to Section 4, then it could become a procedure for doing a compliance/spot check but you are not obliged to do this check on every exam.>>

<Should likely do this on every exam.  Focus on the need for consistency/reproducibility so you must have consistent image “quality” on specified points.  “Did the images meet the QIBA Profile “quality” requirements?”><Flaws are likely to appear “randomly” so our goal is consistently good.  So if you have some lesions that have flawed imaging at a certain step, you basically “skip” the claim for those lesions at that time point, but if a later study is clean again, you can resume those lesions (from the original clean baseline, etc. >

<Would be useful since it covers real world, but it will be challenging since the real world is so “messy”>

3.X.1 Discussion

Artifacts can variably degrade the ability to assess tumor boundaries as discussed in 3.1.4.1, resulting in poor change measures and repeatability.

Clinical conditions can also degrade the ability to assess tumor boundaries, or influence the structure of the tumor itself.  For example, lung collapse due to pneumothorax can result in architectural changes to the lung surrounding a nodule.

3.X.2 Specification

| Parameter | Actor | Specification |
| --- | --- | --- |
| Artifacts | Technologist & Radiologist | Shall confirm the images are free from artifact due to patient motion, as evinced by no perceptible tram tracking appearance of the bronchioles or blurring of the lung architectural contours with lung windows. Shall confirm the images are free from artifact due to dense objects or materials. <<Note that the algorithms will allow/depend on the user to make edits to correct for these things. If we assume that the radiologist can “recover” then we don’t have to be as strict on the QA step?>><<The discussion to lean less on the radiologist and more on the algorithm (to gain the repeatability/consistency) would mean we can’t depend, so we should make the requirements here fairly strong so the algorithm works reliably without intervention>> <<Does this then depend on the radiologist flagging the presence of the artifacts so you know when you are “out of the profile”>><<Can the algorithms detect the presence of artifacts that would be problems?>> |
| Anatomical Coverage |  |  |
| Image Quality |  | <<Differences in uniformity in a few slices in the middle of the scan (compared to the adjacent slices) have been observed that have caused “gaps” in the segmentation results. Humans might be able to read anatomy around the artifacts but algorithm might fall into the trap.>> |
| Clinical Conditions? | Radiologist | Shall confirm that there are no clinical conditions affecting the measurability of the tumors. In the case of lung tumors this would include absence of pneumonia, pneumothorax… <<Certainly a wide range of conditions if we wanted it to be complete. Might be better to be suggestive and give examples rather than enumerate>><another eg atelectasis (localized lung collapse)> |
| Measurability |  | <<Could move the 10mm size here and also elabourate on the measureability characteristics here without making the claim unwieldy>> |
| Shape? |  | <<If we get a measurement of volume we have what we want, but do algorithms tend to fail on certain shapes?>> |
| Uniformity of density? |  |  |

## 3.4. Image Analysis

3.4.1 Discussion

This Profile characterizes each designated tumor by its volume change relative to prior image sets.

This is typically done by determining the boundary of the tumor (referred to as segmentation), computing the volume of the segmented tumor and calculating the difference of the tumor volume in the current scan and in the baseline scan.

Volume Calculation values from a segmentation may or may not correspond to the total of all the segmented voxels.  The algorithm may consider partial volumes, do surface smoothing, tumor or organ modeling, or interpolation of user sculpting of the volume.  The algorithm may also pre-process the images prior to segmentation.

Segmentation may be performed automatically by a software algorithm, manually by a human observer, or semi-automatically by an algorithm with human guidance/intervention, for example to identify a starting seed point, stroke, or region, or to edit boundaries.

If a human observer participates in the segmentation, either by determining while looking at the images the proper settings for an automated process, or by manually editing boundaries, the settings for conversion of density into display levels (window and level) should either be fixed during the segmentation process or documented so that observers can apply consistent display settings at future scans (or a different observer for the same scan, if multiple readers will read each scan, as for a clinical trial).

Tumor Volume Change Variability, which is the focus of the Profile Claim, is a key performance parameter for this biomarker.  The 30% target is a conservative threshold of measurement variation (the 30% change in the claim is the outside of 95% confidence interval of 15% of measurement variability when sample size is 40 or more).  Based on a survey of clinical studies (See Appendix B.2) the 30% target is considered to be reasonable and achievable.  In Table B.1, the range between the minimum and maximum values in the 95% CI of the measurement difference column is mostly within +/- 15%. Considering a large study from Wang et al using 2239 patients [1], the 95% confidence interval ranged [-13.4%,  14.5%].

Methods that calculate volume changes directly without calculating volumes at individual time points are acceptable so long as the results are compliant with the specifications set out by this Profile.

The Image Analysis Tool should be prepared to process both the current data and previous data at the same time and support matching up the appearance of each tumor in both data sets in order to derive volume change values.  Although it is conceivable that they could be processed separately and the results of prior processing could be imported and a method of automated tagging and matching of the tumors could be implemented, such interoperability mechanisms are not defined or mandated here and cannot be depended on to be present or used.

Storing segmentations and measurement results that can be loaded by an Image Analysis Tool analyzing data collected at a later date is certainly a useful practice as it can save time and cost.  For this to happen reliably, the stored format must be compatible and the data must be stored and conveyed.  Although DICOM Segmentation objects are appropriate to store tumor segmentations, and DICOM SR objects are appropriate to store measurement results, these standards are not yet widely enough deployed to make support for them mandatory in this Profile.  Similarly, conveying the segmentations and measurements from baseline (and other time points prior to the current exam) is not done consistently enough to mandate that it happen and to require their import into the Image Analysis Tool.  Managing and forwarding the data files may exceed the practical capabilities of the participating sites.

Image analysis can be performed on any equipment that complies with the specifications set out in this Profile.  However, we strongly encourage performing all analysis for an individual subject on the same platform (manufacturer, model and version) which we expect will further reduce variation.

Medical Devices such as the Image Analysis Tool are typically made up of multiple components (the hardware, the operating system, the application software, and various function libraries within those).  Changes in any of the components can affect the behavior of the device.  In this specification, the “device version” should reflect the total set of components and any changes to components should result in a change in the recorded device version.  This device version may thus be different than the product release version that appears in vendor documentation.

For analysis methods that involve an operator (e.g. to draw or edit boundaries, set seed points or adjust parameters), the operator is effectively a component of the system, with an impact on the reproducibility of the measurements, and it is important to record the operator’s identify as well.  Fully automated analysis software removes that source of variation; although even then, since a human is generally responsible for the final results, they retain the power to approve or reject measurements so their identity should be recorded.

The Tumor Volume Change performance specification below includes the operator performance and is intended to be evaluated based on a typical operator (i.e. without extraordinary training or ability).  This should be kept in mind by vendors measuring the performance of their tools and sites validating the performance of their installation.  Although the performance of some methods may depend on the judgment and skill of the operator, it is beyond this Profile to specify the qualifications or experience of the operator.

Determination of which tumors should be measured is out of scope for this Profile.  Such determination may be specified within a protocol or specified by formal response criteria standards, or may be determined by clinical requirements. Tumors to be measured may be designated by the oncologist or clinical investigator, by a radiologist at a clinical site, by a reader at a central reading facility, or they may be designated automatically by a software analysis tool.

3.4.2 Specification

| Parameter | Actor | Specification |
| --- | --- | --- |
| Common Tumor Selection | Image Analysis Tool | Shall allow all tumors selected for volume measurement at a given time point to be unambiguously labeled. The claim will only be met if the tumor measured at a later time point is the same as that measured at the earlier time points. |
| Multiple Tumors | Image Analysis Tool | Shall allow multiple tumors to be measured, and each measured tumor to be associated with a human-readable identifier that can be used for correlation across time points. |
| Tumor Volume Change Variability | Image Analysis Tool | The following two specifications are essentially the same, with the first applying to the provider of the tool and the second applying to the site where the tool is used. Shall demonstrate the ability to measure tumor volume change (according to Figure 1) on data that meets the criteria of the preceding activities with performance assed using same metrics as identified in overall system Claim at levels of two-thirds of that claimed for the entire system. See 4.3. Assessment Procedure: Tumor Volume Change Variability (Image Analysis Tool). |
|  | Radiologist | Shall (if operator interaction is required by the Image Analysis Tool to perform measurements) demonstrate the ability to measure tumor volume change (according to Figure 1) on data that meets the criteria of the preceding activities with a 95% confidence interval around the measured change of no greater than +/- 30%. |
| Result Verification | Radiologist | Shall review/approve the measurement results as needed. |
| Recording | Image Analysis Tool | Shall record the percentage volume change relative to baseline for each tumor, the device version and the actual model-specific Analysis Software set-up and configuration parameters utilized. Shall be capable of recording the tumor segmentation as a DICOM Segmentation. Shall record the identity of each individual making and/or approving a tumor measurement using the software. |

# 4. Compliance Procedures

To comply with this Profile, participating staff and equipment (“Actors”) shall support each of the activities assigned to them in Table 1.

For each activity, the compliance requirements checklist (sometimes referred to as the “shall language”) for each Actor are documented in Section 3.

Although most of the requirements described in Section 3 are feature-oriented and/or compliance can be assessed by direct observation, some of the requirements are performance-oriented.  The following sub-sections elaborate on the meaning of performance-oriented requirements and how they are intended to be correctly assessed.

Formal claims of compliance by the organization responsible for an Actor shall be in the form of a published QIBA Conformance Statement.  Vendors publishing a QIBA Conformance Statement shall provide a set of “Model-specific Parameters” (as shown in Appendix D) describing how their product was configured to achieve compliance.  Vendors shall also provide access or describe the characteristics of the test set used for compliance testing.

## 4.1. Assessment Procedure: In-plane Spatial Resolution

This procedure can be used by a vendor or an imaging site to assess the In-plane Spatial Resolution of reconstructed images.  It is assumed that the images have been acquired and reconstructed in compliance with Section 3.2.2 and 3.3.2.

This procedure is provided as a reference method.  Sites or vendors wishing to use other methods must first submit to QIBA evidence that the results produced by their proposed method are equivalent to this reference method.

<< allow use of other phantoms if “demonstrated to provide equivalent results to ACR phantom” e.g. manufacturers might have phantoms (and vendor or key site could do the equivalence demonstration work).  An MTF result might be shown to produce an equivalent lp/cm result.  So what is the %MTF (maybe 10%) that correlates to 6lp/cm.>>

Mike can find references in ACR manuals or IEC doc.

Should we have two methods? One for vendor/factory and one for site/spotcheck?

## 4.2. Assessment Procedure: Voxel Noise

This procedure can be used by a vendor or an imaging site to assess the Voxel Noise of reconstructed images.  It is assumed that the images have been acquired and reconstructed in compliance with Section 3.2.2 and 3.3.2.

<Describe the use of the water phantom, an ROI and the standard deviation noise metric>>

## 4.3. Assessment Procedure:

4.1.1 Acquisition Device

This procedure can be used by a vendor or an imaging site to select an appropriate reconstruction kernel??? for an acquisition device model:

- Set the scanning field of view for the patient, the Kyoto Kagaku chest phantom. This setting is to be used to image the ACR phantom. Special handling: In scanning the ACR CT phantom, some manufacturers specify the use of a FOV appropriate to the ACR device (See: http://www.acr.org/accreditation/computed/ct_faq.aspx#thirteen). In this case, follow the manufacturer’s guidance for the ACR phantom. As an example, for the Aquilion16 scanner see the guidance in slide 11 of

http://www.tams-media.com/tams2008/sales/faqs/CT_pdfs/ACR_Guide_Aquilion16.pdf

- Set the beam voltage to 120 kVp

- Set the slice thickness to between 0.75 and 1.25 mm (depending on the available reconstructed slice thicknesses of the scanner)

- Set nominal beam collimation (NxT such as 16 x 0.5mm,  or    128 x 0.6mm, 320 x 0.5 mm) rotation time and pitch such that scan can cover a 35 cm thorax in 15 seconds or less

- ITERATE (hopefully only a few times) on reconstruction kernels to meet spatial resolution spec.

Figure 2: Establishing spatial resolution

- ITERATE (again, hopefully just a few times) on mAs or effective mAs setting, given beam collimation, pitch and rotation time.

Figure 3: Establishing noise spec

- If the scanning FOV is to be changed for the scan of the lung phantom, reset the FOV accordingly and rescan the ACR phantom. Measure the quality parameters, the noise and resolution, with the changed settings.

The quality parameters are expected to change under a changed scanning field of view, as in the special handling. If this is the case, the comparability of the quality of the various scanners is lost. In the analysis of the quantitative measurements of nodule sizes, the data on the actual quality measures may prove to be useful in analyzing device differences.

4.1.2 Technologist

Checklist taken from union of section 3 activities with a requirement indicated for Technologist PLUS UPICT requirements for activities not covered in the Profile

4.1.3 Radiologist

Checklist taken from union of section 3 activities with a requirement indicated for Radiologist PLUS UPICT requirements for activities not covered in the Profile

## 4.4. Assessment Procedure: Tumor Volume Change Variability (Image Analysis Tool)

Note: The procedure in this section is currently only a proposal.

A more detailed procedure and pointers to valid test datasets will be provided in the future.

Until then, there is no approved way to claim conformance to this performance requirement.

Tumor Volume Change Variability performance can be assessed with the following procedure:

- Obtain a designated test image set (see 4.4.1).

- Determine the volume and/or the volume change as appropriate for designated tumors (see 4.4.2).

- Calculate statistical metrics of performance (see 4.4.3).

- Compare against the Tumor Volume Change Variability performance level specified in 3.4.2.

Results shall be reported by a vendor or an imaging site on the performance of an Image Analysis Tool both with and without editing of results by a typical operator.  Note that tumor detection and seeding is unconstrained by this Profile.[this may need further explanation]

### 4.4.1. Test Image Set

The test image set consists of multiple target tumors in multiple body parts in multiple subjects (human or phantom).  The body parts are representative of the stated scope of the Profile (i.e. includes lung nodules as well as metastases such as mediastinal, liver, adrenal, neck, axillary, mesenteric, retroperitioneal, pelvic, etc. described in Appendix B.3).

The target tumors are selected to be measureable (i.e. larger than 10mm diameter [if this is based on RECIST, then RECIST should be referenced], geometrically simple and with sufficiently conspicuous margins) and have a range of volumes, shapes and types to be representative of the scope of the Profile [this is unclear].

The test image set includes at least N target tumors.  Each target tumor has at least T time points.  The tumors span B body parts covered by the scope of the Profile..

The test image set has been acquired according to the requirements of this Profile (e.g. patient handling, acquisition protocol, reconstruction).

Discussion:

We have many test image cases where the true change is known to be 0% (“Coffee break”).

We have many test image cases where the true change is unknown (although change is clearly present).

Are we missing data to show both sensitivity and specificity?[not clear what how this would be specified.  Pixel-based assessment?]

What exactly is our goal with this performance assessment?

Consider a multi- step assessment?

1) Assess (change?) sensitivity (in terms of inherent measurement variation) using “No change” data

2) Assess (volume?) bias using data with a known volume (phantom?)

3) Assess change performance against consensus values (rather than measured/known truth?)

Tumor segmentation performance can be affected by the accuracy or variations in the seed point or axis provided.  Consider preparing the test set with test “inputs” (either with a “click here” dot on the image, or some method for feeding coordinates to the application).[this sort of makes sense, but not completely clear how this would work]

Ideally we want fully realistic images (not phantom) but with known truth for tumor volume change.  Would it be possible to digitally insert tumors into real acquired human images?  [such projects are under consideration]

What is the best way to go about assembling and hosting these datasets?  Such a public dataset is not currently known to exist.  [LIDC and NBIA are examples of such datasets and hosting paradigms].

### 4.4.2 Calculate statistical metrics of performance

The following subsections present the results of the analyses described above:

- Scalar volume measurements

  - Bias and linearity:

  - Variability:

    - Descriptive statistics and transformation.

    - Repeatability across test-retest repetitions within participants, using numeric calculation and Bland-Altman analysis to report UAL, LAL, wSD, RC, wCV, and CCC.

    - Reproducibility:

      - Consider heteroscedasticity to determine if characteristics of this algorithm conform to the set of previously pooled algorithms;

      - Analysis including all tumors / all algorithms:

      - Numeric calculation and Bland-Altman analysis to report UAL, LAL, and RDC.

      - LME model to assess subject, algorithm, and residual variance.

      - Another that partitions tumors based on QIBA Profile “measurability”;[not clear]

      - Another that partitions algorithms according to degree of post-editing.[not clear]

- Comparison of tumor segmentation boundaries:

  - Update or create reference segmentation for each patient and each test-retest repetition.[this would be expensive and highly variable]

  - Calculation of sensitivity, Jaccard index, and Sørensen–Dice coefficients.

  - Merging and plotting of histograms and radial chart.

## 4.5. Performance Assessment: Image Acquisition Site

Note: The procedure in this section is currently only a proposal.

A more detailed procedure and pointers to valid test datasets will be provided in the future.

Until then, there is no approved way to claim conformance to this performance requirement.

Site performance can be assessed with the following procedure:

- Validate image acquisition (see 4.2.1).

- Generate a test image set (see 4.2.2).

- Assess Tumor Volume Change Variability (see 4.1.2, 4.1.3 above).

- Compare against the Tumor Volume Change Variability performance level specified in 3.4.2.

This procedure can be used by an imaging site to evaluate the performance of each of the Actors and Activities in use.  In principle, the final result represents an assessment of the combined performance of all the Actors and Activities at the site.

The procedure presumes that the Actors being used by the site are capable of meeting the requirements described in Section 3 of this document; however it is not a pre-requisite that those Actors have published QIBA Conformance Statements (although that would be both useful and encouraging).

Discussion:

Duke is working on a “platform” that includes a phantom and an analysis tool that may inform the future contents of this section.

Sites that carry out this procedure should really record the parameters they used and document them in something similar to a Conformance Statement.  This would be a useful QA record and could be submitted to clinical trials looking for QIBA compliant test sites.

Are there other criteria that should be worked into this procedure?

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large patient population under consideration.  For imaging it is important to consider the availability of:

- appropriate imaging equipment and quality control processes,

- appropriate injector equipment and contrast media,

- experienced CT Technologists for the imaging procedure, and

- processes that assure imaging Profile compliant image generation at the correct point in time.

A clinical trial might specify “A calibration and QA program shall be designed consistent with the goals of the clinical trial. This program shall include (a) elements to verify that sites are performing correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address issues relating to both radiation dose and image quality (which may include issues relating to water calibration, uniformity, noise, spatial resolution -in the axial plane-, reconstructed slice thickness z-axis resolution, contrast scale, CT number calibration and others). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.”

4.2.1 Acquisition Validation

Review patient handling procedures for compliance with Section 3.1

Establish acquisition protocols and reconstruction settings on the Acquisition Device compliant with Section 3.2 and Section 3.3.  If a QIBA Conformance Statement is available from the Acquisition Device vendor, it may provide parameters useful for this step.

Acquire images of a 20cm water phantom, reconstruct and confirm performance requirements in Section 3.3.2 are met.

Discussion:

UCLA may have more detailed and more complete procedures to recommend for this section.

4.2.2 Test Image Set

Locally acquire a test image set using the protocols established and tested in Section 4.2.1.

The test image set should conform to the characteristics described in Section 4.1.1.

Discussion:

It is highly likely that due to practical constraints the test image set prepared at an individual site would be much less comprehensive than the test image sets prepared by QIBA. Further consideration of what a more limited but still useful test image set would look like.

# References

1.	Wang, Y., R.J. van Klaveren, H.J. van der Zaag-Loonen, G.H. de Bock, et al., Effect of nodule characteristics on variability of semiautomated volume measurements in pulmonary nodules detected in a lung cancer screening program. Radiology, 2008. 248(2): p. 625-31.

2.	Buckler, A.J., P.D. Mozley, L. Schwartz, N. Petrick, et al., Volumetric CT in lung cancer: an example for the qualification of imaging as a biomarker. Academic radiology, 2010. 17(1): p. 107-15.

3.	Mozley, P.D., L.H. Schwartz, C. Bendtsen, B. Zhao, N. Petrick, and A.J. Buckler, Change in lung tumor volume as a biomarker of treatment response: a critical review of the evidence. Ann Oncol, 2010. 21(9): p. 1751-5.

4.	Buckler, A.J., A procedural template for the qualification of imaging as a biomarker, using volumetric CT as an example, in IEEE Applied Imagery Pattern Recognition Workshop2009: Cosmos Club, Washington, D.C. p. 7.

5.	Buckler, A.J., J.L. Mulshine, R. Gottlieb, B. Zhao, P.D. Mozley, and L. Schwartz, The use of volumetric CT as an imaging biomarker in lung cancer. Academic radiology, 2010. 17(1): p. 100-6.

6.	Buckler AJ, S.L., Petrick N, McNitt-Gray M, Zhao B, Fenimore C, Reeves AP, Mozley PD, Avila RS, Data Sets for the Qualification of CT as a Quantitative Imaging Biomarker in Lung Cancer. Optics express, 2010. 18(14): p. 16.

7.	Shankar, L.K., A. Van den Abbeele, J. Yap, R. Benjamin, S. Scheutze, and T.J. Fitzgerald, Considerations for the use of imaging tools for phase II treatment trials in oncology. Clinical cancer research : an official journal of the American Association for Cancer Research, 2009. 15(6): p. 1891-7.

8.	Zhao, B., L.H. Schwartz, and S.M. Larson, Imaging surrogates of tumor response to therapy: anatomic and functional biomarkers. J Nucl Med, 2009. 50(2): p. 239-49.

9.	Maitland, M.L., Volumes to learn: advancing therapeutics with innovative computed tomography image data analysis. Clin Cancer Res, 2010. 16(18): p. 4493-5.

10.	Nishino, M., D.M. Jackman, H. Hatabu, P.A. Janne, B.E. Johnson, and A.D. Van den Abbeele, Imaging of lung cancer in the era of molecular medicine. Acad Radiol, 2011. 18(4): p. 424-36.

11.	Koshariya, M., R.B. Jagad, J. Kawamoto, P. Papastratis, H. Kefalourous, T. Porfiris, C. Tzouma, and N.J. Lygidakis, An update and our experience with metastatic liver disease. Hepatogastroenterology, 2007. 54(80): p. 2232-9.

12.	Jaffe, C.C., Measures of response: RECIST, WHO, and new alternatives. J Clin Oncol, 2006. 24(20): p. 3245-51.

13.	Gavrielides, M.A., L.M. Kinnard, K.J. Myers, and N. Petrick, Noncalcified lung nodules: volumetric assessment with thoracic CT. Radiology, 2009. 251(1): p. 26-37.

14.	Petrick, N., H.J.G. Kim, D. Clunie, K. Borradaile, et al., Evaluation of 1D, 2D and 3D nodule size estimation by radiologists for spherical and non-spherical nodules through CT thoracic phantom imaging, in SPIE 2011.

15.	Kinnard, L.M., M.A. Gavrielides, K.J. Myers, R. Zeng, J. Peregoy, W. Pritchard, J.W. Karanian, and N. Petrick, Volume error analysis for lung nodules attached to pulmonary vessels in an anthropomorphic thoracic phantom. Proc SPIE, 2008. 6915: p. 69152Q; doi:10.1117/12.773039.

16.	Gavrielides, M.A., R. Zeng, L.M. Kinnard, K.J. Myers, and N. Petrick, A template-based approach for the analysis of lung nodules in a volumetric CT phantom study. Proc SPIE, 2009. 7260: p. 726009; doi:10.1117/12.813560.

17.	Winer-Muram, H.T., S.G. Jennings, C.A. Meyer, Y. Liang, A.M. Aisen, R.D. Tarver, and R.C. McGarry, Effect of varying CT section width on volumetric measurement of lung tumors and application of compensatory equations. Radiology, 2003. 229(1): p. 184-94.

18.	Ravenel, J.G., W.M. Leue, P.J. Nietert, J.V. Miller, K.K. Taylor, and G.A. Silvestri, Pulmonary nodule volume: effects of reconstruction parameters on automated measurements--a phantom study. Radiology, 2008. 247(2): p. 400-8.

19.	Borradaile, K. and R. Ford, Discordance between BICR readers. Appl Clin Trials, 2010. Nov 1: p. Epub.

20.	Gavrielides, M.A., R. Zeng, K.J. Myers, B. Sahiner, and N. Petrick, Benefit of Overlapping Reconstruction for Improving the Quantitative Assessment of CT Lung Nodule Volume. Acad Radiol, 2012.

21.	Gavrielides, M.A., R. Zeng, L.M. Kinnard, K.J. Myers, and N. Petrick, Information-theoretic approach for analyzing bias and variance in lung nodule size estimation with CT: a phantom study. IEEE Trans Med Imaging, 2010. 29(10): p. 1795-807.

22.	Gavrielides, M.A., L.M. Kinnard, K.J. Myers, J. Peregoy, W.F. Pritchard, R. Zeng, J. Esparza, J. Karanian, and N. Petrick, A resource for the assessment of lung nodule size estimation methods: database of thoracic CT scans of an anthropomorphic phantom. Opt Express, 2010. 18(14): p. 15244-55.

23.	Das, M., J. Ley-Zaporozhan, H.A. Gietema, A. Czech, et al., Accuracy of automated volumetry of pulmonary nodules across different multislice CT scanners. Eur Radiol, 2007. 17(8): p. 1979-84.

24.	Bolte, H., C. Riedel, S. Muller-Hulsbeck, S. Freitag-Wolf, G. Kohl, T. Drews, M. Heller, and J. Biederer, Precision of computer-aided volumetry of artificial small solid pulmonary nodules in ex vivo porcine lungs. Br J Radiol, 2007. 80(954): p. 414-21.

25.	Cagnon, C.H., D.D. Cody, M.F. McNitt-Gray, J.A. Seibert, P.F. Judy, and D.R. Aberle, Description and implementation of a quality control program in an imaging-based clinical trial. Acad Radiol, 2006. 13(11): p. 1431-41.

26.	Goodsitt, M.M., H.P. Chan, T.W. Way, S.C. Larson, E.G. Christodoulou, and J. Kim, Accuracy of the CT numbers of simulated lung nodules imaged with multi-detector CT scanners. Med Phys, 2006. 33(8): p. 3006-17.

27.	Oda, S., K. Awai, K. Murao, A. Ozawa, Y. Yanaga, K. Kawanaka, and Y. Yamashita, Computer-aided volumetry of pulmonary nodules exhibiting ground-glass opacity at MDCT. AJR Am J Roentgenol, 2010. 194(2): p. 398-406.

28.	McNitt-Gray, M.F., L.M. Bidaut, S.G. Armato, C.R. Meyer, et al., Computed tomography assessment of response to therapy: tumor volume change measurement, truth data, and error. Transl Oncol, 2009. 2(4): p. 216-22.

29.	Keil, S., C. Plumhans, F.F. Behrendt, S. Stanzel, M. Suehling, G. Muhlenbruch, A.H. Mahnken, R.W. Gunther, and M. Das, Semi-automated quantification of hepatic lesions in a phantom. Invest Radiol, 2009. 44(2): p. 82-8.

30.	Guidance for Industry: Standards for Clinical Trial Imaging Endpoints, 2011, U.S. Department of Health and Human Services, Food and Drug Administration.

31.	Bland, J.M. and D.G. Altman, Statistical methods for assessing agreement between two methods of clinical measurement. Lancet, 1986. 1(8476): p. 307-10.

32.	Lin, L.I., A concordance correlation coefficient to evaluate reproducibility. Biometrics, 1989. 45(1): p. 255-68.

33.	Moertel, C.G. and J.A. Hanley, The effect of measuring error on the results of therapeutic trials in advanced cancer. Cancer, 1976. 38(1): p. 388-94.

34.	Lavin, P.T. and G. Flowerdew, Studies in variation associated with the measurement of solid tumors. Cancer, 1980. 46(5): p. 1286-90.

35.	Eisenhauer, E.A., P. Therasse, J. Bogaerts, L.H. Schwartz, et al., New response evaluation criteria in solid tumours: revised RECIST guideline (version 1.1). Eur J Cancer, 2009. 45(2): p. 228-47.

36.	Zhao, B., L.H. Schwartz, C.S. Moskowitz, M.S. Ginsberg, N.A. Rizvi, and M.G. Kris, Lung cancer: computerized quantification of tumor response--initial results. Radiology, 2006. 241(3): p. 892-8.

37.	Zhao, B., G.R. Oxnard, C.S. Moskowitz, M.G. Kris, et al., A pilot study of volume measurement as a method of tumor response evaluation to aid biomarker development. Clin Cancer Res, 2010. 16(18): p. 4647-53.

38.	Schwartz, L.H., S. Curran, R. Trocola, J. Randazzo, D. Ilson, D. Kelsen, and M. Shah, Volumetric 3D CT analysis - an early predictor of response to therapy. J Clin Oncol, 2007. 25(18s): p. abstr 4576.

39.	Altorki, N., M.E. Lane, T. Bauer, P.C. Lee, et al., Phase II proof-of-concept study of pazopanib monotherapy in treatment-naive patients with stage I/II resectable non-small-cell lung cancer. J Clin Oncol, 2010. 28(19): p. 3131-7.

40.	Gietema, H.A., C.M. Schaefer-Prokop, W.P. Mali, G. Groenewegen, and M. Prokop, Pulmonary nodules: Interscan variability of semiautomated volume measurements with multisection CT-- influence of inspiration level, nodule size, and segmentation performance. Radiology, 2007. 245(3): p. 888-94.

41.	Wormanns, D., G. Kohl, E. Klotz, A. Marheine, F. Beyer, W. Heindel, and S. Diederich, Volumetric measurements of pulmonary nodules at multi-row detector CT: in vivo reproducibility. Eur Radiol, 2004. 14(1): p. 86-92.

42.	Boll, D.T., R.C. Gilkeson, T.R. Fleiter, K.A. Blackham, J.L. Duerk, and J.S. Lewin, Volumetric assessment of pulmonary nodules with ECG-gated MDCT. AJR Am J Roentgenol, 2004. 183(5): p. 1217-23.

43.	Hein, P.A., V.C. Romano, P. Rogalla, C. Klessen, A. Lembcke, V. Dicken, L. Bornemann, and H.C. Bauknecht, Linear and volume measurements of pulmonary nodules at different CT dose levels - intrascan and interscan analysis. Rofo, 2009. 181(1): p. 24-31.

44.	Meyer, C.R., T.D. Johnson, G. McLennan, D.R. Aberle, et al., Evaluation of lung MDCT nodule annotation across radiologists and methods. Acad Radiol, 2006. 13(10): p. 1254-65.

45.	Marten, K., F. Auer, S. Schmidt, G. Kohl, E.J. Rummeny, and C. Engelke, Inadequacy of manual measurements compared to automated CT volumetry in assessment of treatment response of pulmonary metastases using RECIST criteria. Eur Radiol, 2006. 16(4): p. 781-90.

46.	Revel, M.P., C. Lefort, A. Bissery, M. Bienvenu, L. Aycard, G. Chatellier, and G. Frija, Pulmonary nodules: preliminary experience with three-dimensional evaluation. Radiology, 2004. 231(2): p. 459-66.

47.	Sohns, C., J. Mangelsdorf, S. Sossalla, F. Konietschke, and S. Obenauer, Measurement of response of pulmonal tumors in 64-slice MDCT. Acta Radiol, 2010. 51(5): p. 512-21.

48.	Werner-Wasik, M., Y. Xiao, E. Pequignot, W.J. Curran, and W. Hauck, Assessment of lung cancer response after nonoperative therapy: tumor diameter, bidimensional product, and volume. A serial CT scan-based study. Int J Radiat Oncol Biol Phys, 2001. 51(1): p. 56-61.

49.	Tran, L.N., M.S. Brown, J.G. Goldin, X. Yan, R.C. Pais, M.F. McNitt-Gray, D. Gjertson, S.R. Rogers, and D.R. Aberle, Comparison of treatment response classifications between unidimensional, bidimensional, and volumetric measurements of metastatic lung lesions on chest computed tomography. Acad Radiol, 2004. 11(12): p. 1355-60.

50.	Jennings, P., S. Aydin, J. Bennett, R. McBride, et al., Inter-laboratory comparison of human renal proximal tubule (HK-2) transcriptome alterations due to Cyclosporine A exposure and medium exhaustion. Toxicol In Vitro, 2009. 23(3): p. 486-99.

51.	Prasad, S.R., K.S. Jhaveri, S. Saini, P.F. Hahn, E.F. Halpern, and J.E. Sumner, CT tumor measurement for therapeutic response assessment: comparison of unidimensional, bidimensional, and volumetric techniques initial observations. Radiology, 2002. 225(2): p. 416-9.

52.	Rohde, S., A.F. Kovacs, J. Berkefeld, and B. Turowski, Reliability of CT-based tumor volumetry after intraarterial chemotherapy in patients with small carcinoma of the oral cavity and the oropharynx. Neuroradiology, 2006. 48(6): p. 415-21.

53.	Lee, S.M., S.H. Kim, J.M. Lee, S.A. Im, et al., Usefulness of CT volumetry for primary gastric lesions in predicting pathologic response to neoadjuvant chemotherapy in advanced gastric cancer. Abdom Imaging, 2009. 34(4): p. 430-40.

54.	Beer, A.J., H.A. Wieder, F. Lordick, K. Ott, M. Fischer, K. Becker, J. Stollfuss, and E.J. Rummeny, Adenocarcinomas of esophagogastric junction: multi-detector row CT to evaluate early response to neoadjuvant chemotherapy. Radiology, 2006. 239(2): p. 472-80.

55.	Griffith, J.F., A.C. Chan, L.T. Chow, S.F. Leung, Y.H. Lam, E.Y. Liang, S.C. Chung, and C. Metreweli, Assessing chemotherapy response of squamous cell oesophageal carcinoma with spiral CT. Br J Radiol, 1999. 72(859): p. 678-84.

56.	Benz, M.R., M.S. Allen-Auerbach, F.C. Eilber, H.J. Chen, S. Dry, M.E. Phelps, J. Czernin, and W.A. Weber, Combined assessment of metabolic and volumetric changes for assessment of tumor response in patients with soft-tissue sarcomas. J Nucl Med, 2008. 49(10): p. 1579-84.

57.	Willett, C.G., M.A. Stracher, R.M. Linggood, L.M. Miketic, J.C. Leong, S.J. Skates, D.C. Kushner, and J.O. Jacobson, Three-dimensional volumetric assessment of response to treatment: stage I and II diffuse large cell lymphoma of the mediastinum. Radiother Oncol, 1988. 12(3): p. 193-8.

58.	Willett, C.G., R.M. Linggood, J.C. Leong, L.M. Miketic, M.A. Stracher, S.J. Skates, and D.C. Kushner, Stage IA to IIB mediastinal Hodgkin's disease: three-dimensional volumetric assessment of response to treatment. J Clin Oncol, 1988. 6(5): p. 819-24.

59.	Wakelee, H.A., P. Bernardo, D.H. Johnson, and J.H. Schiller, Changes in the natural history of nonsmall cell lung cancer (NSCLC)--comparison of outcomes and characteristics in patients with advanced NSCLC entered in Eastern Cooperative Oncology Group trials before and after 1990. Cancer, 2006. 106(10): p. 2208-17.

60.	Gandara, D.R., D. Aberle, D. Lau, J. Jett, T. Akhurst, R. Heelan, J. Mulshine, C. Berg, and E.F. Patz, Jr., Radiographic imaging of bronchioloalveolar carcinoma: screening, patterns of presentation and response assessment. J Thorac Oncol, 2006. 1(9 Suppl): p. S20-6.

61.	Goldstraw, P., J. Crowley, K. Chansky, D.J. Giroux, P.A. Groome, R. Rami-Porta, P.E. Postmus, V. Rusch, and L. Sobin, The IASLC Lung Cancer Staging Project: proposals for the revision of the TNM stage groupings in the forthcoming (seventh) edition of the TNM Classification of malignant tumours. J Thorac Oncol, 2007. 2(8): p. 706-14.

62.	Jemal, A., R. Siegel, E. Ward, Y. Hao, J. Xu, T. Murray, and M.J. Thun, Cancer statistics, 2008. CA Cancer J Clin, 2008. 58(2): p. 71-96.

63.	Zhao, B., G.R. Oxnard, C.S. Moskowitz, M.G. Kris, et al., A Pilot Study of Volume Measurement as a Method of Tumor Response Evaluation to Aid Biomarker Development. Clin Cancer Res, 2010. 16: p. 4647-4653.

64.	Fraioli, F., L. Bertoletti, A. Napoli, F.A. Calabrese, R. Masciangelo, E. Cortesi, C. Catalano, and R. Passariello, Volumetric evaluation of therapy response in patients with lung metastases. Preliminary results with a computer system (CAD) and comparison with unidimensional measurements. Radiol Med, 2006. 111(3): p. 365-75.

65.	Parkin, D.M., F. Bray, J. Ferlay, and P. Pisani, Global cancer statistics, 2002. CA Cancer J Clin, 2005. 55(2): p. 74-108.

66.	Llovet, J.M. and J. Bruix, Systematic review of randomized trials for unresectable hepatocellular carcinoma: Chemoembolization improves survival. Hepatology, 2003. 37(2): p. 429-42.

67.	Lopez, P.M., A. Villanueva, and J.M. Llovet, Systematic review: evidence-based management of hepatocellular carcinoma--an updated analysis of randomized controlled trials. Aliment Pharmacol Ther, 2006. 23(11): p. 1535-47.

68.	Keil, S., F.F. Behrendt, S. Stanzel, M. Suhling, et al., Semi-automated measurement of hyperdense, hypodense and heterogeneous hepatic metastasis on standard MDCT slices. Comparison of semi-automated and manual measurement of RECIST and WHO criteria. Eur Radiol, 2008. 18(11): p. 2456-65.

69.	Therasse, P., S.G. Arbuck, E.A. Eisenhauer, J. Wanders, et al., New guidelines to evaluate the response to treatment in solid tumors. European Organization for Research and Treatment of Cancer, National Cancer Institute of the United States, National Cancer Institute of Canada. J Natl Cancer Inst, 2000. 92(3): p. 205-16.

70.	Llovet, J.M., S. Ricci, V. Mazzaferro, P. Hilgard, et al., Sorafenib in advanced hepatocellular carcinoma. N Engl J Med, 2008. 359(4): p. 378-90.

71.	Cheng, A.L., Y.K. Kang, Z. Chen, C.J. Tsao, et al., Efficacy and safety of sorafenib in patients in the Asia-Pacific region with advanced hepatocellular carcinoma: a phase III randomised, double-blind, placebo-controlled trial. Lancet Oncol, 2009. 10(1): p. 25-34.

72.	Garant, M., M. Trudeau, C. Reinhold, and P.M. Bret, Liver metastasis: comparison of 2 methods for reporting of disease in patients receiving chemotherapy. Can Assoc Radiol J, 1999. 50(1): p. 13-6.

73.	Stillwagon, G.B., S.E. Order, C. Guse, J.L. Klein, P.K. Leichner, S.A. Leibel, and E.K. Fishman, 194 hepatocellular cancers treated by radiation and chemotherapy combinations: toxicity and response: a Radiation Therapy Oncology Group Study. Int J Radiat Oncol Biol Phys, 1989. 17(6): p. 1223-9.

74.	Sohaib, S.A., B. Turner, J.A. Hanson, M. Farquharson, R.T. Oliver, and R.H. Reznek, CT assessment of tumour response to treatment: comparison of linear, cross-sectional and volumetric measures of tumour size. Br J Radiol, 2000. 73(875): p. 1178-84.

75.	Luccichenti, G., F. Cademartiri, M. Sianesi, L. Roncoroni, P. Pavone, and G.P. Krestin, Radiologic assessment of rectosigmoid cancer before and after neoadjuvant radiation therapy: comparison between quantitation techniques. AJR Am J Roentgenol, 2005. 184(2): p. 526-30.

76.	Rohde, S., B. Turowski, J. Berkefeld, and A.F. Kovacs, CT-based evaluation of tumor volume after intra-arterial chemotherapy of locally advanced carcinoma of the oral cavity: comparison with clinical remission rates. Cardiovasc Intervent Radiol, 2007. 30(1): p. 85-91.

77.	Therasse, P., Measuring the clinical response. What does it mean? Eur J Cancer, 2002. 38(14): p. 1817-23.

78.	McHugh, K. and S. Kao, Response evaluation criteria in solid tumours (RECIST): problems and need for modifications in paediatric oncology? Br J Radiol, 2003. 76(907): p. 433-6.

# Appendices

## Appendix A: Acknowledgements and Attributions

This document is proffered by the Radiological Society of North America (RSNA) Quantitative Imaging Biomarker Alliance (QIBA) Volumetric Computed Tomography (v-CT) Technical Committee. The v-CT technical committee is composed of scientists representing the imaging device manufacturers, image analysis software developers, image analysis laboratories, biopharmaceutical industry, academia, government research organizations, professional societies, and regulatory agencies, among others. All work is classified as pre-competitive.

A more detailed description of the v-CT group and its work can be found at the following web link: http://qibawiki.rsna.org/index.php?title=Volumetric_CT.

The Volumetric CT Technical Committee (in alphabetical order):

- Athelogou, M. Definiens AG

- Avila, R. Kitware, Inc.

- Beaumont, H. Median Technologies

- Borradaile, K. Core Lab Partners

- Buckler, A. BBMSC

- Clunie, D. Core Lab Partners

- Cole, P. Imagepace

- Conklin, J. ICON Medical Imaging

- Dorfman, GS. Weill Cornell Medical College

- Fenimore, C. Nat Inst Standards & Technology

- Ford, R. Princeton Radiology Associates.

- Garg, K. University of Colorado

- Garrett, P. Smith Consulting, LLC

- Goldmacher, G. ICON Medical Imaging

- Gottlieb, R. University of Arizona

- Gustafson, D. Intio, Inc.

- Hayes, W. Bristol Myers Squibb

- Hillman, B. Metrix, Inc.

- Judy, P. Brigham and Women’s Hospital

- Kim, HJ. University of California Los Angeles

- Kohl, G. Siemens AG

- Lehner, O. Definiens AG

- Lu, J. Nat Inst Standards & Technology

- McNitt-Gray, M. University California Los Angeles

- Mozley, PD. Merck & Co Inc.

- Mulshine, JL. Rush University

- Nicholson, D. Definiens AG

- O'Donnell, K. Toshiba Medical Research Institute - USA

- O'Neal, M. Core Lab Partners

- Petrick, N. US Food and Drug Administration

- Reeves, A. Cornell University

- Richard, S. Duke University

- Rong, Y. Perceptive Informatics, Inc.

- Schwartz, LH. Columbia University

- Saiprasad, G. University of Maryland

- Samei, E. Duke University

- Siegel, E. University of Maryland

- Silver, M. Toshiba Medical Research Institute – USA

- Steinmetz, N. Translational Sciences Corporation

- Sullivan, DC. RSNA Science Advisor and Duke University

- Tang, Y. CCS Associates

- Thorn, M. Siemens AG

- Vining, DJ. MD Anderson Cancer Center

- Yankellivitz, D. Mt. Sinai School of Medicine

- Yoshida, H. Harvard MGH

- Zhao, B. Columbia University

The Volumetric CT Technical Committee is deeply grateful for the support and technical assistance provided by the staff of the Radiological Society of North America.

## Appendix B: Background Information

## B.1 QIBA

The Quantitative Imaging Biomarker Alliance (QIBA) is an initiative to promote the use of standards to reduce variability and improve performance of quantitative imaging in medicine. QIBA provides a forum for volunteer committees of care providers, medical physicists, imaging innovators in the device and software industry, pharmaceutical companies, and other stakeholders in several clinical and operational domains to reach consensus on standards-based solutions to critical quantification issues. QIBA publishes the specifications they produce (called QIBA Profiles), first to gather public comment and then for field test by vendors and users.

QIBA envisions providing a process for developers to test their implementations of QIBA Profiles through a compliance mechanism. Purchasers can specify conformance with appropriate QIBA Profiles as a requirement in Requests For Proposals (RFPs). Vendors who have successfully implemented QIBA Profiles in their products can publish QIBA Conformance Statements.  The Conformance Statements are accompanied by “Model-specific Parameters” (as shown in Appendix D) describing how to configure their product for alignment with the Profile.

General information about QIBA, including its governance structure, sponsorship, member organizations and work process, is available at http://qibawiki.rsna.org/index.php?title=Main_Page.

QIBA has constructed a systematic approach for standardizing and qualifying volumetry as a biomarker of response to treatments for a variety of medical conditions, including cancers in the lung (either primary cancers or cancers that metastasize to the lung [18]).

## B.2 CT Volumetry for Cancer Response Assessment: Overview and Summary

Anatomic imaging using computed tomography (CT) has been historically used to assess tumor burden and to determine tumor response to treatment (or progression) based on uni-dimensional or bi-dimensional measurements. The original WHO response criteria were based on bi-dimensional measurements of the tumor and defined response as a decrease of the sum of the product of the longest perpendicular diameters of measured tumors by at least 50%. The rationale for using a 50% threshold value for definition of response was based on data evaluating the reproducibility of measurements of tumor size by palpation and on planar chest x-rays [33, 34]. The more recent RECIST criteria introduced by the National Cancer Institute (NCI) and the European Organisation for Research and Treatment of Cancer (EORTC) standardized imaging techniques for anatomic response assessment by specifying minimum size thresholds for measurable tumors and considered other imaging modalities beyond CT. As well, the RECIST criteria replace longest bi-directional diameters with longest uni-dimensional diameter as the representation of a measured tumor [35]. RECIST defines response as a 30% decrease of the largest diameter of the tumor. For a spherical tumor, this is equivalent to a 50% decrease of the product of two diameters. Current response criteria were designed to ensure a standardized classification of tumor shrinkage after completion of therapy. They have not been developed on the basis of clinical trials correlating tumor shrinkage with patient outcome.

Technological advances in signal processing and the engineering of multi-detector row computed tomography (MDCT) devices have resulted in the ability to acquire high-resolution images rapidly, resulting in volumetric scanning of anatomic regions in a single breath-hold. Volume measurements may be a more sensitive technique for detecting longitudinal changes in tumor masses than linear tumor diameters as defined by RECIST. Comparative analyses in the context of clinical trial data have found volume measurements to be more reliable, and often more sensitive to longitudinal changes in size and thus to treatment response, than the use of a uni-dimensional diameter in RECIST. As a result of this increased detection sensitivity and reliability, volume measurements may improve the predictability of clinical outcomes during therapy compared with RECIST. Volume measurements could also benefit patients who need alternative treatments when their disease stops responding to their current regimens [36-39].

The rationale for volumetric approaches to assessing longitudinal changes in tumor burden is multi-factorial. First, most cancers may grow and regress irregularly in three dimensions. Measurements obtained in the transverse plane fail to account for growth or regression in the longitudinal axis, whereas volumetric measurements incorporate changes in all dimensions. Secondly, changes in volume are believed to be less subject to either reader error or inter-scan variations. For example, partial response using the RECIST criteria requires a greater than 30% decrease in tumor diameter, which corresponds to greater than 50% decrease in tumor volume. If one assumes a 21 mm diameter spherical tumor (of 4.8 cc volume), partial response would require that the tumor shrink to a diameter of less than 15 mm, which would correspond to a decrease in volume all the way down to 1.7 cc. The much greater absolute magnitude of volumetric changes is potentially less prone to measurement error than changes in diameter, particularly if the tumors are spiculated or otherwise irregularly shaped. As a result of the observed increased sensitivity and reproducibility, volume measurements may be more suited than uni-dimensional measurements to identify early changes in patients undergoing treatment.

Table B.1 Summarizing the precision/reproducibility of volumetric measurements from clinical studies reported in the literature

| Scan | Reader | # of Readers | # of Patients | # of Nodules | Tumor Size, Mean (range) | Organ System | Volumetry, 95% CI of Measurement Difference | Volumetry, Measurement Difference % | 1D Measurement, 95% CI of Measurement Difference | 1D, Mean Measurement Difference % | Slice Thickness /Recon Interval, mm | Author, Year |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| repeat scans | intra-reader | 1 | 20 | 218 | 9.85 mm | lung, mets | -21.2 to 23.8% | 1.30% |  |  | 1.0/0.7 | Gietama et al. 2007 [40] |
| repeat scans | intra-reader | 3 | 32 | 32 | 38 mm (11–93 mm) | lung, NSCLC | -12 to 13.4% | 0.70% | -7.3% to 6.2% | -0.60% | 1.25/1.25 | Zhao et al. 2009 [8] |
| same scan | intra-reader | 1 | 10 | 50 | 6.9 mm (2.2–20.5 mm) | lung, mets | -3.9 to 5.7% | 0.90% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [41] |
| same scan | inter-reader | 2 | 10 | 50 | 6.9 mm (2.2–20.5 mm) | lung, mets | -5.5 to 6.6% | 0.50% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [41] |
| repeat scans | not specified | not specified | 10 | 151 | 7.4 (2.2–20.5 mm) | lung, mets | -20.4 to 21.9% | 1.50% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [41] |
| repeat scans | not specified | not specified | 10 | 105 | <10 mm | lung, mets | -19.3 to 20.4% | 1.70% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [41] |
| same scan (5 sets, 1 set/phase) | intra-reader ? (consensus by 2 readers), 3 x reading | 2 | 30 | 73 | ~1–9 mm [25.3 (0.2–399 mm3)] | lung, noncalcified nodules | coefficient of variance as large as 34.5% (95% CI not reported) | not reported | not reported | not reported | 0.75/0.6 | Boll et al. 2004 [42] |
| same scan | inter-reader | 2 | 33 | 229 | 10.8 mm (2.8–43.6 mm), median 8.2 mm | lung, primary or mets | -9.4 to 8.0% | 0.70% | -31.0 to 27% | -2.00% | 1.0/0.8 | Hein et al. 2009 [43] |
| same scan | inter-reader, inter-algorithms (6 readers x 3 algorithms) | 6 | 16 | 23 | not reported | lung, nodules | 55% (upper limit) | not reported | not reported | not reported | 1.25/0.625 | Meyer et al. 2006 [44] |
| same scan | intra-reader | 2 | 50 | 202 | 3.16–5195 mm3, median 182.22 mm3 | lung, mets | % not reported | 0.15 to 0.22% | % not reported | 2.34–3.73% (p<0.05 1D vs 3D) | 0.75/0.70 | Marten et al. 2006 [45] |
| same scan | inter-reader | 2 | 50 | 202 | 3.16–5195 mm3, median 182.22 mm3 | lung, mets | % not reported | 0.22 to 0.29% | % not reported | 3.53–3.76% (p<0.05 1D vs 3D) | 0.75/0.70 | Marten et al. 2006 [45] |
| same scan | inter-reader | 2 | 2239 | 4225 | 15–500 mm3 (effective diameter 3.1–9.8 mm) | lung, nodules | -13.4 to 14.5% | 0.50% | not reported | not reported | 1.0/0.7 | Wang et al. 2008 [1] |
| same scan | intra-reader | 2 | 24 | 52 | 8.5 mm (<5 to 18 mm) | lung, noncalcified nodules | 8.9 % (upper limit) | not reported | not reported | not reported | 1.25 or 2.5/not specified | Revel et al. [46] |
| same scan | inter-reader (3 readers x 3 measurements) | 3 | 24 | 52 | 8.5 mm (< 18 mm) | lung, noncalcified nodules | 6.38 % (upper limit) | not reported | not reported | not reported | 1.25 or 2.5/not specified | Revel et al. [46] |

Abbreviations: 1D = unidimensional; mets = metastasis; CI = confidence interval

The above table provides a basis for the 30% value in the Profile Claim.  The range between the minimum and maximum values in the 95% CI of the measurement difference column is mostly within +/- 15%. Considering a large study from Wang et al using 2239 patients [1], the 95% confidence interval ranged [-13.4%,  14.5%]. Thus, 30% is a conservative threshold of measurement variation. For example, the 30% change in the claim is the outside of 95% confidence interval of 15% of measurement variability when sample size is 40 or more.

## B.3 Detailed Literature Review by Indication

To date, volumetry has been evaluated in lung, liver, head and neck, esophagus, and rectal cancers, sarcoma, and lymphoma (Appendix 1, Tables 1–7). Most studies compared volumetry with either unidimensional RECIST or bidimensional WHO classifications. Volumetry showed a high degree of concordance with uni- or bidimensional assessment in some studies [47, 48]; others showed considerable discordance between these methods in response classifications [49-52]. Correlation of volumetric assessment with pathologic response was examined in four studies (two esophageal, one gastric cancer, and one sarcoma) in patients who had or were having neoadjuvant chemotherapy. Among those four studies, volumetric assessment was correlated with pathologic response in two studies (one esophageal and one gastric study) [53, 54], whereas no such correlation was found in an esophageal study [55] and a sarcoma study [56]. Two of the above neoadjuvant studies also followed esophageal cancer patients for OS or PFS, but neither study showed correlation with volumetric assessment [54, 55]. In addition, two small studies [57, 58] with lymphoma patients showed that patients with greater reduction in tumor volume at 1–2 months of chemotherapy had lower probability of relapse at one year.

#### Lung Cancer (Tables B.2 and B.3)

Lung cancer typically spreads as it advances from localized disease to the neighboring lymph nodal structures of the lung (regional metastatic spread). The most advanced stage is metastasis to a distant site such as the brain or liver. In clinical trials, depending on the initial stage at diagnosis, either progression of localized disease or the discovery of a new site of metastatic dissemination is the basis for declaring failure of the efficacy of a new drug. In virtually all lung cancer clinical trials, there are situations when either a quantitative or a qualitative endpoint may be relevant, but it is likely that quantitative endpoints will be most frequently informative in trials.

With advanced disease, there is a tendency toward more frequent disease progression at a distant metastatic site rather than progression due to extension from the primary tumor [59]. These patterns of disease progression impact clinical trial design in measuring drug response. However, there are exceptions to the pattern just described, such as bronchioalveolar carcinoma. This more indolent cancer tends to spread extensively within the lung but seldom to distant sites [60].

Staging, Therapeutic Options, and Response Assessment by Imaging Approaches

Staging defines the extent of lung cancer dissemination at the time of initial patient diagnosis. The schema for staging lung cancer has been updated recently to more accurately cluster patients who benefit from particular therapeutic interventions with predictable outcomes [61]. How staging relates to lung cancer drug therapy approaches, the imaging approaches used in those stages, and issues relative to image requirements is summarized in Table B.2 [62].

Table B.2. Summary of Image Processing Issues Relative to Lung Cancer Stage

| Stage | Percent of Cases | Percent 5-year Survival | Imaging Focus/ Therapy Focus | Imaging Tool | Issues | Thoracic Segmentation | Hi-Res |
| --- | --- | --- | --- | --- | --- | --- | --- |
| I | 16 | 49 | Primary tumor/ Neo and adjuvant RX | MDCT | Small cancers surrounded by air | Can be straightforward | Needed |
| II/III | 35 | 15.2 | Primary, hilar, and mediastinal lymph nodes/Combined modality | MDCT, PET | Larger tumors and nodes abut other structures | Often challenging | Optional |
| IV | 41 | 3 | Primary/regional nodes and metastatic sites/ Chemotherapy | MDCT, PET, bone, brain scans | Tumor response often determined outside of the chest | Often challenging | Optional |

The imaging goal may vary in different disease stages. For example, with Stage IV lung cancer, the disease progression could be due to new growth in the primary lung tumor and/or metastasis of the cancer to a distant site, and not growth of the primary cancer site. In Stage II and III lung cancer, disease progression is often manifested by increased tumor involvement in regional lymph nodes. CT imaging would typically be used to assess potential disease progression in either the primary tumor or in the lymphatic tissue. The development of new sites of metastatic disease in a Stage IV clinical trial will require a different imaging approach. To assess for new sites of metastatic disease, CT may be used to look for thoracic, hepatic, or retroperitoneal sites of metastasis, and PET scans will frequently be used to assess the progression of metastatic disease across the entire body. Common both to improving size-based measures (i.e., moving from linear diameters to volume) as well as more computationally sophisticated measures (e.g., tissue density in CT, mechanistic measures in PET) is a need for means to qualify performance across stakeholders involved in the application of these measures.

The potential utility of volumetry in predicting treatment response in lung cancer patients has been investigated by several groups. Jaffe pointed out that the value of elegant image analysis has not been demonstrated yet in clinical trials [12]. Value depends, at least in part, on the extent to which imaging endpoints meet criteria as substitute endpoints for clinical outcome measures. In this review, however, value is limited to the ability of imaging to predict either beneficial biological activity or progressive disease sooner than alternative methods of assessment, so that individual patients can move on to other treatment alternatives, or at the very least, stop being exposed to toxicity without benefit. In this context, value is predominantly a function of sensitivity and accuracy.

In 2006, Zhao and colleagues [36] reported a study of 15 patients with lung cancer at a single center. They used MDCT scans with a slice thickness of 1.25 mm to automatically quantify unidimensional LDs, bidimensional cross products, and volumes before and after chemotherapy. They found that 11/15 (73%) of the patients had changes in volume of 20% or more, while only one (7%) and 4 (27%) of the sample had changes in uni- or bidimensional line-lengths of >20%. Seven (47%) patients had changes in volume of 30% or more; no patients had unidimensional line-length changes of 30% or more, and only two patients (13%) had changes in bidimensional cross products of 30% or more. The investigators concluded that volumetry was substantially more sensitive to drug responses than uni- or bidimensional line-lengths. However, this initial data set did not address the clinical value of increasing the sensitivity of change measurements.

In a follow-up analysis [63], the same group used volumetric analysis to predict the biologic activity of epidermal growth factor receptor (EGFR) modulation in NSCLC, with EGFR mutation status as a reference. In this population of 48 patients, changes in tumor volume at three weeks after the start of treatment were found to be more sensitive and equally specific when compared to early diameter change at predicting EGFR mutation status. The positive predictive value of early volume response for EGFR mutation status in their patient population was 86%. The investigators concluded that early volume change has promise as an investigational method for detecting the biologic activity of systemic therapies in NSCLC.

In 2007, Schwartz and colleagues [38] unidimensionally and volumetrically evaluated target lesions, including lymph node, liver, peritoneal, and lung metastases, in 25 patients with metastatic gastric cancer being treated with combination therapy, and reported that volumetry predicted clinical response earlier than unidimensional RECIST by an average of 50.3 days.

In 2008, Altorki and colleagues [39] reported that volumetry is substantially more sensitive than changes in unidimensional diameters. In a sample of 35 patients with early-stage lung cancer treated with pazopanib, 30 of 35 (85.7%) were found to have a measurable decrease in tumor volume; only three of these 35 subjects met RECIST criteria for a PR.

In a retrospective analysis of 22 patients with locally advanced lung cancer treated with radiation and chemotherapy, assessment of treatment response by volume change was found to be in agreement with that by RECIST and WHO criteria (K 0.776; 95% CI 0.357–1.0 for agreement with both RECIST and WHO) [48] in 21 of 22 patients.

In another retrospective analysis of 15 patients with lung metastases from colorectal cancer, renal cell, or breast carcinoma, volumetric assessment of 32 lung lesions at baseline and after 1–4 months standard chemotherapy or radiotherapy showed fair to poor agreement with either RECIST or WHO assessment for response classification [49].

In another retrospective analysis of 68 patients with primary or metastatic lung malignancies, volumetric assessment of treatment response was found to be highly concordant with RECIST (K 0.79–0.87) and WHO assessment (K 0.83–0.84) [47]. The intraobserver reproducibility of volumetric classification was 96%, slightly higher than that of RECIST and WHO. The relative measurement error of volumetric assessment was 8.97%, also slightly higher than that of unidimensional and bidimensional assessment.

In another retrospective analysis of nine patients with lung metastases who were undergoing chemotherapy, volumetric assessment of treatment response agreed in all but one case with RECIST assessment at the patient level (K 0.69); at the lesion level, volumetric and RECIST assessment agreed on 21 of the 24 lesions (K 0.75). The level of agreement between volumetric and RECIST assessment was equivalent or superior to that of inter-observer agreement using the RECIST criteria [64].

#### Primary Liver Cancer and Metastatic Lesions in the Liver (Table B.4)

Hepatocellular carcinoma (HCC) is the most common form of liver cancer in adults [65]. The majority of patients have underlying hepatic dysfunction, which complicates patient management and trial design in the search for effective treatment [66, 67]. Despite advances in many aspects of HCC treatment, >70% of HCC patients present with advanced disease and will not benefit from existing treatment modalities, including liver transplantation, surgical resection, and loco-regional therapies. At present, only one systemic agent, i.e., sorafenib, is approved for advanced HCC patients. There remains a great need for safe and effective systemic therapies for HCC patients who progressed on or do not tolerate sorafenib and for patients with more advanced hepatic dysfunction. The liver is also a common site of metastatic spread; metastatic involvement of the liver can occur with many neoplasms, including lung, colorectal, esophageal, renal cell and breast, and stomach cancers, pancreatic carcinoma, and melanoma [11, 68].

Evidence that radiologic responses reflect clinical outcomes has recently emerged in patients who were receiving systemic therapy for advanced liver cancer. In a phase 3 trial, sorafenib, a small molecule kinase inhibitor, prolonged the survival of patients with advanced liver cancer to 10.7 months as compared with 7.9 months for the placebo group. The time to radiologic progression as defined by RECIST [69] was also significantly prolonged in the sorafenib group, in parallel with the survival advantage [70]. This survival advantage conferred by sorafenib was later confirmed in the Asian population [71].

Volumetric CT has been investigated in only a few studies in patients with metastatic liver lesions [51, 72] or HCC [73] (Appendix 1) as discussed below. These studies compared volumetry with RECIST and/or the bidimensional WHO method in classifying treatment response, and found considerable discordance between volumetry and RECIST or WHO assessment [51, 72].

Prasad and colleagues [51] compared volumetric with unidimensional (RECIST) and bidimensional (WHO) measurements in assessing response to treatment in 38 patients with liver metastases from breast cancer in a phase 3 trial. PR was defined as >65% reduction in volume; PD was defined as >73% increase in volume; and stable disease was defined as changes in volume between those in PR and PD. Patients were treated with docetaxel or capecitabine plus docetaxel, and tumors were measured at baseline and six months posttreatment. Response assessment using uni- and bidimensional methods are highly concordant (37 of 38 patients). Volumetric assessment of tumor burden was discordant with uni- and bidimensional results in 12 (32%) and 13 (34%) patients, respectively.

In another retrospective analysis of 10 patients with liver metastases from colorectal (8), esophageal (1), and gastric (1) cancers who were receiving chemotherapy, 26 pairs of pre- and posttreatment CT scans were evaluated by bidimensional criteria (WHO) and volumetry. Stable disease in the volumetric analysis was defined as between an increase in volume of less than 40% and a reduction in volume of less than 65%. Discordance between the bidimensional assessment and volumetry was found in 19–35% of the cases in disease status categories [72].

Stillwagon and colleagues [73] used volumetric measurements to assess the response to radiation and chemotherapy in 194 patients with unresectable HCC. PD was defined as 25% increase in volume; PR was defined as 30% reduction in volume; and stable disease was defined as less than 25% increase or less than 30% decrease in tumor volume.

#### Lymphoma (Table B.5)

Lymphomas comprise ~30 distinct diseases. Volumetric assessment of lymphoma has been found to correlate with treatment outcome in two early studies [57, 58] using non-helical scanners. Agreement with RECIST and WHO assessment was also found to be excellent in another study [74].

In a study of eight patients with Stage I and II diffuse large cell lymphoma of the mediastinum followed for 12 to 68 months (mean 29 months), tumor volume was assessed before and at 1 to 2 months after chemotherapy. The relative tumor volume reduction was higher in those who remained in remission than in patients who had relapsed (89% and 73% reduction, respectively). However, whether this difference was statistically significant was not reported. It was also noted that the initial tumor volume prior to chemotherapy was also greater in the group who later relapsed [57].

In a study of 12 patients with stage IA to IIB mediastinal Hodgkin’s disease who were followed for 12 to 84 months (mean 35 months) after treatment, patients with a >85% reduction in volume at 1 to 2 months after six cycles of chemotherapy had a lower incidence of mediastinal relapse (0/6, 0%) compared with those having 85% of less reduction (4/6, 67%) [58].

In a study of 16 patients with lymphoma or germ cell tumors, volumetric assessment of response to chemotherapy agreed completely with the WHO criteria in classifying responses of the lesions (20 lesions), and agreed in 18 of the 20 (90%) lesions with RECIST criteria [74].

#### Colorectal and Gastric Cancers (Table B.6)

Data suggest that volumetry may be valuable in assessing response to neoadjuvant therapy in gastric and colorectal cancers. In a prospective phase 2 study in 33 patients with resectable advanced gastric cancer who had four cycles (eight weeks) of neoadjuvant chemotherapy before surgical resection, volume reduction of primary gastric cancer correlated with histopathologic grades of regression, but the unidimensional reduction of maximum thickness and standardized uptake value (SUV) of FDG-PET did not. The optimal cut-off value of the tumor volume reduction was determined to be 35.6%, resulting in a positive predictive value and negative predictive value of 69.9% and 100%, respectively [53].

In a study of 15 patients with rectosigmoid cancer prospectively enrolled in neoadjuvant radiation therapy, using a reduction of >65% in tumor volume as the threshold for PR, volumetric analysis disagreed with the WHO criteria in classifying treatment response in one patient and with the RECIST assessment (measuring the maximal wall thickness) in four patients [75].

#### Head and Neck Cancer (Table B.7)

Head and neck cancers are clinically heterogenous, comprising multiple anatomic sites of origin with distinct natural histories and prognoses. Cure rates are low (30–50%) in locally advanced disease.

The role of volumetry in response assessment in head and neck cancer is unclear. In two retrospective studies of 129 patients with early or late stages of oral cavity or oropharynx carcinoma, assessment of response by volumetry had low agreement (38–56%) with clinical assessment by inspection and palpation [52, 76]. In the first study of 42 patients with early-stage oral cavity or oropharynx carcinoma, volume assessment of response at three to four weeks after local chemotherapy had low agreement with clinical assessment by inspection and palpation according to WHO criteria (38%) in classifying treatment response. It is noted that the lesion volume was calculated manually, assuming lesions were ellipsoid-shaped [52].

In the second retrospective study reported by the same group, 87 patients with advanced oral cavity or oropharynx carcinoma were assessed by lesion volume before and three weeks after local chemotherapy. Volume assessment of treatment response agreed with clinical assessment by WHO criteria in 49 of 87 patients (56%) [76].

#### Sarcoma (Table B.8)

The response to treatment in sarcoma is difficult to objectively measure and quantify anatomically as shown by the limited usefulness of RECIST in this setting [77, 78]. Assessment of tumor dimensions in sites such as bone, bowel, and peritoneal metastases is problematic; in addition, tumor volume reductions that can be measured by standard criteria may occur slowly or not at all (e.g., due to persistence of necrotic or fibrotic tissue).

Volumetry has not demonstrated a value in response assessment in sarcoma. In a study of 20 patients with locally advanced high-grade soft-tissue sarcoma prospectively enrolled in neoadjuvant therapy, volume assessment before and after pre-operative treatment failed to correlate with histopathologic response and was unable to differentiate histopathologic responders (n=6) from non-responders (n=14). In contrast, changes in FDG uptake measured by SUV (both mean and maximum) using PET were predictive of histopathologic response at a high accuracy (area under response operating characteristics (ROC) curve = 1.0 and 0.98, respectively) [56].

Table B.3. Evaluation of Response to Therapy by Volumetry in Lung Cancer

| Disease Stage/ Therapy | Number of Patients Evaluated | VIA Response Measurement/Timing | Comparator | Results | Statistical Analysis | Reference |
| --- | --- | --- | --- | --- | --- | --- |
| NSCLC, locally advanced/ radio ± chemo (mostly carboplatin/ paclitaxel | 22 | PR –65% | RECIST, WHO | Good concordance between 3D, 2D, 1D (cases). CR 4/4/4, PR 16/15/15, NR 2/3/3. | Kappa values. 3D vs 2D Kappa 0.776 (95% CI 0.357–1.0, substantial agreement); 3D vs 1D Kappa 0.776 (95% CI 0.357–1.0, substantial agreement); 1D vs 2D Kappa 1.0 (perfect agreement) | Werner-Wasik et al. 2001[48] |
| NSCLC, early stage gefitinib 3 wks, neoadjuvant | 48 | –24.9% (dichotomizing cut-off) | EGFR mutation sensitizing tumor to tyrosine kinase inhibitor; volume change -65% (RECIST deduced); optimal cut-off 1D (–7%) | Optimal cut-off of 3D changes 24.9%; sensitivity 90%, specificity 89% for classifying tumor w/o EGFR sensitizing mutation; PPV 86%, NPV 92%. 3D (24.9%) superior to 1D (optimal and RECIST). | Youdens' index (sensitivity + specificity −1) for determination of optimal dichromatic cut-off value; Wilcoxon rank-sum test for significance of difference | Zhao et al 2010 [63] |
| Lung mets from colorectum, renal cell, breast; standard chemo or radio | 15 | Stable disease -65% to +44%; 2 follow-ups, at 1–4 months | RECIST, WHO | Kappa 3D vs 1D 0.818 (Visit 1 to V2), 0.429 (V2 to V3); 3D vs 2D 0.412 (V1 to V2), 0.118 (V2 to V3); fair agreement 3D vs 1D; poor 2D vs 3D | Kappa values | Tran et al 2004 [49] |
| NSCLC (16), SCLC (9), lung mets of various origins (43); treatment not specified | 68 | Stable disease –65% to +44%; 3 months for lung cancer, time varied for mets | RECIST, WHO | Kappa 1D vs 3D 0.79-0.87, Kappa 2D vs 3D 0.83-0.84 | Kappa values | Sohns et al. 2010 [47] |
| Lung mets, unspecified origin; chemo | 9 (24 nodules) | Stable disease –65% to +73%; | RECIST | At nodule/lesion level, disagreement 3 in 24 nodules (Kappa 0.75); at patient level, disagree 1/9 (Kappa 0.59) | Kappa values | Fraioli et al. 2006 [64] |
| NSCLC, stage I or II, operable and resectable/ gefitinib > 21 days | 15 | –20% and –30%; 26.4 days since baseline scan | RECIST and WHO | 3D more sensitive in detecting changes. > –20%: 3D: 11/15 (73%); 1D 1/15 (7%) (p< .01); 2D 4/15 (27%)(P= .04); > –30%: 3D, 7/15 (47%); 1D 0/15 (p= .02); 2D, 2/15 (13%) (p= .06). | P values | Zhao et al. 2006 [36] |
| Mets to lymph node, liver, peritoneal and lung originated from primary gastric cancer or Gastroespoha-geal junction adenocarcinoma/irinotecan, cisplatin and bevacizumab | 25 | 3D, –65%/ 6-week follow-up for 10 cycles. 1D and 3D comparison made at the time with maximal clinical response | RECIST | 8/25 (72%) responders by both RECIST and 3D; 3D identified responders a mean of 50.3 days earlier than did RECIST criteria | There was a statistically significant (p<0.01) change in ratio of volume measurement change to RECIST measurement change for responding versus stable patients. | Schwartz et al. 2007 [38] |
| NSCLC, Stage I/II Resectable/ neoadjuvant, pazopanib 800mg qd for 2 to 6 weeks | 35 | Volume change, response criteria not specified/1 week after last dose | RECIST | 3D: 30/36 (86%) –1% to –86%. 2/35 > –50% (86% and 75%; 23/35 (66%) > –10%; 12/35 > –30%; 1D 3/25 PR (reduction 86%, 75%, and 36%). Discordance between 3D and RECIST, not head-to-head comparison in % change. 3D superiority unclear. | Not specified | Altorki et al. 2010 [39] |

Table B.4. Evaluation of Response to Therapy by Volumetry in Liver Cancer

| Disease Stage/ Therapy | Number of Patients Evaluated | VIA Response Measurement/Timing | Comparator | Results | Statistical Analysis | Reference |
| --- | --- | --- | --- | --- | --- | --- |
| Hepatic mets from GI systemic chemo +/- intra-arterial chemo | 10 (37 scans) | Stable disease –65% to +40% | WHO | Discordance between 2D and 3D: 19% and 35% (w/o or w 5% variation interval). Conclusion: 2D and 3D not interchangable. 2D tended to identify PD when 3D indicated no change. All lesions hypodense. | Not specified | Garant et al. 1999 [72] |
| Hepatic mets from breast docetaxel vs capecitabine + docetaxel | 38 | Stable disease –65% to +73% | RECIST, WHO | Treatment response concordance 1D and 2D; discordance 1D vs 3D, and 2D vs 3D | Not specified | Prasad et al. 2000 [51] |

Table B.5. Evaluation of Response to Therapy by Volumetry in Lymphoma

| Disease Stage/ Therapy | Number of Patients Evaluated | VIA Response Measurement/Timing | Comparator | Results | Statistical Analysis | Reference |
| --- | --- | --- | --- | --- | --- | --- |
| Lymphoma or germ cell; chemo | 16 (20 lesions) | Volume change/timing not specified | RECIST, WHO | 3D agreed completely with 2D, agreed in 18/20 (lesions) with 1D | Coefficient of variation calculated | Sohaib et al. 2000 [74] |
| Diffuse large cell lymphoma of the mediastinum; multiagent chemo | 8 | Volume change; 1–2 months (CT follow-up) | Relapse/ remission/ death | Patients were followed for minimum 1 yr or until death, mean 29 months (13–68 months). Reduction of tumor volume greater in pts in remission than in relapse (89% vs 73%, respectively). | No statistical analysis performed | Willett et al. 1988 [57] |
| Mediastinal Hodgkin's, stage IA to IIB; multiagent chemo | 12 | Volume change; 1–2 months (CT follow-up) | Relapse/ remission/ death | Patients were followed for minimum 1 yr or until death, mean 35 months (12–84 months). a >85% reduction in volume at 1 to 2 months after six cycles of chemotherapy had a lower incidence of mediastinal relapse (0/6, 0%) compared with those having 85% of less reduction (4/6, 67%) | No statistical analysis performed | Willett et al. 1988 [58] |

Table B.6. Evaluation of Response to Therapy by Volumetry in Colorectal and Gastric Cancers

| Disease Stage/ Therapy | Number of Patients Evaluated | VIA Response Measurement/Timing | Comparator | Results | Statistical Analysis | Reference |
| --- | --- | --- | --- | --- | --- | --- |
| Gastric, resectable advanced; chemo, neoadjuvant: oxaliplatin, infusional 5-FU, leucovorin; 8 wks (4 cycles) | 33 | Volume change cut-off 35.6%; 4, 8 wks | Pathologic response | Volume reduction of primary tumor or index node correlated best with histopathologic grades for regression, followed by short diameter of index node. No correlation with thickness of primary gastric cancer, or SUV by PET/CT. The optimal cut-off value of the tumor volume reduction 35.6%. | Spearman rank for correlation determination; ROC for determination of optimal cut-off value | Lee et al. 2009 [53] |
| Rectosigmoid; neoadjuvant radiation | 15 | PR –65%; timing not specified | Maximal wall thickness (RECIST), WHO | Discordance w RECIST and WHO (4/15 and 1/15, respectively) | Student’s t test for paired data; Pearson’s correlation test. p < 0.05 | Luccichenti et al. 2005 [75] |

Table B.7. Evaluation of Response to Therapy by Volumetry in Head and Neck Cancer

| Disease Stage/ Therapy | Number of Patients Evaluated | VIA Response Measurement/Timing | Comparator | Results | Statistical Analysis | Reference |
| --- | --- | --- | --- | --- | --- | --- |
| Oral cavity and oropharynx, carcinoma T1/2; chemo (cisplatin), intra-arterial | 42 | CR –90%, PR -50%, stable disease –50% to +25%, PR >+25%; 4 wks | Clinical inspection and palpation of lesions, classified per WHO criteria | Concordance in classifying response categories 16 of 42 pts (38%) | Not reported for concordance analysis | Rohde 2006 [52] |
| Oral cavity and oropharynx, carcinoma T3/4; chemo (cisplatin), intra-arterial | 87 | CR –90%, PR –50%, stable disease –50% to +25%, PR >+25%; 4 wks | Clinical inspection and palpation of lesions, classified per WHO criteria | Concordance in classifying response categories 49/87 pts (56%); Kappa value was not reported. | Kappa for agreement between clinical and radiological remission rates | Rohde 2007 [76] |

Table B.8. Evaluation of Response to Therapy by Volumetry in Sarcoma

| Disease Stage/ Therapy | Number of Patients Evaluated | VIA Response Measurement/Timing | Comparator | Results | Statistical Analysis | Reference |
| --- | --- | --- | --- | --- | --- | --- |
| Sarcoma, locally advanced high-grade, soft tissue; chemo (ifosfamide/doxorubicin or gemcitabine/docetaxel)± radiation | 20 | Volume change/timing not specified | Pathologic response | Volume reduction not significant pre- and post-treatment; not predictive of histopathologic response (6 responders, 14 nonresponders), AUC = 0.48 | ROC curve | Benz 2008 [56] |

Abbreviations:

1D = unidimensional measurement; 2D = bidimensional measurement; 3D = volumetric measurement; AUC = area under the curve; CI = confidence interval; CR = complete response; EGFR = epidermal growth factor receptor; FU = fluorouracil; Mets = metastasis; NSCLC = non small cell lung cancer; OS = overall survival; PFS = progression free survival; PR = partial response; PR = partial response; RECIST = Response Evaluation Criteria in Solid Tumors; ROC = response operating characteristics; SCLC = small cell lung cancer.

## Appendix C: Conventions and Definitions

Acquisition vs. Analysis vs. Interpretation: This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be computational steps that transform the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge. (The transformation of knowledge into wisdom is beyond the scope of this document.)

Image Analysis, Image Review, and/or Read: Procedures and processes that culminate in the generation of imaging outcome measures, such tumor response criteria. Reviews can be performed for eligibility, safety or efficacy. The review paradigm may be context specific and dependent on the specific aims of a trial, the imaging technologies in play, and the stage of drug development, among other parameters.

Image Header: that part of the image file (or dataset containing the image) other than the pixel data itself.

Imaging Phantoms: devices used for periodic testing and standardization of image acquisition. This testing must be site specific and equipment specific and conducted prior to the beginning of a trial (baseline), periodically during the trial and at the end of the trial.

Time Point: a discrete period during the course of a clinical trial when groups of imaging exams or clinical exams are scheduled.

Tumor Definition Variability: the clarity of the tumor boundary in the images.  It originates from the biological characteristics of the tumor, technical characteristics of the imaging process, and perhaps on the perception, expertise and education of the operator.

Technical Variability - originates only from the ability to drawing unequivocal objects. In other words, the perception of tumor definition is supposed absolutely clear and similar for any given operator when attempting to assess “Technical” variability.

Global Variability - partitioned as the variability in the tumor definition plus the “Technical” variability.

Intra-Rater Variability - is the variability in the interpretation of a set of images by the same reader after an adequate period of time inserted to reduce recall bias.

Inter-Rater Variability - is the variability in the interpretation of a set of images by the different readers.

Repeatability – considers multiple measurements taken under the same conditions (same equipment, parameters, reader, algorithm, etc) but different subjects.

Reproducability – considers multiple measurements taken where one or more conditions have changed.

## Appendix D: Model-specific Instructions and Parameters

For acquisition modalities, reconstruction software and software analysis tools, Profile compliance requires meeting the Activity specifications above; e.g. in Sections 3.2, 3.3 and 3.4.

This Appendix provides, as an informative annex to the Profile, some specific acquisition parameters, reconstruction parameters and analysis software parameters that are expected to be compatible with meeting the Profile requirements.   Just using these parameters without meeting the requirements specified in the Profile is not sufficient to achieve compliance.  Conversely, it is possible to use different compatible parameters and still achieve compliance.

Additional parameter sets may be found in QIBA Conformance Statements published by vendors and sites.  Vendors claiming product compliance with this QIBA Profile are required to provide such instructions and parameters describing the conditions under which their product achieved compliance.

Sites using models listed here are encouraged to consider these parameters for both simplicity and consistency. Sites using models not listed here may be able to devise their own settings that result in data meeting the requirements.  Tables like the following may be used by sites that wish to publish their successful/best practices.

In any case, sites are responsible for adjusting the parameters as appropriate for individual subjects.

Discussion:

It would likely be useful to include a description of the imaging subject in the following tables.

In terms of standardization, it may make sense to ask vendors to publish parameters for a known reference phantom as a stable benchmark for sites to adjust for individual patient variations.

Table D.1 Model-specific Parameters for Acquisition Devices

| Acquisition Device | Settings Compatible with Compliance |
| --- | --- |
| <Vendor> <Model> <Version> | Submitted by: |
|  | kVp Number of Data Channels (N) Width of Each Data Channel (T, in mm) Gantry Rotation Time in seconds mA Pitch Scan FoV |

Table D.2 Model-specific Parameters for Reconstruction Software

| Reconstruction Software | Settings Compatible with Compliance |
| --- | --- |
| <Vendor> <Model> <Version> | Submitted by: |
|  | Reconstructed Slice Width, mm Reconstruction Interval Display FOV, mm Recon kernel |

Table D.3 Model-specific Parameters for Image Analysis Software

| Image Analysis Software | Settings Compatible with Compliance |
| --- | --- |
| <Vendor> <Model> <Version> | Submitted by: |
|  | a b c d |

## Appendix E: Metrology Definitions and Methods

Two statistical analyses were conducted, based on the type of data: 1) variability of scalar volume measurements, including individual participant performance across test-retest repetitions as well as the performance across algorithms, and 2) comparison of segmentation boundaries relative to reference standard segmentations.  The former allows us to compare the performances of these imaging algorithms by measuring agreement of the computed result when the algorithm is held constant as well as when measured by different algorithms, regardless of the similarity in the contours that give rise to the scalar volumes; the latter provides the means by which differing algorithms may be evaluated in terms of the specific segmentation task they are performing which gives rise to the computed scalar volumes.

Variability of Scalar Volume Measurements

The models we use assume that the variance is constant across the range of the response. As a first step we determine whether the input data is skewed, which indicates a non-constant variance and perform a transformation suited to the distribution presented so as to achieve constant variance.  For this data, the measurement variation was not constant across the range of volumes; it increased with increasing volume measurements as will be illustrated below.  Typically, input data is log normal, hence the selection of log transformation as appropriate. To meet the assumptions of the analyses, a log-transformation was applied to volume.  As a result, residuals approximately follow a normal distribution which validates the conclusions of the model outputs. Whereas analyses were conducted on the log-scale, data is presented on the original scale, where possible.

Based on the transformed data, we undertook two analyses of volume measurement variability in this study, repeatability and reproducibility [22]. Repeatability refers to variability of measuring tumor volume when repeated measurements are acquired on the same subject under identical or nearly identical conditions. Thus assessment of repeatability approximates the “pure” measurement error of tumor volume measurement. Specifically, we assessed repeatability as the variance of tumor volume measurement when the marker was obtained from repeated imaging of subjects with intentionally short interval so that biological features could be reasonably assumed to have remained unchanged. We assessed repeatability for each of the several participating groups. We further assessed reproducibility, as the variability in tumor volume measurements under the condition where algorithms are not held constant.

We used visual as well as numeric methods to assess variability. Plotting test-retest replications (for repeatability) or pair-wise combinations of algorithms (reproducibility) appear as a straight line of unity in the presence of agreement.  Numerically, we denote the measurement of the jth algorithm for the ith subject at the kth replication as Yijk, where j=1,…,11, i=1, …,31, and k=1, 2.  We used a simple general model , where Yijk and (ijk are the observed value and measurement error and where ( is the population mean.  is conditional on the mean of infinite replications made on subject i by algorithm j. Both repeatability and reproducibility were assessed numerically as well as graphically by the Bland and Altman method [45, 46]. The method produces an Upper Agreement Limit (UAL) and the Lower Agreement Limit (LAL) which provides a range within which we expect 95% of the differences between replicate measures of a given algorithm (repeatability) or pair-wise measures of by two algorithms (reproducibility), are expected to lie. Repeatability was represented as differences between each test-retest repetition and plotted against the averages of the two volume measurements.

Based on these analyses, we compute multiple metrics because each provides complementary insight into performance. RC is the least significant difference between two repeated measurements on a case taken under the same conditions,

.

The interpretation of RC is that the difference between any two normally-distributed measurements on the subject is expected to fall between –RC and RC for 95% of replicated measurements [47]. The 95% tolerance interval for 95% of differences between replicated measurements is , where

and   is the thpercentile of the distribution with   degrees of freedom.

The within-subject standard of deviation (wSD) is estimated as square root of the averaged sample variances across tumors, where the sample variance is computed from the replications for each tumor. This wSD assumes that the within-tumor variance is the same across all tumors. The within-subject coefficient of variance (wCV) is a relative measure of repeatability, which we calculate as wSD/mean and thus is proportional to the magnitude of the tumor’s size.

Concordance correlation coefficient (CCC) was computed as in [48]. CCC is a measure of agreement that is a product of the correlation coefficient, penalized by a bias term that reflects the degree to which the regression line differs from the line of agreement.

Reproducibility was analyzed similarly but instead of the two repetitions, pairwise comparisons were made between algorithms.  In this case, the LOA by Bland and Altman provides a range within which we expect 95% of the differences in measurements between two algorithms to lie. The LOA are calculated as   ( t(n-1); (/2 sd(Yi1k – Yi2k) (1+1/n), where , where i, j, k, t, and n are as defined above in the repeatability section but now j varies pair-wise. Linear Mixed Effects (LME) modeling is used to separate the variability due to subject, algorithm, subject-by-algorithm interaction, and residual.  Each of these terms was considered as a random effect in the model. Model assumptions were evaluated with a Q-Q and observed-versus-fitted plots.  Use of this model determined the relative contribution to variability by the algorithm as assessed as the sum of the variability to algorithm summed with the subject-by-algorithm interaction as compared with the residual due to other factors in order to inform the QIBA claim by measuring to what extent algorithm versus other variance contributes to overall error.

As done with repeatability, we compute multiple metrics from the reproducibility analysis because each provides complementary insight into performance. The reproducibility coefficient (RDC), the interclass correlation coefficient (ICC), and the magnitude of variance explained by algorithm versus the residual variance which originates in other factors. Similar to RC, RDC is calculated as the least significant difference between two measurements taken under different conditions, in our project, by two different algorithms. The ICC is a measure of the agreement between the participating groups’ measurements of the CT volumes. The ICC is a relative index; it depends on the between-tumor variability. Since the between-tumor variability differs in magnitude for small and large tumors, the ICCs of the small and large tumors are not comparable.

In addition to computing the metrics on all tumors, two stratified reproducibility analyses were performed, one by the degree of automation used by the algorithm, and a second by retrospectively dividing into two types: (a) tumors that could be classified as meeting the conditions described in the “Claims” section of the QIBA Profile [43], and (b) tumors that did not meet these conditions. Specifically, the claims section of the QIBA profile states that the claims are only applicable “when the given tumor is measurable (i.e. tumor margins are sufficiently conspicuous and geometrically simple enough to be recognized on all images….) .and the longest in-plane diameter of the tumor is 10 mm or greater”.  Therefore, tumors described as meeting the QIBA Profile were those that were judged to have clearly identified tumor margins; all tumors used in this study exceeded the 10 mm diameter threshold.

Comparison of Segmentation Boundaries

Whereas the nature of clinical data makes actual ground truth unavailable, we can form a reference truth if one assumes that those pixels with the highest agreement among participants as being part of the tumor (or not part of the tumor), may collectively be said to be a reference segmentation. We first produced a reference segmentation using the Simultaneous Truth And Performance Level Estimation (STAPLE) method [49]. This filter performs a pixel-wise combination of an arbitrary number of input images, where each of them represents a segmentation of the same image, i.e., the segmentations performed by participant algorithms. Each input segmentation is weighted based on its "performance" as estimated by an expectation-maximization algorithm, described in detail in [50]. We then compare each individual segmentation result to this reference data, using Sensitivity (SE) or true positive rate, calculated as follows. If we define a confusion matrix C where Cuv is the number of voxels segmented with label u while the true label is v. For any label w, we calculate true positive (TP), true negative (TN), false positive (FP), and false negative (FN) as:

Typically SE is accompanied by Specificity, otherwise known as the true negative rate.  However, this quantity has a strong dependence on the size of the field of view which is constant for all participants so we omit reporting this as it is not informative. Rather, the otherwise unused TP and FN computations  are used in the calculation of  two additional spatial overlap measures, the Jaccard index [51], and Sørensen–Dice coefficients [52, 53]:

While at some point it may be evident which is the more important, for this work we compute and present all three types of numeric comparisons, collectively described as “overlap metrics.”

# Footnotes

warning that the profile data is not to be applied to all cases.

The .1 values would need to be replaced with appropriate numbers such that the bias term includes a de minimis portion of the overall error with respect to the variability (which everyone believes is the larger component of error).

# Comments

I think that the number of issues is very large and that these are only a  very incomplete subset.  Given the effort and complexity of actually making this list comprehensive, it should be removed because it is confusing and inaccurate in many cases.

Make simpler statement by first address bias, then linearity

Intercept is extrapolation down to origin

If this is done on humans, need to specify maximum time between scans

Delete one

The <fill in> values here would be informed by the 1A groundwork project, with a “sanity check” from literature sources.

?

The <fill in> values here would be informed by the 1A groundwork project, with a “sanity check” from literature sources.

Same as u6

The <fill in> values here would be informed by the 1C groundwork project, with a “sanity check” from literature sources.

The <fill in> values here would be informed by the first 3A groundwork project, with a “sanity check” from literature sources.

Should we specify the minimum time between baseline and follow-up scans to allow changes to occur?

I can’t remember why this metric has been adopted instead of the regular relative difference? What is the issue with asymmetric thresholds ? Is there any reference establishing this metric as more reliable or more sensitive?

For repeatability of the change you need two baseline and two follow-up scans.

The <fill in> values here would be informed by the 1B and second 3A groundwork projects, sanity checked from literature sources.

Takes values from the second 1B groundwork project

Again, we need 4 scans for this, two baseline and two follow-up. Need to specify which scanners should be different.

This could be a third 3A chjallenge, using 1C data

Take values from second 3A

Need to address greater complexity here.  A 40slice baseline scan that takes 14 seconds compared with a dual source scanner that takes 2 seconds at time point B might get equivalent results with different rates.

Need a metric that does not nail down to rate. E.g. HU opacification of artery in the field of view within 25% of baseline result.

Do we need a shall for oral contrast? Will deviation break the claim?

Do we need to point to papers that talk more about the relative size of the risks?  Be clear that the risks are not being ignored.

The radiation risks are small, and age and gender dependent.

We firmly believe in the benefits of imaging and quantitative imaging and strongly encourage all actors to perform imaging at the lowest possible dose.

Assuming 6-8 is easily achievable, then there is little reason to loosen.  And algorithms are not particularly asking for us to tighten to better resolution?

**Still would like to get some quantitative description of the “uniformity” from center out to edge of FOV.

6lp/cm came from ACRIN6678 experience (for lung screening)  In ACR the rods are about 8 cm from isocenter.

Rick A might have data showing issues with non-uniformity, but we haven’t seen the details yet.  The issues might be more significant when the scanner is at the outer boundary of it’s performance envelope.

The implication of basing it on the ACR phantom is that the resolution is evaluated at several points around an 8cm radius ring centered on the isocenter.  Is it fair to assume that resolution does not degrade drastically further from the isocenter?

Find out from clinicians how far out they need “coverage”.  Where in the FOV are the lesions likely to be?

VN is sensitive to many factors like IR, kernel, slice thick, kv, mas.  Number came perhaps from 6678 too.

Need an “upper” limit on the noise value to constrain the drive for low dose and higher resolution (lower slice width) from degrading the algorithm performance.  The “lower” limit prevents sacrificing dose/res too much.  And having a range makes it more likely that the two scans (baseline and change measure) are not too different in characteristics.

Really may have to do a survey of the 3A participants.  Make sure the email is clear about the performance scenario (1cm lesions) we are asking about.  Probably fairly noise insensitive, but they will need details to consider accurately.  Could offer 10-50 and ask if it is acceptable?

In the universe of clinically indicated CT exams,

Probably need a better number here.  Mike can explore a bit

Note the interactions between slice thickness, noise, resolution.  Maybe we need to build a “bounding box” that deals with them in concert rather than in isolation.  Each spec should prevent going “out of bounds” in one way or another.  Want variance reduction, not perfection.

Do we need some upper bounds on some of these? E.g. we’ve allowed 0.5mm slices.  Is that wise?

Need to start planning for when we allow Iterative.  Maybe now? And validate it in the Technical and Clinical Confirmation work.

Liver work may provide some input and work at UCLA.  Do a bit of a literature review for a proxy Groundwork Project we can reference.  Duke?

May need a requirement clarification here to highlight that each given “selected” tumor must be tracked over time, but it is not a requirement that sites agree on what tumors are selected.  I.e. this is about longitudinal labelling, not about selection.

The claim will only be met if the tumor measured at a later time point is the same as that measured at the earlier time points.

Decide whether the first two rows have become duplicative and we can drop some text.

- Enumerate the atomic requirements

Should we soften/remove this requirement and allow/accept systems that are only capable of measuring a single tumor for a patient?  In the first requirement we use the word “all” which implies multiple.

Need to either make this similar to analysis tool language (but with an allocation derived in the groundwork reader studies) or identify qualifications based on team sense of what is morst appropriate.

Can we reword this requirement to be conditional on whether the application uses segmentations to measure change?  Some applications might be able to function without creating segmentations.

Scanner resolution does vary, but there will be limits on the equipment at sites to better characterize the resolution.  Vendors probably would have such eqt/skill.

Design has an impact but so does parameters

Need some text around what it means to “resolve” – it’s subjective.  “Reasonably clear”.

Note that it’s more about being able to visualize 6 as a threshold (rather than finding the exact number you can resolve at the limit).

Could have a detailed test for vendor and looser test for site QA.  The site needs simplicity to be practical, but if there is a benefit to characterizing the device in more detail, the vendor is likely capable of a more sophisticated test.  This might let us have our cake and at least eat a bit of it too. (If there is value).  Could the vendor measurement mean that the site does not need to measure this at all?  Or is the vendor measurement about a baseline and an evaluation of the (non)uniformity of resolution so the site spot check can be put in context. (Check in with Rick and Mike and Kristin)(Numerically, how fast does it drop off?)

Note that this procedure also does recon.  An important question is whether in fact we can reasonably tease-apart the compliance of an acquisition device separate from recon software.  To do so we would need to articulate a procedure to evaluate raw projection data.  Likewise, for the reconstruction actor, we would need a dataset (rather, multiple) of raw data to allow independent recon testing.

Alternatively, we can retreat from separating these actors and combine them.  All of section 3 can remain the same, including separate section, but just have the actor be the same.  There are pros and cons to such a combination that we will want to discuss as a team.

Presently I have not yet done a detailed check on how this set of procedures, which was used in the 1C project, compares with how we came out on specs after the peer review process which took place afterwards.  I have placed it here to give an outline of the content for team discussion and assuming that we want to go this way the actual text would be updated.

I have also not yet incorporated the important issue of non-uniformity across the FOV.  What I envision is that this test will need to be done in a helical acquisition mode with plots on the effect for all reasonable tumor locations, rather than only at isocenter.  This should account for the issue of non-uniformity if we specify that it should never be out of spec even in worst-case positioning.

Does image analysis tool encompass display capabilities? Particularly, shouldn’t we consider the role of Windowing and its variable use by readers? May be this aspect deserve some specifications or at least to be commented somewhere.

Of course we have an open discussion about this, I have not edited this part for now.

Maybe we should consider allowing analysis tools to be certified as compliant with the protocol for specific body parts only. This may expedite the actual use of good tools while the vendor optimizes them for different body parts or while QIBA is assembling test data sets for all organs.

If the purpose of the test is to evaluate the real-life performance of a segmentation tool, using phantoms may not serve the purpose unless the phantoms have been validated to be realistic enough to serve whatever purpose that the test is intended to compare.

This should be representative of the tumors in the population, rather than specifically “geometrically simple”.  Most real-life tumors are not “geometrically simple”.  A segmentation tool performs relatively well for “geometrically simple” tumors does not necessary performs similarly well in a random set of tumors from the population.

Criteria such as RECIST recommends to sum up lesions for assessing tumoral burden. Shouldn’t we consider such Nb of lesions parameter in the data set design and subsequent analysis?

How is this change verified?  This is a critical issue for establishing ground truth.

This is one of the things we test by having two compliant readers work on the same data. If an algorithm is too sensitive to the location of  seed points (placed by experts), does it deserve to pass? I think not.

Using input coordinates (e.g., by input file) of a seed point will avoid the uncertainty of manual “click here” that may occur even with a dot on the image.  Also, should include several seed points randomly picked within a small “central” region of the tumor to obtain the segmentation results; the variability of the segmentation results from slightly different seed points for a given segmentation algorithm should be a part of the comparisons.  An algorithm that has large variability due to small difference in seed point location is a disadvantage in practice.

This may also be achieved by digitally shrinking or expanding the boundaries of segmented real tumors by known amounts (volume changes) to generate one of the data sets in a change test.

How about the NCI National Biomedical Imaging Archive at : https://imaging.nci.nih.gov/ncia/login.jsf

There is also: https://public.cancerimagingarchive.net/ncia/login.jsf

Would this potentially be a role for ACRIN?  Seems the most logical, but of course funding is needed.  Or would RSNA consider serving in this role?

At this pass I placed a short outline of the procedure here and the narrative explanation in a new appendix.

I can fill this one in but first let's converge on the nature of the section

Is there a criterion for acceptance?

Aren’t we testing a specific algorithm for compliance with the protocol? Same for 1.2.3.4. Each test should stand on its own and not compared to others.

The precision profiles defined above do not include this measure. Why do we need it?

May include additional measures such as average distance  and maximum distance between the segmented boundaries.  Definitions of some of these distances can be found in the literature.

I can expand this for the bias and linearity as well but first let's converge as a group on the nature of this appendix.

# Text boxes

Note to users – when referencing this QIBA Profile document, please use the following format:

CT Volumetry Technical Committee.  CT Tumor Volume Change Profile, Quantitative Imaging Biomarkers Alliance.  Version 2.2.  Reviewed Draft. QIBA, August 8, 2012.  Available at:  http://rsna.org/QIBA_.aspx