QIBA scans for performance protocols

Phantom: ACR (section 3&4)

Scanner: GE  HD750

Data channel: 64x0.625 mm

kVp: 120

Pitch: 0.984:1

Rotation time: 1.0 second

Scan length: 112.5 mm

Recon thickness: 1.25 mm

Scan FOV1:  32 cm

| mA Kernel | 80 CTDI=5.69 | 95 CTDI=6.76 | 150 CTDI=10.68 | 200 CTDI=14.24 | 250 CTDI=17.79 |
| --- | --- | --- | --- | --- | --- |
|  | Res. (lp/cm) | noise | Res. (lp/cm) | noise | Res. (lp/cm) | noise | Res. (lp/cm) | noise | Res. (lp/cm) | noise |
| standard | 6 | 19.10 | 6 | 18.19 | 6 | 14.91 | 6 | 12.68 | 6 | 11.14 |
| detail | 7 | 24.86 | 7 | 23.17 | 7 | 18.46 | 7 | 15.53 | 7 | 13.78 |
| ASIR (std, 50%) | 6 | 14.10 | 6 | 12.32 | 6 | 10.27 | 6 | 8.97 | 6 | 7.52 |

Scan FOV2:  50 cm

| mA Kernel | 80 CTDI=6.52 | 95 CTDI=7.74 | 150 CTDI=12.22 | 200 CTDI=16.30 | 250 CTDI=20.37 |
| --- | --- | --- | --- | --- | --- |
|  | Res. (lp/cm) | noise | Res. (lp/cm) | noise | Res. (lp/cm) | noise | Res. (lp/cm) | noise | Res. (lp/cm) | noise |
| standard | 6 | 20.54 | 6 | 18.52 | 6 | 15.28 | 6 | 12.63 | 6 | 11.80 |
| detail | 7 | 25.55 | 7 | 23.14 | 7 | 18.06 | 7 | 16.09 | 7 | 14.92 |
| ASIR (std, 50%) | 6 | 13.25 | 6 | 12.91 | 6 | 9.82 | 6 | 8.56 | 6 | 8.12 |

Resolution images (please zoom in to view)

Scan FOV1:  32 cm

| mA Kernel | 80 CTDI=5.69 | 95 CTDI=6.76 | 150 CTDI=10.68 | 200 CTDI=14.24 | 250 CTDI=17.79 |
| --- | --- | --- | --- | --- | --- |
| standard |  |  |  |  |  |
| detail |  |  |  |  |  |
| ASIR (std, 50%) |  |  |  |  |  |

Scan FOV2:  50 cm

| mA Kernel | 80 CTDI=6.52 | 95 CTDI=7.74 | 150 CTDI=12.22 | 200 CTDI=16.30 | 250 CTDI=20.37 |
| --- | --- | --- | --- | --- | --- |
| standard |  |  |  |  |  |
| detail |  |  |  |  |  |
| ASIR (std, 50%) |  |  |  |  |  |

# Text boxes

Standard: 95 – 130 mAs

Detail: 155 – 200 mAs

ASiR:  unknown, but lower than 80 mAs

Standard: 105 – 140 mAs

Detail: 150 – 205 mAs

ASiR:  unknown, but lower than 80 mAs