## Sheet1
Reference | Setting | Nodule sizes | CT Scanner | Scan number | kVp | effective mAs | mAs | mA | Rotation time (sec) | Pitch | Detector collimation (mm) | Section thickness (mm) | Section Increment (mm) | Kernel | FOV (mm) | Software | Segmentation | Error
Goo et al. Radiology 2005; 235:850-856 | Synthetic phantom | 3.2, 4.8, 6.4, 12.7 mm | Siemens Sensation 16 | 1 | 120 | x | 127.5 | 255 | 0.5 | x | 0.75 | 0.75, 1.0, 2.0, 3.0, 5.0 | 0.5, 1.0, 2.0 | B60f | 100, 200, 300 | Infinitt Rapidia | Automated | For 1 mm sections: 4.8 mm nodule-14%; 6.4 mm nodule-9.4%; 12.7 mm nodule-5.4%
Das et al. Invest Radiol 2007; 42:297-302 | Synthetic phantom | 3-10 mm | Siemens Volume Zoom, Sensation 16, Sensation 64 | 1 | x | "standard and low dose protocols" | x | x | x | 1 | VZ-1 and 2.5 mm; S16-0.75 and 1.5 mm; S64-0.6 and 1.0 mm | "thinnest possible" | "reasonable overlap" | B50f | 300 | Siemens LungCare | Automated | 4.5-13% depending on scanner and nodule type
Ravenel JG et al. Radiology 2008;247:400-408 | Synthetic phantom | 3-16 mm | GE Lightspeed (16) | 1 | 120 | 120 | 60 | 120 | 0.5 | 0.532 | X | 0.625, 1.25, 2.5, 5.0 | x | STD, lung, bone, bone+, edge, detail, soft | 96, 150, 260, 360 mm | GE VCAR | Automated | For nodules >6 mm: 5.5% at 36 cm FOV (spherical; -2.3% at 36 cm FOV (lobulated)
Chen B et al. Phys Med Biol 2012; 57:1335-1348 | Synthetic phantom | 5-10 mm | GE Lightspeed VCT | 1 | 120 | CTDIvol 6.22 mGy | x | x | x | 1.375 | x | 0.625, 1.25, 2.5 | X | Standard | x | GE LungVCAR | Automated | 5 and 10 mm acrylic nodules-6-7% at 0.625 and 1.25 mm thickness; 10 mm polypropylene nodule-2-3% at 0.625 and 1.25 mm thickness
Willemink MJ et al. AJR 2012; 199:1241-1246 | Synthetic phantom | 3, 5, 8, 10, 12 mm | Philips Brilliance iCT (256) | 1 | 80, 100, 120 | 19-76 | 25, 50, 100 | x | 0.33 | 0.76 | 0.625 | 0.9 | X | C, iDose iterative (levels 2, 4, 6) | x | Philips IntelliSpace | Automated | For nodules 5-12 mm: -24% to 13%
Chen, B et al. Med Phys 2013; 40(11) | Synthetic phantom | 4.76 and 9.53 mm | GE Discovery CT750 HD | 5 | 120 | CTDI 7.5 mGy | X | X | 0.5 | 1.375 | x | 0.625, 1.25, 2.5 | equal to slice thickness | Detail | 360 | GE Lung VCAR and TeraRecon iNtuition | Automated | For 0.625 and 1.25 mm slice thickness: Lung VCAR 4-8% for 9.53 mm nodules and 5-10% for 4.76 mm nodules; iNtuition -3 to -10 % for 9.53 mm nodules and 5-23% for 4.76 mm nodules
Gavriledes MA. Acad Radiol 2013;20:173-180 | Synthetic phantom | 5, 8, 10, 20 mm | Siemens Somatom 64 | 10 | x | x | 25, 100, 200 | x | x | 0.9 and 1.2 | 0.6 | 0.75, 1.5, 3.0 | Contiguous and 50% overlap | B40f and B60f | 400 | Custom | Matched filter-based estimation | For 5-10 mm nodules, 0.75 and 1.5 mm slice thickness: 10-25% with contiguous recon, 5-9% with overlapping recon
Coenen A et al. Jpn J Radiol online 18 Aug 2013 | Synthetic phantom | 5, 8, 10, 12 mm | GE Discovery 750 HD | 2 doses | 120 | 60 and 250 | 60 and 250 | 120 and 500 | 0.5 | 0.984 | 0.625 | 0.625 | x | Bone and 0, 50%, and 100% iterative | 200 | GE Lung VCAR | Automated | 5 mm-19%; 8 mm-4.6%; 10 mm; 3.8%; 12 mm; 2.4%
Xie X et al. AJR 2014; 202:W202-W209 | Synthetic phantom | 5.1-88.4 mm^3 (<5.6 mm) | Siemens Sensation 64 and Philips Brilliance 64 | 3 | 120 | Siemens 13.3; Philips 14.4 | 20 | 40 | 0.5 | Siemens 1.5; Philips 1.39 | Siemens 0.6; Philips 0.625 | 1 | 0.7 | Siemens B30f; Philips B | 300 | Siemens LungCARE | Automated | -39%