## Checklist Resolutions
Checklist items (See Comment Resolutions tab for feedback comments) |  |  | Site A |  | Site B |  | Site C |  | Site D |  | Committee Discussion | Resolution | Status
|  |  | Conform | Opinion | Conform | Opinion | Conform | Opinion | Conform | Opinion |  | 0 | TBD | To be discussed
[12 empty] | 0 | OK | No action requested
[12 empty] | 0 | Discuss | Need to decide resolution
[12 empty] | 0 | TODO | Resolution decided
Pre-Aacquistion and QA |  | Actor | [9 empty] | 0 | Done | Profile update completed
3 | Site Conformance Check
3.0.2.1 | Ultrasound Scanner | QA Manager | Y | Y | Y | Y | Y | Y
3.0.2.2 | Technologist / Sonographer / Radiologist | QA Manager | Y | Y | Y | Y | Y | Y
3.1 | Staff Qualification
3.1.2.1 | Operator Training | Tech or Rad | Y | Y | Y | Y | Y | Y |  |  |  |  | Site A | University of Rochester | GE
3.3 | Pre-delivery | [11 empty] | Site B | UT Southwestern | Siemens and Philips
3.3.2.1 | System Conformance | MFR | Y | Y | Y | Y | Y | Y |  |  |  |  | Site C | University of Alabama Birmingham | Philips
3.3.2.2 | Hardware & transducer MFR specified parameters | MFR | Y | Y | Y | Y | Y | Y |  |  |  |  | Site D | N/A
3.3.2.3 | SWS Measurement Consistency | MFR | Y | Y | Y | Y | Y | Y
3.4 | Installation (including software and hardware upgrades)
3.4.2.1 | Software verification | QA Manager |  |  | Y | Y | Y | Y
3.4.2.2 | SWS Measurement Concordance | QA Manager | N |  | N | FN | N | FW
3.5 | Periodic Quality Assurance (QA)
3.5.2.1 | US Imaging QA | QA manager | Y | Y | Y | Y | Y | Y
3.5.2.2 | SWS Measurement Consistency & System QA Testing | QA manager | N |  | N | FN | N | NF
3.5.2.3 | Operator Training & Qualification Testing | QA manager | Y | Y | Y | Y | Y | FW
3.5.2.4 | US Imaging / SWS Phantom Characterization & Stability Testing | QA Manager | N |  | N | FN | N | NF
Subject and Data Acquisition
3.7 | Subject Selection
3.7.2.1 | Prior Surgery | Tech or Rad | Y | Y | Y | Y | Y | Y
3.7.2.2 | Breathing | Tech or Rad | Y | Y | Y | FW | Y | Y
3.8 | Subject Handling
3.8.2.1 | Fasting State | Tech or Rad | Y | Y | Y | FW | Y | Y
3.8.2.2 | SWS Acquisition System | Tech or Rad | Y | Y | Y | Y | Y | Y
3.9 | SWS Image Acquisition (SWEI) and Point SWS Measurement
3.9.2.1 | Patient Position | Tech or Rad | Y | Y | Y | Y | Y | Y
3.9.2.2 | Transducer Position | Tech or Rad | Y | FW | Y | Y | Y | Y
3.9.2.3 | Ultrasound Image | Tech or Rad | Y | Y | Y | Y | Y | Y
3.9.2.4 | Measurement Region of Interest (ROI) Use and Placement | Tech or Rad | Y | Y | Y | FW | Y | Y
3.9.2.5 | Respiration | Tech or Rad | Y | FW | Y | Y | Y | Y
3.9.2.6 | Liver Movement | Tech or Rad | Y | Y | Y | Y | Y | Y
3.9.2.7 | Measurement Location Consistency | Tech or Rad | Y | Y | Y | Y | Y | Y
3.9.2.8 | Number of Measurements | Tech or Rad | Y | Y | Y | Y | Y | Y
3.9.2.9 | SWS Imaging - Quality Assessment | Tech or Rad | Y | Y | Y | Y | Y | Y
3.9.2.10 | SWS Imaging - Measurement ROI Placement | Tech or Rad | Y | Y | Y | Y | Y | Y
<Paste the Section # and Title of each assessment procedure> <Opinion here reflects the practicality of performing the procedure itself. Opinion on the associated requirement (pass/fail threshold value) is shown above in the corresponding checklist row. Comments on specific details of the procedure go on the Comment Resolutions tab>

## Comment Resolutions
Commenter | Section | Line # | Priority | Issue (See also Checklist tab) | Proposal | Committee Discussion | Owner (opt.) | Resolution (w Rationale if rejected) | Status
[8 empty] | 11 | TBD | To be discussed
[8 empty] | 0 | OK | No action requested
IMPORTANT: Read the Instructions Tab FIRST | [7 empty] | 0 | Discuss | Need to decide resolution
[8 empty] | 0 | TODO | Resolution decided
[8 empty] | 0 | Done | Profile update completed
|  | Use line# of PC PDF, they're stable |  | Describe your issue. Don’t write a book. Do include enough so it's clear what you see as a problem. | Suggest new wording or describe a way to address your issue. The committee may simply accept your suggested text. Even if they don’t, it gives a good sense of what you’re looking for. Leaving this blank means you can’t imagine how to resolve the issue | Cmte Notes as needed. | BC member who will lead resolution of this issue <Delete this column if not used> | How the comment was/will be resolved. BC may simply accept & change as proposed, may accept issue & resolve differently, or may reject issue (with a rationale for why).
UAB | 3.4.2.2 |  |  | UAB: Old phantom deteriorated; new phantom ordered will have areas where theoretically elastography can be measured. QA performed as per ACR accreditation standards. To conform, will depend on length of time to perform and if mandated by accreditation programs |  |  |  |  | TBD | To be discussed
UAB | 3.5.2.2 |  |  | No; not practical to purchase additional verasonics phantom due to limited institutional funds. Perhaps equipment manufacturer or central respository can loan. Who will do this? What will the manufacturer do when a change of > 5 % if found without institution payment (for software upgrade, for example) |  |  |  |  | TBD | To be discussed
UAB | 3.5.2.3 |  |  | Yes; No fornal continuing annual QC but large volume of elastography studies per sonographer |  |  |  |  | TBD | To be discussed
UAB | 3.5.2.4 |  |  | We don't have a phantom, institution unlikely to pay for additional outside testing of phantom if acquired. Possibly QA manager could perform, now just visual evaluation of phantom while testing. |  |  |  |  | TBD | To be discussed
UAB | 3.8.2.2 |  |  | Comment: May have to use alternate manufacturer if manufacturer selection changes |  |  |  |  | TBD | To be discussed
UTSW | 3.3.2.1, 3.3.2.2, 3.3.2.3 |  |  | Required personal outreach to manufacturer; had to create a standard form to facilitate; | Post standard conformance form to Wiki page for others to use in the future |  | DF | No adjustment to Profile needed. This has already been discussed with the Process committee | TBD | To be discussed
UTSW | 3.7.2.2 |  |  | Pt unable to hold breathe for 4-8 seconds. Lots of Liver movement. | Document in Tech Notes and report to Radiologist |  | DF | No adjustment to Profile needed. Already allows for patient-assocaition variabilty that should be reported to Radiologist for QC | TBD | To be discussed
UTSW | 3.9.2.4 |  |  | Pt Liver had a lot of reverb, had to sample deeper | Document in Tech Notes and report to Radiologist |  | DF | No adjustment to Profile needed. Already allows for patient-assocaition variabilty that should be reported to Radiologist for QC | TBD | To be discussed
UTSW | 3.8.2.1 |  |  | Patient presented non-facting and refused to reschedule | Document in Tech Notes and report to Radiologist |  | DF | No adjustment to Profile needed. Already allows for patient-assocaition variabilty that should be reported to Radiologist for QC | TBD | To be discussed
UR | 3.9.2.2 |  |  | Parallel was not ideal for this exam | Document in Tech Notes and report to Radiologist |  | SMcA | No adjustment to Profile needed. Already allows for patient-assocaition variabilty that should be reported to Radiologist for QC | TBD | To be discussed
UR | 3.9.2.5 |  |  | Suspend tidal resp did not work | Document in Tech Notes and report to Radiologist |  | SMcA | No adjustment to Profile needed. Already allows for patient-assocaition variabilty that should be reported to Radiologist for QC | TBD | To be discussed

## Instructions
IMPORTANT: Do not edit this template spreadsheet.
FIRST: Use File->Make a copy... to copy this template into a new file for your comments
UPDATE the file name (upper left corner of your new sheet window) which says "Copy of ..." to include your Profile Name and remove Copy of
NOW you can edit your copied document.
See http://qibawiki.rsna.org/index.php/Public_Comment_Process for more guidance on the comment resolution process
Checklist Resolutions Legend
The "Checklist Resolutions" tab is for collating the checklist responses submitted by all the Technical Confirmation sites | Site Conformance with the requirement during Technical Feasibility testing
| Y | Conformed
The "Comment Resolutions" tab is for collating comments submitted by all the Technical Confirmation sites | N | Did not conform
Both tabs support Biomarker Committee assessment of the feedback and potential modification of the Profile | Site Opinion of the requirement
| R | Do routinely
| FW | Feasible, will do to conform | If site reports didn't conform but feasible, ask for details on why not done
TODO Incorporate QIBA User Feedback Template content. | FN | Feasible, but will not do to conform | Ask for details on why not
| NF | Not feasible | Ask for details on why not