# Profile: DCEMRI Quantification

QIBA DCEMRI Sub-committee

Date: July 15, 2010

Draft Version 0.6

I. CLINICAL CONTEXT       M. Schnall

A growing understanding of the underlying molecular pathways active in cancer has led to the development of novel therapies targeting VEGF, EGFR-tk, PI3-k, mTOR , Akt  and other pathways. Unlike the conventional cytotoxic chemotherapeutic agents, many of the molecularly-targeted agents are cytostatic, causing inhibition of tumor growth rather than tumor regression. One example is anti-angiogenesis agents, which are presumed to act through altering tumor vasculature and reducing tumor blood flow. In this context, conventional endpoints such as tumor shrinkage may not be the most effective means to measure therapeutic responses. Functional imaging is an important candidate biomarker to predict and monitor targeted treatment response and to document pharmacodynamic response.

Dynamic Contrast Enhanced magnetic resonance imaging (DCE-MRI) represents an MRI-based method to assess tumor vascularity by tracking the kinetics of a low-molecular weight contrast agent intravenously administered to patients that highlights the tumor vasculature. The emerging importance of angiogenesis as a cancer therapy target makes assays of vascularity important to clinical research and future clinical practice related to targeted cancer therapy. There are multiple literature reports of the application of DCE-MRI to predict and detect changes associated with angiogenesis targeted therapy (Wedam, 2006, Rosen, 2004; Dowlati, 2002; Stevenson, 2003, Morgan et al, JCO 2003, Flaherty et al Cancer Biol Ther 2008, Liu et al JCO 2005, Drevs et al JCO 2007).  Further, there is interest in the application of quantitative DCE-MRI to characterize contrast enhancing lesions as malignant in several organ systems including breast and prostate.

QIBA recognizes the potential importance of DCE-MRI as a functional imaging biomarker of angiogenesis.  As a result, DCE-MRI QIBA committee has been formed to define the basic standards for DCE-MRI measurements and quality control.

References

Dowlati A, Robertson K, Cooney M, et al. A phase I pharmacokinetic and translational study of the novel vascular targeting agent combretastatin a-4 phosphate on a single-dose intravenous schedule in patients with advanced cancer. Cancer Research 2002;62:3408-3416

Drevs J, Siegert P, Medinger M, Mross K, Strecker R, Zirrgiebel U, Harder J, Blum H, Robertson J, Jürgensmeier JM, Puchalski TA, Young H, Saunders O, Unger C. Phase I clinical study of AZD2171, an oral vascular endothelial growth factor signaling inhibitor, in patients with advanced solid tumors. J Clin Oncol. 2007 Jul 20;25(21):3045-54

Flaherty KT, Rosen MA, Heitjan DF, Gallagher ML, Schwartz B, Schnall MD, O'Dwyer PJ Pilot study of DCE-MRI to predict progression-free survival with sorafenib therapy in renal cell carcinoma. Cancer Biol Ther. 2008 Apr;7(4):496-501

Liu G, Rugo HS, Wilding G, McShane TM, Evelhoch JL, Ng C, Jackson E, Kelcz F, Yeh BM, Lee Jr FT, Charnsangavej C, Park JW, Ashton EA, Steinfeldt HM, Pithavala YK, Reich SD, Herbst RS (2005) Dynamic Contrast-enhanced Magnetic Resonance Imaging as a pharmacodynamic measure of response after acute dosing of AG-013736, an oral angiogenesis inhibitor, in patients with advanced solid tumors: results from a phase I study. J Clin Oncol 23: 5464–5473

Morgan B, Thomas AL, Drevs J, Hennig J, Buchert M, Jivan A, Horsfield MA, Mross K, Ball HA, Lee L, Mietlowski W, Fuxuis S, Unger C, O'Byrne K, Henry A, Cherryman GR, Laurent D, Dugan M, Marme D, Steward WP (2003) Dynamic contrast-enhanced magnetic resonance imaging as a biomarker for the pharmacological response of PTK787/ZK 222584, an inhibitor of the vascular endothelial growth factor receptor tyrosine kinases, in patients with advanced colorectal cancer and liver metastases: results from two phase I studies. J Clin Oncol 21: 3955–3964

Rosen M, Veronese M, Lee Richard, Schwatz B, O’Dwyer P, Flaherty, K. Dynamic Contrast- enhanced MRI (DCE-MRI) of Primary and Metastatic Renal Cell Carcinoma in Humans: Measurement of Tumor Vascularity as a Means of Assessing Anti-Vascular Effects of BAY-43- 9006 in Vivo, RSNA Annual Meeting, Chicago, December, 2004

Stevenson JP, Rosen M, Sun W, et al. Phase I trial of the antivascular agent combretastatin A4 phosphate on a 5-day schedule to patients with cancer: magnetic resonance imaging evidence for altered tumor blood flow. Journal of Clinical Oncology 2003;21:4428-4438

Wedam SB, Low JA, Yang SX, Chow CK, Choyke P, Danforth D, Hewitt SM, Berman A, Steinberg SM, Liewehr DJ, Plehn J, Doshi A, Thomasson D, McCarthy N, Koeppen H, Sherman M, Zujewski J, Camphausen K, Chen H, Swain SM (2006) Antiangiogenic and antitumor effects of bevacizumab in patients with inflammatory and locally advanced breast cancer. J Clin Oncol 24: 769–777

II. CLAIMS

Can measure quantitative tumor vascular properties - Ktrans and IAUGC - from Dynamic Contrast Enhanced MRI at 1.5T using low molecular weight Gd-based agents  within a 20% test-retest coefficient of variation for tumors that are more than 2cm fixed and 3cm moving (reword tumor size part)

III. PROFILE DETAIL/PROTOCOL

- Executive Summary     Jeff E.

Word about what is the state of art in research and clinical trials.

Why would standardization help.

Few sentences on what this profile is for.

- Context of the Imaging Protocol within the Clinical Trial     Jeff E.

Alterations in quant. Parameters in response to therapy

- Site Selection, Qualification and Training      Gudrun

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large patient population under consideration. For DCE-MRI use as quantitative imaging biomarker it is essential to put some effort into an imaging capability assessment prior to final site selection for a specific trial. For imaging it is important to consider the availability of:

- appropriate imaging equipment,

- appropriate injector equipment and contrast media  (cp 5.),

- experienced MR technicians for the imaging procedure and

- processes that assure imaging protocol compliant image generation at the correct point in time.

Imaging equipment qualification :

1.5 T MR machines with 60-70 cm bores need to be available. The scanner needs to be under quality assurance and quality control according to local requirements (including maintenance schedule). The used software version should be identified. It might be beneficial to identify and qualify a second scanner at the site, if available. If this is done prior to the study start there will be no difficulties later on in case the first scanner is temporarily unavailable.

Injector Qualification

A power injector is required for DCE-MRI studies. It needs to be properly serviced and calibrated.

MR Technicians

MR technicians running DCE-MRI procedures should be MR certified according to local regulations. Technician should be experienced in DCE-MRI. The person should be experienced in clinical study related imaging and should be familiar with GCP. A qualified backup person is needed that should fulfill the same requirements. Contact details for both technicians should be available in case of any questions.

Imaging qualification process

The above mentioned details can be obtained using a simple questionnaire as pre-qualification step.

If appropriate equipment and personnel is available a site visit is recommended. During the site visit study related imaging protocols are discussed and, ideally, all scan parameters are entered at the MR scanner.

To qualify the scanner a phantom imaging process is recommended. QIBA DCE-MRI phantom or other qualified phantoms should be used (?? Should we provide a reference list?).

Before the phantom scan the expected outcome needs to be defined. Then the phantom(s) should be scanned using the imaging sequences as defined in the imaging protocol. Data analysis as needed should be performed to generate the results that can be compared to the originally defined desired outcome. After passing phantom imaging a volunteer scan might be considered. All imaging sequences except contrast media based imaging will contribute to further imaging capability assessment of the site.

All those results shall be documented and if passed constitute the site qualification for DCE-MRI procedure. This process ensures study specific training of the site personnel and needs to be documented and signed.

Phantom scan should be repeated every 3 months during the course of the study.  Ongoing image quality inspection on a per scan basis is essential.

Any changes to scanner equipment, software versions etc. need to be documented and will result in the need of imaging qualification renewal.

- Subject Scheduling        Alexander G.

- Utilities and Endpoints of the Imaging protocol within the Clinical Trial

  - Claim:  This image acquisition and processing protocol should be sufficient to measure quantitative tumor vascular properties – Ktrans and IAUC – from Dynamic Contrast enhanced MRI at 1.5T using low molecular weight Gadolinium (Gd) based agents within a 20% test-retest coefficient of variation for tumors (2cm in size fixed, and 3cm moving)

- Management of Pre-enrollment Imaging Tests

  - The history of prior medical imaging procedures that might, or might not, be used as part of the selection criteria for enrolling patients in a clinical trial that uses this protocol is outside the scope of QIBA DCE-MRI committee.  However, only image acquisition and processing protocols that conform to, or exceed, the minimum design specifications described in this protocol are sufficient for quantifying tumor vascular parameters with the precision of measurement specified in the profile claims document.  In practice, this will often require “baseline” scans to be repeated according to these guidelines when the objective is to quantify longitudinal changes within subjects.

- Timing of Imaging Tests within the Clinical Trial Calendar

  - The DCE MRI committee believes that all baseline evaluations should be performed as close as possible to the treatment start.  Otherwise, these imaging procedures are not time sensitive.  The interval between follow up scans within patients may be determined by current standards for good clinical practice or the rationale driving a clinical trial of a new treatment

- Subject Selection Criteria related to Imaging

  - Please see Appendix (?) for absolute contraindications to MRI.  These include, but are not limited to the following:

    - Pregnancy

    - Presence of a defibrillator device

    - Presence of a pacemaker

    - Any metallic fragments within the vitreous of the eye

  - The QIBA DCE-MRI committee acknowledges that there are potential and relative contraindications to MRI in patients suffering from claustrophobia. Methods for minimizing this risk are at the discretion of the physician caring for the patient.

  - The QIBA DCE-MRI committee acknowledges that there are potential risks associated with the use of contrast material.  The default recommendations for intravenous contrast that follow assume there are no known contra-indications in a particular patient other than the possibility of an allergy to gadolinium.  The committee assumes that local standards for good clinical practice (cGCP) will be substituted for the default in cases where there are known risks, e.g., patients with chronic renal failure.

- Subject Preparation      Alexander G.

- Interval Timing (e.g., oral and/or IV intake, vigorous physical activity, timing relative to non-protocol-related medical interventions, etc.).

  - There are no specific patient preparation procedures for the MRI scans described in this protocol. The DCE-MRI committee acknowledges that there are specifications for other procedures that might be acquired contemporaneously, such as requirements for fasting prior to FDG PET scans or the administration of oral contrast for abdominal CT. Those timing procedures may be followed as indicated without adverse impact on these guidelines

- Specific Pre-Imaging Instructions

  - Prior to Arrival

    - The local standard of care for acquiring MRI scans may be followed. For example, patients may be advised to wear comfortable clothing, leave jewelry at home, etc

  - Upon Arrival (including ancillary testing associated with the imaging and downstream actions relative to such testing)

    - Detail: Staff shall prepare the patient according to the local standard of care.

      - Patients should be assessed for any removable metal objects on their bodily surfaces that will be in the field of view.

      - Patient should be "comfortably positioned", in "comfortable clothes to minimize patient motion and stress (which might affect the imaging results) and any unnecessary patient discomfort.

- Detail: Bladder State

  - Ideal: micturation immediately prior to being placed in the MRI machine

  - Target: empty bladder

  - Acceptable: any

  - The target here is purely for patient comfort.

- Note: Factors that adversely influence patient positioning or limit their ability to cooperate should be recorded in the corresponding DICOM tags and case report forms, e.g., agitation in patients with decreased levels of consciousness, patients with chronic pain syndromes that limit their ability to cooperate with requirements for breath holding or remaining motionless, etc.

- Imaging-related Substance Preparation and Administration    Alexander G.

Agent (low molecular weight Gd based ECF agents), Equipment, dose 0.1 mmol/kg), injection site, timing, injection rate (2-3 cc/sec), gauge

“It should be noted that there is limited data with newer contrast agents (such as protein binding agents, ….”

- Imaging Agent Preparation and Specification (Contrast agent or radiopharmaceutical)

  - The DCE-MRI committee acknowledges that the use of intravenous contrast material is often medically indicated for the diagnosis and staging of cancer in many clinical settings.

- Contrast administration: (agent, dose, route)

  - Ideal:

    - Each subject should have an intravenous catheter with a gauge no smaller than 20 gauge which should be placed in the right antecubital fossa.  Injection through a port-a-catheter, or permanent indwelling catheter is not allowed.

    - Contrast should be administered in a dynamic fashion, preferably with a power injector. At baseline and at each subsequent time-point, the same dose of contrast and rate of contrast administration should be performed as clinically safe.  The rate of administration should be no less than 3 cc/sec.

  - Target: Same rate and dose of contrast administration, and exact same start time of scans relative to contrast administration. Sites should use the same brand of contrast each time they scan a particular patient.

  - Acceptable: Exactly the same contrast agents and administration procedures must be used in each examination.  If the patient cannot get an i.v., then the patient should be withdrawn from the study.

- Contrast Dose Reduction Based On Creatinine Clearance: (renal function)

  - An extracellular Gadolinium based contrast agent (e.g. Gd-DTPA) will be utilized.

  - Patient’s renal creatine clearance should be estimated, and estimation of gloemerular filtration rate (eGFR) determined in adults through well known and adopted formulas.

    - If eGFR < 60, then the subject should withdraw from the study.

- Individual Subject Imaging-related Quality Control      Mark Rosen

Mitigate gross patient motion (claustrophobia, anxiety)

Metal artifact (within FOV) – exclusion

Patient size

Must get good anatomic image

- Imaging Procedure     Sandeep/Ed J.

This section describes the imaging protocols and procedure for conducting a DCEMRI exam. Suitable localizer (scout) images must be collected at the start of exam, and used to confirm correct coil placement as well as selection of appropriate region to image.

The DCEMRI exam will consist of the following three components:

- Ratio Map Protocol, for the acquisition of data useful for signal intensity normalization.

- Variable Flip Angle T1 Mapping protocol, for the acquisition of data that allows the computation of a pre-contrast T1 map.

- DCEMRI protocol, a 3D T1-weighted gradient echo sequence, which collects dynamic data during the passage of contrast agent.

These are detailed below.

  - DCEMRI Protocol:

3D fast spoiled gradient recalled echo (FSPGR) or equivalent

Body XMT

Phased array receive coil

Extended dynamic range

No parallel imaging

No magnetization preparation

Coronal Oblique acquisition (including aorta)

Frequency encode S/I

TE as short as possible (<1 ms)

TR as short as possible (3-6 ms)

Temporal resolution ≤10 sec

Flip angle: 30 deg

±31.25 kHz receiver bandwidth (i.e., 250 Hz/pixel)

Region appropriate FOV (recommend 42 x 34 cm)

80% phase encoding FOV (34 cm for a 42 cm FE FOV)

Partial Fourier (“fractional echo”) as needed

As many slices per acquisition as possible (~12 prior to zero fill)

(minimum slices: 10; lesion must fit be contained in the imaging volume, and at least 6 analyzable slices should be obtained)

5-8 mm slice thickness with slice interpolation

256 x 160 acquisition matrix

  - Ratio Map Protocol:

All parameters the same as for dynamic protocol except:

Single acquisition phase

15 degree flip angle

8 NEX (averages)

Scan time: ~1:04 min

Repeat second time with body receive coil.

  - Variable Flip Angle T1 Mapping Protocol:

All parameters the same as for dynamic protocol except:

Single acquisition phase

For GE scanners: Set Turbo Mode to 0 (under User CVs). This will lengthen TE and TR

(and scan time), but will force the TE and TR values to be the same for all flip angles.

2, 5, 10, 15, 20, 25, 30 deg flip angles, ordered from high to low.

4 NEX (averages)

0:43 min/flip angle (on GE scanner using Turbo Mode of 0; TE=1.04 ms, TR=5.13 ms)

Ensure that auto prescan is not performed between each flip angle acquisition so that system gain settings are identical for each flip angle acquisition. Auto-prescan should only be performed at the beginning of the series.

- Image Post-processing   Sandeep

There are no specific image post-processing requirements in this profile. No user-selected post-processing filters or image normalization methods should be used prior to data analysis as described in the next steps. If phased-array receiver coils are used, image combining and reconstruction should be according to standard manufacturer algorithms.

- Image Analysis    Ed Ashton

Analysis of DCE-MRI data is carried out in a series of distinct steps:

- Generate a ratio image using Body and Phased Array coil images.

- Generate a T1 map using Variable Flip Angle Images.

- Apply time-series motion correction to the dynamic data.

- Correct signal intensities in the dynamic data using the ratio image.

- Convert SI(t) in the dynamic data to [Gd](t).

  - Method 1:  Use a signal formation model in conjunction with the previously calculated T1 map.

  - Method 2:  Use a look-up table generated from gadolinium concentration phantom data.

- Calculate an arterial input function.

- Identify the region or regions of interest in the dynamic data.

- Calculate vascular parameters.

Each of these steps is addressed in detail below.

## 9.1 Generate a Ratio Image

The intent of this step is to optimize the information available for calculation of vascular parameters by combining the signal uniformity available using the inherent body coil with the improved SNR available using a torso phased array coil or similar.  The required inputs for this process are two images at each slice location - one acquired with the inherent body coil, and another acquired with the selected phased array (PA) coil.  These images should be prescribed identically to the dynamic series, and should be acquired immediately prior to the dynamic series, using the protocol as detailed in Section ****.  The output is an image in which each pixel contains a scale factor which, when applied to the dynamic data (which will be acquired using the PA coil) will ideally be sufficient to correct any non-uniformity in signal intensity due to the coil.  The ratio image is calculated as follows:

- Normalize the range of the two data sets IB and IPA.

  - Identify the minimum value in each image.

  - Subtract the minimum value from each image.

  - Identify the maximum value in each image.

  - Multiply IB by the ratio maxPA / maxB.

- Calculate the ratio image IR=IB/IPA.

- Reduce noise due to low signal regions (e.g. air space) and poorly registered regional transitions.

  - Set the ratio = 1.0 at all points where the signal value in either IB or IPA is less than 10.

  - Process the resulting IR using a median filter with support 21x21.

## 9.2 Generate a T1 Map

The intent of this step is to provide a complete map of pre-contrast T1 values for the imaged slab.  These values will then be used in the signal formation model based conversion of changes in signal intensity to gadolinium concentration.  The required inputs for this process are two or more images acquired with different flip angles, preferably covering a broad range above and below the Ernst angles for the acquired TR and range of T1 values expected in the human body. These images should be prescribed identically to the dynamic series, and should be acquired immediately prior to the dynamic series.  The output of this step is an image of T1 values which can be co-registered to the dynamic series and used in subsequent calculations.  The T1 values at each voxel location are calculated as follows:

- Create a vector x containing the signal intensity at each flip angle divided by the tangent of the flip angle.

- Create a vector y containing the signal intensity at each flip angle divided by the sine of the flip angle.

- For the n acquired flip angles create a set of points (x0,y0)… (xn,yn).

- Fit a line with slope s to the set of points defined in Step 3.

- T1 = -TR/ln(s).

## 9.3 Apply Motion Correction to the Dynamic Data

The intent of this step is to correct for patient motion that occurs between acquired phases of the dynamic data due to respiration, swallowing, and other involuntary movements.  This step is not intended to correct ghosting artifacts that can occur along the phase encoding direction within a particular image due to patient motion during acquisition.  These artifacts are more or less intractable unless the motion is regular and easily modeled, and are best addressed by adjusting the phase/frequency encoding scheme to minimize their impact on structures of interest.

In general, simple rigid shift or affine transform based registration methods will not be adequate for this step, due to the fact that the movement in question is typically limited to specific regions within the image – for example, the liver in a coronal scan of the abdomen may move substantially with respiration while the bulk of the body remains relatively motionless.  Fully deformable registration methods based on optical flow may provide good results in some cases [1, 2].  However, these methods will frequently fail for the phases immediately surrounding the contrast injection.  Semi-automated registration in which a user identifies the target tumor and only information drawn from that region is used to generate phase to phase shifts provides an alternative approach.  This allows rigid shift methods using mutual information [3], which tend to be more robust than optical flow methods, to be employed.  Finally, registration may be carried out manually.

## 9.4 Correct Signal Intensities Using the Ratio Image

The intent of this step is to correct for signal intensity variations introduced by non-uniform spatial sensitivity in the phased array coil.  This step is carried out by multiplying the signal value at each pixel in each image in the dynamic data by the value in the corresponding pixel in the ratio image calculated in 9.1.

## 9.5 Convert SI(t) in the Dynamic Data to [Gd](t)

The intent of this step is to convert the arbitrary signal intensity units in the dynamic data into units of gadolinium concentration.  This step should be applied after the regions of interest for analysis have been defined, but prior to the calculation of vascular parameters.  Two methods for accomplishing this are defined below.

### 9.5.1 Conversion Using a Signal Formation Model

Gadolinium concentration at each image pixel is given by:

In this equation, T10 is taken from the T1 map calculated in 9.2.  T1 is given by:

where TR is the repetition time for the dynamic series, α is the flip angle for the dynamic series, and A and B are given by:

### 9.5.2 Conversion Using a Look-Up Table

This method requires that a phantom containing a range of concentrations of gadolinium and a range of baseline T1 values (generally obtained via different concentrations of copper sulfate or a similar compound) is scanned using the dynamic protocol on each scanner that will be used for the study.  Data from these phantoms can then be used to construct a look-up table relating baseline T1, signal delta, and gadolinium concentration.  This look-up table can then be used during the analysis process to directly convert signal delta to gadolinium concentration prior to parameter calculation.

## 9.6 Calculate an Arterial Input Function

The intent of this step is to generate an accurate, patient-specific arterial input function to serve as an input to the vascular model.  One way to accomplish this is to have an analyst draw a manual ROI within an artery, and use the mean enhancement curve within that ROI as the subject-specific AIF, as described by Vonken et al [4].  It has been demonstrated previously that this method has significant variability associated with it [5], due primarily to the spatially- and temporally-varying flow artifacts found in major arteries.  A better option is to make use of an automated search technique to generate a locally optimal AIF.  Several methods of accomplishing this have been described previously [5, 6].

## 9.7Identify Regions of Interest in the Dynamic Data

The intent of this step is to define the boundaries of the tumors of interest in the dynamic data in order to determine which voxels will contribute to the parameter set for each tumor.  This step will generally be carried out by a trained and qualified radiologist using either manual tracing or a semi-automated boundary identification tool [5].

## 9.8 Calculate the Vascular Parameters

The intent of this step is to generate the parameter set which will be used to characterize the tissues of interest.  Parameters will be calculated based on the standard Tofts model [7], which is derived from the Kety equations [8].  The vascular bed is modeled as a linear system, such that:

with impulse response h(t) given by:

where Ktrans is the volume rate constant between blood plasma and extra-cellular extra-vascular space (EES) and kep is the rate constant between the EES and blood plasma.  Given the tissue uptake curve Ct(t)  and the AIF Cp(t)  ,  Ktrans and  kep are estimated using a gradient-descent energy minimization scheme.

A full parameter set will be calculated for each voxel within the defined tumor boundaries.  Parameters may be reported out either as mean and median values per tumor or as histograms.

## References

- Horn B, Schunck B, Determining optical flow.  Artificial Intelligence, 1981.

- Sharp G, Kandasamy N, Singh H, Folkert M.  Gpu-based streaming architectures for fast cone-beam CT image reconstruction and demons deformable registration.  Phys Med Biol, 2007.

- Pluim J, Maintz J, Viergever M.  Mutual Information Based Registration of Medical Images:  A Survey.  IEEE Trans on Medical Imaging, 2003.

- Vonken E, Osch M, Bakker C, Viergever M.  Measurement of cerebral perfusion with dual-echo multi-slice quantitative dynamic susceptibility contrast MRI.  J Magn Reson Imaging, 1999.

- Ashton E, McShane T, Evelhoch J.  Inter-operator variability in perfusion assessment of tumors in MRI using automated AIF detection.  Lect Notes Comput Sci 3749, 2005.

- Rijpkema M, Kaanders J, Joosten F, van der Kogel A, Heerschap A.  Method for quantitative mapping of dynamic MRI contrast agent enhancement in human tumors.  J Magn Reson Imaging, 2001.

- Tofts P, Kermonde A.  Measurement of the blood-brain barrier permeability and leakage space using dynamic MR imaging.  1. Fundamental Concepts.  Magn Reson Med, 1991.

- Kety S.  Peripheral blood flow measurement.  In: Methods in Medical Research, vol 8.  Potter VR (Ed.), Year Book Medical Publishers, 1960.

- Image Interpretation    Michael Knopp.

Tumor ROI definition

Discard end slices

Statistics to calculate from resulting maps

Single operator

- Archival and Distribution of Data      Sandeep

Archival and data distribution procedures are recommended so that all analysis results can be recomputed for verification and validation purposes. In addition to saving of all original images in DICOM formats, the following information must be archived along with the image data:

- Detailed specification of the input function selection. This may include a binary mask of pixels selected for arterial input function, or may consist of a tabulated text file containing RAS coordinates (or x,y,z) co-ordinates of the AIF pixel locations.

- Saving all parameters and users inputs required for image registration if used. Time-series image registration may be used to align data spatially over time. Any parameters which control the performance of the registration algorithm (metric used, optimization parameters, user click points, etc) must be stored in suitable format.  It is preferred to save the registration transform parameters so that identical registration can be reproduced in a multi-center environment.

- All regions of interest where analysis is performed and statistics are performed should be saved.

When data is distributed, it should be properly anonymized according to HIPAA regulations and data must only be distributed after meeting all local IRB specified informed consent protocols.

- Quality Control          Mark Rosen

List all sources of artifact and variation and procedures to mitigate them.

Prospective:

Contrast delivery

Subject motion

Acq plane

Image artifacts (wrap, metal, etc)

Retrospective:

Adherence to imaging protocols

Registration methods

AIF methods (variability in peak)

Patient and coil positioning

Using reference to recent pre-DCE-MRI imaging (usually CT study), tumor burden and location assessment should be made.  Optimally, review of actual imaging by a radiologist involved in the DCE-MRI study planning should be made.  Based on this review, putative target tumor lesion(s) should be identified prior to the baseline DCE-MRI study, and the preferred anatomic region for DCE-MRI (chest, abdomen, pelvis, extremity) should be made.    Methods for selecting target lesions are defined in section XX above.

DCE-MRI subject should be placed supine in the scanner position.  Rarely, prone positioning can be used for patients who profess a preference was such positioning to ease their concerns with claustrophobia.  More preferably, oral anxiolytics can be used for such patients prior to each DCE-MRI study.

When patient condition allows, placement of arms over the head may allow for avoidance of undesirable wrap artifact for temporally optimized 3D spoiled gradient echo sequences.  However, such positioning often cannot be sustained by patients without excessive discomfort.  In such cases, arms placed anteriorly over the cest or at the sides may be preferable.  For larger patients, side-down arm positioning may require adjustment of DCE-MRI imaging FOV to avoid undesirable wrap artifact.

Torso array coil placement in the area of anatomic target lesion(s) is then performed.  Ideally, both anterior and posterior coils are centered over the expected target lesion location.

Tumor selection  and Image QC prior to DCE-MRI sequence initiation

Axial imaging pre-gadolinium should be performed using available sequences per instiutional routine.  Imaging in end-expiratory breath held positioning is encouraged, especially when imaging the chest or upper abdomen.  Both T1 and T2 imaging should be performed and inspected to identify putative target lesions that may be unacceptable for DCE-MRI due to the presence or large areas which previously unidentified cystic or hemorrhagic areas.

Once the target lesion is identified, the DCE-MRI volume (oblique coronal or oblique sagittal slab) should be prescribed, using the axial T1W or T2W slice through the center of the target lesion, with the center of the slab bisecting both the target tumor and the representative artery (usually the aorta).  It is recommended for central laboratory quality assessment that the grid lines representing each slice of the DCE-MRI 3D slab be overlaid on this representative pre-Gad T1 or T2 image, and saved as a separate series for exportation with the remainder of the DCE-MRI study.  This overlay image should serve as a guide for reproducible placement of the DCE-MRI 3D slab at subsequent patient visits.

A representative 3D SPGR sequence is then obtained using the local torso array coils.  Breath-hold (end-expiration) imaging is recommended to minimize the artifact from this image set.  This sequence should be inspected for additional artifact (especially phase wrap) to identify whether such artifact occur in areas that overlap either the target tumor lesion or the arterial vessel.   As the image contrast in this sequence is often low, an anatomic expert should be available to aid the technologist in this assessment.  If unacceptable wrap is seen, adjustments to the imaging protocol (increase in the FOV, decrease in the degree of rectangular phase reduction—if allowed without unacceptable increase in temporal resolution) should be undertaken with repeat imaging if the initial DCE-MRI slab imaging is seemed unacceptable.  Patient repositioning may also be attempted at this point if deemed necessary.  If, despite all effort, acceptable DCE-MRI 3D slab image set cannot be obtained, then an alternate target tumor should be selected, if possible.

Centralized post acquisition DCE-MRI quality control and assessment

In practice, DCE-MRI image sets must be assessed for quality before quantitative analyses are performed.  While it is beyond the scope of this profile to detail estimates of DCE-MRI technical “success” in clinical practice or in clinical trial settings, it is recognized that within a clinical trial cohort, a fraction of DCE-MRI data sets may be considered suboptimal due to errors in patient selection, DCE-MRI exam prescription, magnet performance during the DCE-MRI study, or other factors, any or all of which may lead to spurious errors in quantification of tumor Ktrans or IAUCG.  Recognition of these factors will help to mitigate reporting of spurious DCE-MRI changes in patients, and will enable further evaluation of results of DCE-MRI technical assessment.

In both technical assessment and early phase clinical trials, censuring of data for reporting should be discouraged, as this may lead to overt or inadvertent biases in the ensuing results.  However, it may be appropriate to assess individual data sets for overall technical success.  In such cases, in vivo results—be they for reproducibility assessment or reponse to therapeutic interventions—may be reported for both the entire cohort (analogous to an intention-to-treat reporting) and for the subset of subject with technically “adequate” studies.  The remainder of this section refers to methods for evaluating the technical adequacy of data sets.

- Confirming patient/lesion suitability for DCE-MRI

As discussed in section XX above, patients suitable for DCE-MRI analysis must possess at least one tumor of sufficient size.  This requires that there be at least one tumor (designated the DCE-MRI target lesion) be >=2cm in longest diameter in anatomically fixed locations (defined in section XX) or >=3cm or greater in anatomically non-fixed areas.  Furthermore, the target lesion for DCE-MRI analysis should be well removed from areas subject to large degrees of cardiac pulsatility artifact (e.g. hilar/subhilar lung regions, lateral left lobe of the liver), as these artifacts may yield unacceptable signal intensity fluctuations, rendering DCE-MRI quantitation of questionable validity.

Depending on the specifications of individual protocols, use of target lesions that violate specified acceptable target lesion attributes should be rejected.   In rare instances, a second non-target lesion meeting requirements for targetlesion status may be visible within the analyzable DCE-MRI volume.  In such cases, these lesions may be analyzed as a substitute for the intended DCE-MRI target lesion.

- Confirming fidelity of DCE-MRI exam prescription

The DCE-MRI exam includes a set of images used to quantify baseline (pre-gadolinium) tumor R1 (1/T1) values (termed T1 mapping series), a second set used for removal of local coil influences and a dynamic gadolinium enhanced series.  These series must adhere to required prescription constraints in order for DCE-MRI quantification to be valid.

Unless specifically indicated (e.g. flip angle variation in a T1-mapping series), all imaging parameters (e.g. TR, TE) and the slab imaging geometry must be held constant thoughout the DCE-MRI series.  Gross deviations from accepted parameter values or geometric prescription will cause the quantification to be invalid.  Furthermore, automatic gain recalibration must be disabled to ensure that signal intensities are scaled similarly throughout the exam prescription.

Dynamic enhanced imaging series must also be assessed for adequate quality.  The dynamic enhanced series must be run with adequate temporal resolution (discussed in section XX) to ensure accurate quantification of the arterial input function and tumor enhancement curve.  Baseline (pre-gadolinium) imaging of sufficient duration must be obtained to ensure that the correct baseline tumor signal intensity is obtained.  The total scan duration must also be adequate to ensure sampling of the tumor enhancement curve.

- Fidelity of contrast administration

During the dynamic enhanced imaging, the gadolinium bolus must be adequately administered.  Bolus administration is required to for adequate dynamic range of contrast enhancement.

- Fidelity of DCE-MRI slab placement

The ideally placed slab should bisect the target tumor and the representative arterial vessel used to evaluate the arterial input function (usually the descending thoracic or abdominal aorta).  In the event that the center slice of the slab does not bisect these two structures, the slab must be wide enough such that the center of both the target tumor and the aorta do not lie at the last 1-2 edge slices of the slab.  The number of acceptable slices will therefore depend on the number of slices in the slab prescription.  If 12 slices are prescribed, then the tumor and the aorta must be identified between slices 3 and 10.

- Reproducibility of DCE-MRI study prescription

Repeat DCE-MRI examination, whether for determination of test-retest reproducibility of DCE-MRI quantification or for evaluation of change in tumor DCE-MRI characteristics in the setting of an anti-tumor therapy, must be performed with strict adherence to protocol characteristics as defined by the initial imaging session.  The following parameters should be assessed in evaluating the adequacy of the DCE-MRI study

- Patient positioning should be similar between the two sessions.  Prone positioning for alleviation of claustrophobia should not be undertaken unless such positioning was performed at the initial session.  Arm positioning (by sides or above the head) should adhere to original positioning.  Number of and placement of local receive coils should  be stable between DCE-MRI studies.

- Patients should be land-marked in a near-identical manner as performed for the initial scan.  Stability of patient landmarking should be assessed by evaluation of the location of a specific fixed area of anatomy (such as a certain spinal level or other rigid osseous landmark).  The z-position of this landmark should not deviate by more than XX mm between the two (or more) examinations.

- Study parameters should not have varied between the two sessions. Specifically, the TR, flip angle(s), and DCE-MRI slab acquisition parameters should be identical.  When oblique slab prescription is utilized, the obliquity of the slab prescription should be as close as possible to that used in the original study.  At best, the obliquity of the slab plane should not deviate by more than XX degrees from the prior placement.

- Reproducibility of slab geometry placement by the technologist should be assessed via evaluation of the DCE-MRI image sets.  Slice location overlays on representative axial images should be used to determine if the 3D oblique slab was positioned reproducibly.  While slight variations in slab obliquity are expected and acceptable, the relationship between the slab and the center of the target lesion and artery should be relatively constant for each DCE-MRI session.  In addition, z-centering of the slab should be constant between studies, as gross variation in z-positing of the center of the slab relative to the tumor and artery may lead to unacceptable deviations in signal intensities as portions of the image extend toward the ends of the magnet bore.  For larger bore magnet systems, the tumor center should be within XX mm of isocenter in the z position.  For short bore magents, the postion of the tumor in relation to isocenter should be XX mm.

- Assessing image artifacts during DCE-MRI study

As the DCE-MRI study optimizes contrast and temporal resolution, volumetric coverage during the DCE-MRI study is minimized.  Such MRI prescriptions may lead to image artifacts, including wrap (aliasing).  These artifacts will be relatively constant during the MRI study.  Such artifacts are acceptable as long as they do not interfere with he evaluation of signal intensity for the target DCE-MRI lesion and/or the arterial vessel.  Fixed aliasing artifacts across one or both of these structures will render the DCE-MRI image set non-analyzable

Motion artifacts are inevitable in the DCE-MRI studies.  As the dynamic enhanced study is performed during free respiration, tumors located in the chest and upper abdomen will undergo motion during the study, and as such, respiratory artifacts will be present.  However, the degree of such artifacts should be such that they do not obscure the tumor or representative arterial vessel during the course of the dynamic enhanced imaging.  If individual timepoints demonstrate such artifacts, they should be discarded from the dynamic enhanced series.

Additional artifacts, due to magnet instability or random electronic noise may also be seen during the course of DCE-MRI imaging studies.  The DCE-MRI series should be assessed for such artifacts.    Gross patient motion artifacts may also be encouneteed in the course of a DCE-MRI study.  Such artifacts arise when the patient undergoes voluntary or involuntary shifting of position during the DCE-MRI study.  All DCE-MRI image sets should be assessed for evidence of such motion.  If gross motion within or between DCE-MRI series is noted, the DCE-MRI series should be rejected as non-analyzable.

One method for assessing for random magent instability or gross patient artifacts is to evaluate the dynamic signal intensity of a region of interest in non-enhancing adipose tissue in the posterior aspect of the patient, as this area of anatomy should be relative motion-free during the course of the dynamic enhanced imaging.  If the signal intensity is seen to deviate by more than 10% of the mean signal intensity, then either random machine noise or gross patient motion should be suspected.  If the deviation is transient (e.g. noise “spike”) the time point or points affected may be discarded an the remainder of the DCE-MRI series evaluated.  If the deviation is fixed over a portion of the dynamic enhanced image set, this is usually a sign of gross patient motion. Such data sets should be rejected as non-analyzable.

- Monitoring and reporting quality assessment in DCE-MRI studies

As discussed earlier in this section, it is expected that within a given cohort, a subset of DCE-MRI studies may provide poor quantified data due to errors discussed above.  In general, protocols including DCE-MRI assessment should discuss specific methods for handling and reporting such data sets.  It is the opinion of the committee that no data sets should be deemed analyzable for quantitative DCE-MRI without explicit documentation of the reasons for such failure.  Furthermore, depending on the nature of the study endpoints, DCE-MRI analysis should proceed when feasible for all such data sets, regardless of image QC measures.  In such cases, result reporting my include cohort-wide and subset analysis, the latter including only those data sets that meet pre-specified quality control standards.  At a minimum, the number of DCE-MRI cases and patients excluded from analysis, and the reasons for such exclusions, should be specified in all result reporting.

- Imaging-associated Risks and Risk Management    Sandeep

Standard MRI and Gd-based contrast agent contra-indications

APPENDICES

- Acknowledgements and Attributions

- Background Information

- Conventions and Definitions

- Documents included in the imaging protocol (e.g., CRFs)

- Associated Documents (derived from the imaging protocol or supportive of the imaging protocol)

- TBD

- Model-specific Instructions and Parameters

GE prescan details

IV. COMPLIANCE SECTION

Testability/Test plan;

Phantom data sets; simulation images; clinical test data

V. ACKNOWLEDGEMENTS

# Comments

Are there are other highly cited review papers that we should include here ?

Reword this sentence and make it compatible with the wording of claims in CT profile for example.

What is this?

Define more:  Actual DCE-MRI experience, or experience more generally in body MRI including clinical multi-phase contrast enhanced imaging?  More rigorous requirements severely limits DCE-MRI capable sites currently.

This needs more details on exactly which phantom, and the phantom imaging experiment details. These are not included later in the protocols section as that describes the clinical protocol.

Numerical definitions for “passing” phantom imaging?

I have used the following mock DCE-MRI in volunteers:  Use gallbladder as target “tumor, prescribe oblique coronal slab through GB center and aorta.

Ennumerate attributes of volunteer scan to assess?

Please check the organization/scope of this section. Does the Claim need to be repeated here ?

For pts on clinical trials in the metastatic setting, CT CAP is pretty standard, and is ideal for judging tumor burden, distribution, and suitability for DCE-MRI, as well as planning target lesions region for DCE-MRI set-up.

Define minimum.  Usually no more than 28-30 days for standard oncology trials.  Arguably, two weeks better for functional assays such as DCE-MRI?

ECOG status as ana added issue.  Pts who care ECOG 2 often cannot lie still in magnet.  ECOG 3 can’t make it in to the study.  In my opinion, OK to be a bit more stringent than usual contraindications for MRI in general

Other electronic devices (spinal stimulator

We include non-MRI compatible metal anywhere in CNS

Don’t we also have exclusion criterion based on type/size/location of tumor ? Also, patient should not have had another contrast enhanced MRI exam within the previous XX hours.

More recent FDA guidelines published.

Not sure this rises to the level of ideal/target/acceptable UPICT specs

Mark’s point below is very important. We should exclude such cases.

Covered by excluding ECOG >=2?

I have strong feelings here that we overstate the importance of high injection rate and minimize the downsides.

I do not necessarily disagree for the purposes of defining the ideal test-retest cohort, but I point out that from my experience, this eliminates ~10% of typical phase I-II patients.  Also, many PICC lines (though not ports) are rated for injection up to 5 cc/sec.

Objection!

We have to explicitly state the timing of the injection relative to the start of DCEMRI imaging. This is vital to getting some clean baseline pre-contrast images. We need to specify that contrast injection needs to occur after other anatomic and VFA sequence, and ratio map etc have been acquired.

I thought if ratio image is used, then T1 maps are not needed. If T1 maps are used, then ratio correction is not needed. We seem to be mixing both.

To yield new IB ?

In my opinion, motion correction will have a huge impact on quantification. We need to be more definitive on what we are recommending here, and also some means of QC to check whether data is adequately free of motion after correction.

Are we doing this regardless of whether T1 maps are used or not?

Expression for E10

How about conversion using known fixed T1 value of tissue?

We need to recommend an approach and explain the algorithm here. How about population averaged AIF and/or reference tissue AIF  methods?

Part of section 4 or 7?

Include common features:  minimum size, location, non-cystic, away from pulsatile structures.  Guidance for methods to identify ideal target lesion when several potential targets are available?

Part of Section 7?

Arguably, this is what goes in section 12, but some of this may need to come before current section 9.

Caveats for SAR-required reduction in flip or increases in TR?

At least during the T1 mapping series?

Do we have a preferred minimum?

Required minium post gad imaging?

Potentially controversial.  Worthy of group discussion