QIBA Profile: FDG-PET/CT as an Imaging Biomarker Measuring Response to Cancer Therapy

Version 0.7.3

May 31, 2012

Table of Contents

# I. Executive Summary	3

# II. Clinical Context and Claims	4

# Claim: 	5

# III. Profile Details	6

# 1. Subject Handling	7

# 2. Image Acquisition	8

# 3. Image Data Reconstruction & Post-processing   	18

# 4. Image Analysis	19

# 5. Image Interpretation Reporting	25

# 6. Quality Control	26

# IV. Compliance	25

# 7. Image Acquisition Site	24

# 8. Acquisition Device	25

# 9. Reconstruction Software	28

# 10. Analysis Software	31

# 11. Version tracking	32

# References	35

# Appendices	35

# Appendix A: Acknowledgements and Attributions	35

# Appendix C: Conventions and Definitions	35

# Appendix G: Model-specific Instructions and Parameters	36

# Definitions

# QIBA:  Quantitative Imaging Biomarkers Alliance. The Quantitative Imaging Biomarkers Alliance (QIBA) was organized by RSNA in 2007 to unite researchers, healthcare professionals and industry stakeholders in the advancement of quantitative imaging and the use of biomarkers in clinical trials and practice.

# Profile:

# RSNA:

# FDG:

# PET:

# PET/CT:

# UPICT:

# SNM:

# AAPM:

# EANM:

# SUV: (be specific about technical definition here)

SUVmax etc.

# ROI:

DICOM:

ACR:

# Profile requirements

Illustrative example:

Parameter	Entity/Actor	Normative text: Clear boxes are current requirements

Shaded boxes are intended for future requirements

| Lesion Analysis: Multiple Voxels | Analysis Tool | Shall provide tools to measure and report SUV mean and max normalized to body weight |
| --- | --- | --- |
|  |  | Shall provide tools to measure and report SUV mean SUV max and SUV peak, normalized to body weight and/or lean body mass |

# Above table items are normative (i.e. required in order to be compliant with the QIBA protocol). The intent of the normative text is to be prescriptive and detailed to facilitate implementation. In general the intent is to specific the final state or output, and not how that is to be achieved.

All other text is considered informative only

# I. Executive Summary

This QIBA Profile documents specifications and requirements to provide comparability and consistency for quantitative FDG-PET across scanners in oncology. It can be applied to both clinical trial use as well as individual patient management. The document, developed through the efforts of the QIBA FDG-PET Technical Subcommittee, has shared content with the FDG-PET UPICT protocol, as well as additional material focused on the devices used to acquire and analyze the FDG-PET data.

Figure: Illustration of the Profile components

Summary for Clinical Trial Use

The QIBA FDG-PET/CT Profile defines the technical and behavioral performance levels and quality control specifications for whole-body FDG-PET/CT scans used in single- and multi-center clinical trials of oncologic therapies. While the emphasis is on clinical trials, this process is also intended to apply for clinical practice. The specific claims for accuracy are detailed below in the Claims.

The specifications that must be met to achieve compliance with this Profile correspond to acceptable levels specified in the FDG-PET UPICT Protocol.  The aim of the QIBA Profile specifications is to minimize intra- and inter-subject, intra- and inter-platform, and inter-institutional variability of quantitative scan data due to factors other than the intervention under investigation. FDG-PET/CT study(ies) performed according to the technical specifications of this QIBA Profile in clinical trials can provide qualitative and/or quantitative data for single time point assessments (e.g., diagnosis, staging, eligibility assessment, investigation of predictive and/or prognostic biomarker(s)) and/or for multi-time point comparative assessments (e.g., response assessment, investigation of predictive and/or prognostic biomarkers of treatment efficacy).

# II. Clinical Context and Claims

FDG is a glucose analogue. The rationale for its use in oncology is based on the typically increased glucose metabolism of tumors compared to normal tissue.  FDG is transported into tumor cells via glucose transport proteins, usually up-regulated in tumor cells.  Once internalized FDG is phosphorylated to FDG-6-phosphate, which does not enter glycolysis and becomes metabolically trapped.  Uptake is not specific for tumor cells and there are some normal tissues that show high background uptake.

Applications and Endpoints for Clinical Trials

FDG-PET/CT imaging can be used for a wide range of clinical indications and research questions. These are addressed more completely in the FDG-PET/CT UPICT Protocol (UPICT section 1.1). This QIBA Profile specifically addresses the requirements for measurement of tumor FDG uptake with PET/CT as an imaging biomarker for evaluating therapeutic response.

Biomarkers useful in clinical research for patient stratification or evaluation of therapeutic response would be useful subsequently in clinical practice for the analogous purposes of initial choice of therapy and then individualization of therapeutic regimen based on the extent and degree of response as quantified by FDG-PET/CT.

The technical specifications described in the profile are appropriate for quantification of tumor FDG uptake and measuring longitudinal changes within subjects. However, many of the profile details are generally applicable to quantitative FDG-PET/CT imaging in other applications.

FDG-PET scans are sensitive and specific for detection of most malignant tumors. [Fletcher et al. Recommendations on the use of 18F-FDG PET in oncology. J Nucl Med (2008) vol. 49 (3) pp. 480-508]. Coverage for oncology imaging procedures in the US by the Centers for Medicare and Medicaid Services are explicitly listed in the National Coverage Determination (NCD) for Positron Emission Tomography (PET) Scans (220.6). FDG-PET scans reliably reflect glucose metabolic activity of cancers and this metabolic activity can be measured with high reproducibility over time. Longitudinal changes in tumor FDG activity during therapy predict clinical outcomes earlier than changes in standard anatomic measurements. Therefore, tumor metabolic response or progression as determined by tumor FDG uptake can serve as a pharmacodynamic endpoint in well-controlled Phase I and Phase IIA studies as well as an efficacy endpoint in Phase II and III studies.  In tumor/drug settings where the preceding phase II trials have shown a statistically significant relationship between FDG-PET response and an independent measure of outcome, changes in tumor FDG activity may serve as the primary efficacy endpoint for regulatory drug approval in registration trials.

## Claim:

Part 1. For single-center studies. If FDG-PET imaging is performed in accordance with the imaging protocol described in section III 'Profile Details' and while also using equipment, personnel, facilities, and other resources according to the technical specifications listed in section IV 'Compliance Specifications', then

Tumor glycolytic activity as reflected by the standard uptake value (SUV) can be measured from FDG-PET/CT scans with a ≤ 30% test-retest coefficient of repeatability, with a 95% limit of agreement, for solid tumors at least 2 cm in diameter and with a minimum baseline SUV of 1.5.

Part 2. For multi-center studies. If FDG-PET imaging is performed as described above, and uniformly across imaging sites, then it is anticipated that the same claims in Part 1 can be met.

Part 3. For single-center studies changes in SUV (defined as ...) greater that XX% are representative of true biological change, with a with a 95% confidence interval, for solid tumors at least 2 cm in diameter and with a minimum baseline SUV of X.X

# III. Profile Details

The following figure provides a graphical depiction that describes the marker at a technical level.

Figure 1: The assay method for computing and interpreting glycolytic metabolic activity using PET/CT may be viewed as a pipeline using either one or two or more scan sequences.

- Patients or subjects are prepared for scanning (e.g. 6 hr fasting). FDG is administered. Patient waits quietly for bio-distribution and uptake of FDG (e.g. 60 min)

- Scan data from the PET and CT exams is acquired.

- Data correction terms are estimated and PET (and CT) images are reconstructed.

- Quantitative measurements are performed.

- Images are reviewed for qualitative interpretation.

Note that steps 4 and 5 may occur in either order or at the same time.

Images may be obtained at a multiple time points, notably at a minimum of two time points before and after therapeutic intervention for a response assessment as is considered by this document. The change is typically assessed as a percentage according to the formula: (post-treatment metabolic activity – pre-treatment metabolic activity) / pre-treatment metabolic activity.  Response criteria are then applied to categorize the response assessment (e.g. CMR, PMR, SMD, PMD).  These response criteria are beyond the scope of this document.

The following sections provide details describing various issues, parameters and specifications for each issue or parameter required for compliance:

| Section | Title | Performed by |
| --- | --- | --- |
| 1 | Subject Handling | Personnel (including Technologists and Schedulers) at an Image Acquisition Facility |
| 2 | Image Data Acquisition | Technologist at an Image Acquisition Facility using an Acquisition Device |
| 3 | Image Data Reconstruction | Technologist at an Image Acquisition Facility using Reconstruction Software |
| 4 | Image Analysis | Imaging Physician (or supervision of Image Analyst) using one or more Analysis Software tools |
| 5 | Image Interpretation | Imaging Physician based on the synthesis of information obtained by Image Analysis using a pre-defined Response Assessment Criteria |

This document organizes acquisition, reconstruction and post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be a computational step that transforms the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge.

Relation of this Profile to Expectations for QIBA Profile Compliance

The requirements included here are intended to establish a baseline level of capabilities.  Providing higher levels of performance or advanced capabilities is both allowed and encouraged.  Furthermore the QIBA Profile is not intended to limit equipment suppliers in any way with respect to how they meet these requirements. Institutions meeting the stated criteria are considered to be QIBA Compliant.

## 1. Subject Handling

## 1.1  Subject Scheduling

The study should include specific directions as to the management of subjects with abnormal fasting blood glucose measurements whether known to be diabetic or not.  While it is known the high levels of circulating blood glucose and impact the measurement of FDG uptake, there is a paucity of scientific data to suggest the critical cutoff for abnormally high blood glucose measurements and if these patients should be excluded from clinical trials that use FDG-PET/CT scan data. it is important to define how such subjects and the data from their imaging studies are managed to ensure comparability of imaging data within and among clinical trials.  Specifically, consideration should be given to the exclusion of subjects with abnormal fasting blood glucose when quantitative FDG-PET/CT is being used as the study’s primary endpoint. Reference FDG-PET/CT UPICT Protocol for Diabetic Scheduling and Management discussion (UPICT Section 4.2.2).

1.1.1 Timing of Imaging Test Relative to Intervention Activity (UPICT Section 1.2)

The study protocol should specifically define an acceptable time interval that should separate the performance of the FDG-PET/CT scan from both (1) the index intervention and (2) other interventions (e.g. chemotherapy, radiotherapy or prior treatment).  This scan (or time point) is referred to as the “baseline” scan (or time point). The time interval between the baseline scan and the initiation of treatment should be specified as well as the time intervals between subsequent FDG-PET studies and cycles of treatment.  Additionally, the study protocol should specifically define an acceptable timing variance for performance of FDG-PET/CT around each time point at which imaging is specified (i.e., the acceptable window of time during which the imaging may be obtained “on schedule”).  The timing interval and window are dependent upon 1) the utility for the FDG-PET/CT imaging within the clinical trial, 2) the clinical question that is being investigated and 3) the specific intervention under investigation.  Suggested parameters for timing of FDG-PET/CT within oncologic trials are more completely addressed in the FDG-PET/CT UPICT Protocol section 1.2.

1.1.2. Timing Relative to Confounding Activities (UPICT Section 3.2)

Activities, tests and interventions that might increase the chance for false positive and/or false negative FDG-PET/CT studies should be avoided prior to scanning. The allowable interval between the potentially confounding event and the FDG-PET/CT exam will be dependent on the nature of the confounder. For example, a percutaneous or excisional biopsy of a suspicious mass may cause focally increased FDG-PET activity or might lead to the appearance of a non-malignant mass (e.g., hematoma) on the CT portion of the study. A percutaneous ablation procedure of a known malignant focus may cause focally increased FDG-PET activity and/or an immediate post-ablation increase in the apparent volume of the ablation target lesion. The time of onset and the duration of the increased FDG-PET activity and/or the change in lesion volume might be different for these two different confounding factors.

If iodinated contrast is to be used for the CT portion of the PET/CT study, conflict with other tests and treatments should be avoided congruous with community standards of care (e.g., thyroid scan).

1.1.3. Timing Relative to Ancillary Testing (UPICT Section 3.3)

Avoid scheduling tests that might confound the qualitative or quantitative results of the FDG-PET/CT study within the time period prior to the scan. For example, a glucose tolerance test should not be scheduled during the 24 hours prior to the performance of FDG-PET/CT. Similarly, other tests that might involve increasing plasma glucose, insulin, or corticosteroid levels should also be avoided. Exercise cardiac stress testing should be avoided during the twenty-four (24) hours prior to the performance of FDG-PET/CT. Similarly, other tests that might involve vigorous exercise and thereby increase muscle metabolic function should also be avoided.

## 1.2 Subject Preparation (UPICT Section 4)

Management of the subject can be considered in terms of three distinct time intervals (1) prior to the imaging session (prior to arrival and upon arrival), (2) during the imaging session and (3) post imaging session completion.  The pre-imaging session issues are contained in this section while the intra-imaging issues are contained in section 2.1 on image data acquisition.

1.2.1. Prior to Arrival (UPICT Section 4.1)

The main purpose of subject preparation is to reduce tracer uptake in normal tissue (kidneys, bladder, skeletal muscle, myocardium, brown fat) while maintaining and optimizing tracer uptake in the target structures (tumor tissue). For more detail, refer to the FDG PET UPICT Protocol (Section 4.1) that addresses  (1) Dietary, (2) Fluid Intake, and (3) Other activities that may affect tissue FDG uptake.

- Dietary

  - Diabetic management – Refer to FDG-PET/CT UPICT Protocol sections 1.7.2 & 4.2.2

  - Fasting status - Subjects should not eat any food (either enteral or parenteral) for at least six  hours prior to the anticipated time of FDG administration.

- Fluid Intake: Adequate hydration (before and after FDG administration) is important both to ensure a sufficiently low FDG concentration in urine (fewer artifacts) and to reduce radiation exposure to subjects. Adequate hydration is especially important when contrast CT imaging will be used. Whichever hydration strategy is used (how much and when to administer), the protocol should be uniform among sites during a trial. Specific hydration recommendations are presented in the FDG-PET/CT UPICT Protocol (reference Section 4.2.1). The fluid administered should not contain glucose or caffeine.

- Other Activities: To minimize FDG uptake in muscle, the subject should avoid strenuous or extreme exercise before the PET exam for a minimum of at least 6 hours (preferably for a time period of 24 hours).

The compliance issues around these parameters are dependent upon adequate communication and oversight of the Scheduler or Technologist at the Image Acquisition Facility with the subject. Communication with the subject and confirmation of compliance should be documented.

1.2.2. Upon Arrival (UPICT Section 4.2)

Upon arrival 1) confirmation of subject compliance with pre-procedure instructions and 2) the occurrence of potentially confounding events (see listing in Section 4.2.1 of FDG-PET/CT UPICT Protocol) should be documented on the appropriate case report forms.

There should be documentation of subject-specific risk factors including, but not limited to, previous contrast reactions (if iodinated contrast is to be used).

1.2.3 Preparation for Exam (UPICT Section 4.2.3)

In order to avoid heterogeneous physiological distribution of the FDG, it is critical that subject preparation after arrival and prior to imaging is standardized among all sites and subjects throughout the conduct of the clinical trial.

- The waiting and preparation rooms should be relaxing and warm (> 75° F or 22° C) during the entire uptake period (and for as long as reasonably practicable prior to injection, at least 15 minutes is suggested as acceptable). Blankets should be provided if necessary.

- The subject should remain recumbent or may be comfortably seated; activity and conversation should be kept to an absolute minimum. For example, the subject should be asked to refrain from speaking, chewing, or reading during the uptake period. For brain imaging the subject should be in a room that is dimly lit and quiet for FDG administration and subsequent uptake period.

- The subject may use the toilet, but if possible not for the 30 minutes immediately after injection of FDG. The subject should void immediately (5 – 10 minutes) prior to the FDG-PET/CT image acquisition phase of the examination.

- Bladder catheterization is not routinely necessary; but if necessary the catheter should be placed prior to injection of FDG. Bladder catheterization may be important for the evaluation of pelvic tumors (e.g., cervix or prostate cancer).

- Following the administration of FDG, the subject should drink 500 ml of water (or receive by intravenous administration 250 - 500 ml of non-glucose containing fluid). Fluid intake may need to be modified for those subjects on fluid restriction.

- For specific areas of anatomic interest (e.g., tumors located in the lower abdomen, pelvis or kidney) intravenous diuretic agents may be used (e.g., 20 – 40 mg of furosemide given nearly contemporaneously (within 10 – 15 minutes) with the administration of FDG). If bladder catheterization is performed, IV diuretics should be administered as described here so as to ensure that the concentration of activity in the renal collecting systems and bladder is relatively dilute.

- Sedation is not routinely required, but is not contraindicated provided that the sedative used does not interfere with the uptake of FDG. Sedation may have utility in specific clinical circumstances such as brain or head and neck tumors, claustrophobic subjects, or children.

- The amount of fluid intake and use of all medications (e.g., diuretic, sedative) must be documented on the appropriate case report form.

- Subjects undergoing a CT scan should empty their pockets and remove any clothing containing metal and any metallic jewelry from the body parts to be scanned, changing into a hospital gown if necessary.

Parameter	Entity/Actor	Specification

| Height & Weight | Technologist | The Technologist shall measure and document subject height & weight and enter this information into the scanner during the PET/CT acquisition. Subject height and body weight shall be measured at the time of the baseline study with standardized measurement devices and with the subject inan examination gown or light clothing. If patient cannot be moved from the bed, the date and source of information should be documented. |
| --- | --- | --- |
|  |  | The Technologist shall measure subject height & weight and enter this information into a common data format mechanism used for recording all needed information (Appendix D). This information is then input into the scanner using manual or direct entry into the scanner during the PET/CT acquisition. |
|  |  | To be superseded by direct electronic entry to DICOM fields: See section 8 |

- Diabetic Monitoring and Management (UPICT Section 4.2.2)

The subject’s blood glucose level should be measured (using CLIA-approved, CLIA cleared, or equivalent (outside US) glucose measurement  device or laboratory) within the preceding 2 hours (ideally within 1 hour, especially in subjects with diabetes) of FDG administration and documented.

| Blood glucose level measurement | Technologist or Lab Technologist | Within 2 hours preceding FDG administration the Technologist or Lab Technologist shall measure and document time of subject blood glucose collection. Glucose measurement should be performed using a CLIA approved, CLIA cleared, or equivalent (outside US) glucose measurement device. Deviations from this process shall be documented. |
| --- | --- | --- |
| Blood glucose level documentation | Technologist or Lab Technologist | The Technologist or Lab Technologist shall enter the results of the blood glucose assay and the time of blood draw on a case report form or similar subject information sheet. |
|  |  | The Technologist or Lab Technologist shall enter the results of the blood glucose assay into a common format mechanism used for recording all needed information (Appendix D). |
|  |  | To be superseded by direct electronic entry to DICOM fields: See section 8. |
| Blood glucose level threshold | Technologist | The Technologist shall enforce the glucose thresholds for imaging as defined in the Protocol; if not, then reason for non-compliance shall be provided and documented on case report form or similar subject information sheet. |
|  |  | The Technologist shall document non-compliance information into a common format mechanism used for recording all needed information (Appendix D). |
|  |  | To be superseded by direct electronic entry to DICOM fields: See section 8. |

## 1.3. Imaging-related Substance Preparation and Administration  (UPICT Section 5)

1.3.1. Radiotracer Preparation and Administration

1.3.1.1  Radiotracer Description and Purpose

The FDG radiopharmaceutical must meet USP (or comparable International equivalent) specifications or meet other current specifications as defined by the FDA, EMEA or other appropriate regulatory agency approval.

1.3.1.2 Radiotracer Dose Calculation and/or Schedule (UPICT Section 5.2)

The administered 18F-FDG dose administered ranges between about 5 – 20 mCi (185 – 740MBq).  The administered dose depends upon multiple factors including but not limited to (1) country where exam is performed, (2) subject size and age (3) scanning mode and (4) percentage of scan bed (slice) overlap.  The exact dose and the time at which dose is calibrated should be recorded.  Residual dose remaining in the tubing, syringe or automated administration system or any dose spilled during injection should be recorded.

| Administered FDG Radiotracer Dose | Technologist | The Technologist shall Assay the pre-injection FDG dose (i.e. radioactivity) and time of measurement, Record the time that FDG was injected into the subject, Assay the residual dose in the syringe (and readily available tubing and components) after injection and record the time of measurement. These values shall be entered into the scanner during the PET/CT acquisition. For scanners that do not provide for entry of residual dose information, the net injected radioactivity should be manually calculated by decay correcting all measurements to the time of injection and then subtracting the residual radioactivity from the pre-injection radioactivity. The net injected radioactivity is then entered into the scanner during the PET/CT acquisition. All data shall be documented. |
| --- | --- | --- |
|  |  | All data should be entered into the common data format mechanism (Appendix D) This information is then input into the scanner using manual or direct entry into the scanner during the PET/CT acquisition. |
|  |  | To be superseded by direct electronic entry to DICOM fields: See section 8 |

1.3.1.3 Radiotracer Administration Route (UPICT Section 5.4)

FDG should be administered intravenously through a large bore (≥21 gauge) indwelling catheter placed anatomically remote (e.g., contralateral extremity to site of disease if at all possible) to any site(s) of suspected pathology, preferably in an antecubital vein. Intravenous ports should not be used, unless no other venous access is available. If a port is used addition flush volume should be used. As reproducible and correct administration of FDG is required for quantification purposes, extravasation or paravenous administration should be avoided.  If an infiltration is suspected, the event and expected quantity should be recorded and the infiltration site should be imaged. The approximate amount of infiltration should be estimated from the images where possible. If the infiltration is greater than 5% of the administered dose and the quantitative result from the FDG-PET/CT study is a primary or secondary endpoint, the data point might be censored from review or the subject might not be included in the study. The injection site should be documented on the appropriate case report form.

Presuming that the IV access site is properly functioning, the same route of administration may be used for iodinated contrast as is used for FDG.

| FDG Administration | Technologist | Technologist shall administer FDG intravenously through a large bore (≥21 gauge) indwelling catheter placed anatomically remote to any sites of suspected pathology, preferably in an antecubital vein. Intravenous ports should not be used, unless no other venous access is available. In the case of manual administration, a three-way valve system should be attached to the intravenous cannula so as to allow at least a 10 cc normal (0.9% NaCl) saline flush following FDG injection. |
| --- | --- | --- |
| Suspected infiltration | Technologist | Technologist shall 1. Record the event and expected amount of FDG: [Minor (estimated less than 5%), Moderate (estimated more than 5% less and than 20%), Severe (estimated more than 20%)] 2. Image the infiltration site. |
|  |  | Record the event and expected amount of FDG into the common data format mechanism (Appendix D) |

1.3.2 Contrast Material Preparation and Administration

The use of contrast material during FDG-PET/CT imaging is complex and analyzed in detail in the QIBA-UPICT FDG-PET in Section III.2."Image Data Acquisition for specific scanning strategies". In summary, the presence of iv and/or oral contrast material improves the detection of lesions with CT and may improve the anatomic localization, interpretation, and analysis of the FDG-PET/CT exam. However, the presence of contrast material may affect the attenuation correction applied for the PET scan with consequent bias in measured SUVs.

| CT Contrast Agent | Technologist | Technologist shall record the type and amount of CT Contrast Agent. 1. Was oral contrast used: Type [Positive, Negative], amount (volume in cc). 2. Was IV contrast used?, amount (volume in cc), time of injection. |
| --- | --- | --- |
|  |  | Record the event and expected amount of CT Contrast Agent into the common data format mechanism (Appendix D). |

## 2. Image Data Acquisition

This section summarizes the imaging protocols and procedures that shall be performed for an FDG-PET/CT exam. Detailed descriptions are included in the referenced FDG-PET/CT UPICT protocol sections.

It is recommended that all FDG-PET/CT scans for an individual subject be performed on the same PET/CT scanner hardware and software throughout the trial.  In the event of equipment malfunction, follow-up scans on an individual participant can be performed on a different scanner of the same model and software version provided it has met the scanner qualification requirements .  The follow up scans should be performed with identical acquisition parameters as the first (baseline), inclusive of all the parameters required for both the CT and PET acquisitions.

The FDG-PET/CT UPICT Protocol (Reference Section 7.1.1) describes scanning strategies that can be used in a clinical trial. For strategy 1, there is no intent to obtain a diagnostic CT scan at the FDG-PET imaging session, however a low-dose CT scan is needed for attenuation correction. For strategy 2, a diagnostic CT scan is obtained. There are further considerations that must be followed for each of the two strategies. The workflow chosen for a given protocol should be described in the protocol and should be tailored commensurate to the level of expectation of the obtained data (e.g. qualitative or quantitative SUV analysis).

Strategy 1: For FDG-PET/CT in which the CT is used for attenuation correction and localization only (no diagnostic CT intent):

- CT Scout (topogram), followed by

- CT for anatomic localization and attenuation correction, followed by

- PET Emission scan acquisition

Strategy 2: For FDG-PET/CT in which a diagnostic CT is performed in conjunction with FDG-PET, one of two strategies shall be used. Either (2a) follow Strategy 1 and then, with no or minimal patient motion after the PET Emission scan acquisition, perform an additional IV contrast-enhanced diagnostic CT or (2b) perform a contrast-enhanced diagnostic CT before following the workflow described in Strategy 1.

| Scanning Strategy (Workflow) | Technologist | Technologist will follow Profile compliant workflow strategy, which will be compatible with Acquisition Device capability. The same workflow used at baseline shall be used at all subsequent time points. Use consistent with common data format mechanism |
| --- | --- | --- |

For both strategies, there are several common issues specific to the CT exam that may have an impact on quantitative FDG-PET output, which need attention and protocol specification. These include (1) contrast material administration, (2) respiratory motion compensation instructions and (3) CT scanning technique (kVp, mAs and pitch). Here is a summary of the acceptable level of behavior/procedure for each of these three issues.

At a minimum, all these issues should be addressed in the clinical trial protocol, ideally with consistency across all sites and all subjects (both inter-subject, and intra- and inter-facility) with the target of consistency across all time points for each given subject. The actual details of imaging for each subject at each time point should always be recorded.  Any particular clinical trial should NOT allow some sites to implement one strategy and other sites to implement the alternative.

CT Exam Variables and Specifications:

- Contrast Agents - The presence of a positive contrast agent (IV or oral), by affecting the CT attenuation map, can result in a small variability of quantitative SUV evaluation. If this were the only consideration, then ideal would be to prohibit CT contrast administration. However, in some clinical situations (dependent upon tumor type, tumor behavior or level of anatomic interest), the benefit of CT contrast agents may outweigh the small errors induced in SUV measurement that may include increased SUV variability. Each protocol should specify the desired approach for the given study. Most importantly, for each subject, the same approach should be followed for all imaging time points.

In cases where CT contrast agents are used, there are two main strategies:

Strategy 1: No IV; dilute positive oral contrast allowed

Strategy 2: Use negative or dilute positive oral contrast for the non-attenuation CT scan. Ensure that the diagnostic CT acquisition (which may be performed with IV contrast) is performed consistently for a given subject across all time points.

| CT Contrast agent | Technologist | CT contrast agents shall be given commensurate with the workflow strategy as selected from above move text to box |
| --- | --- | --- |

- Respiratory Motion - Respiratory motion causes SUV errors by two mechanisms: motion blurring and errors in attenuation correction due to mismatches between CT-based attenuation map and emission data. Various strategies could be used to minimize, document and compensate for respiratory motion.  Shallow breathing shall be performed during CT AC acquisition (see UPICT Protocol section 7.1.1). The subject should (a) be monitored and if breathing pattern is not consistent with shallow breathing expectation, coached in the breathing protocol and (b) should remain motionless throughout the scan.

| Respiratory motion minimization | Technologist | The Technologist shall observe subject breathing; if the subject is not breathing commensurate with shallow breathing expectation, provide verbal instruction to the subject to perform shallow breathing prior to and during CT and PET scans Add scanner based-corrections (validated by mfr.) |
| --- | --- | --- |

The Technologist shall document physical factors that adversely influence subject positioning or limit their ability to comply with instructions (e.g. breath-hold, remaining motionless, etc.).  Ideally there would be a standardized method for recording such information.

| Breathing and motion non-compliance | Technologist | The Technologist shall document issues regarding subject non-compliance with breathing and motion Use consistent with common data format mechanism |
| --- | --- | --- |
|  |  | Ideally, the Technologist shall document issues regarding subject non-compliance with breathing and motion using a standardized method for recording such information |

- CT Technique

The actual kVp and exposure (CTDI, DLP) for each subject at each time point should be recorded. CT dose exposure should be appropriately reduced wherever possible and particularly in smaller patients and children.

| CT Technique: Dose Exposure | Technologist | The Technologist shall ensure that CT dose exposure is appropriate and is reduced whenever possible, such as smaller subjects and children. Protocols defined by Image Gently/ Wisely Should be recorded and documented Use consistent with common data format mechanism |
| --- | --- | --- |

Regarding CT radiation exposure, the lowest radiation dose necessary to achieve the diagnostic objective should be used. For a given protocol, the purpose of performing the CT scan (attenuation correction only or attenuation correction and anatomic localization) should be determined. The CT technique (tube current, rotation speed, pitch, collimation, kVp, and slice thickness) used should result in as low as reasonably achievable exposure needed to achieve the intended goal of imaging working with the scanner manufacturer (Appendix reference) to achieve this objective. The technique used for an imaging session should be repeated for that subject for all subsequent time points assuming it was properly performed on the first study.

2.1 Required Characteristics of Resulting Data

The PET/CT exam will consist of two components, the PET emission scan and the CT transmission scan.  From these data sets, the non-attenuation-corrected PET images are reconstructed for quality control purposes and attenuation-corrected PET images are reconstructed for qualitative interpretation and quantitative analysis. Instrument specifications relevant to the Acquisition Device are included in Section IV Compliance – Acquisition Device.

2.1.1 PET Data Structure (UPICT Section 7.1.2)

The matrix, slice thickness, and reconstruction zoom shall yield a target voxel size of 3 – 4 mm in all three dimensions, although not necessarily isotropic and not achieved by re-binning, etc., of the reconstructed images.

| PET Matrix/Voxel size | Technologist | The Technologist shall perform the scan such that he matrix, slice thickness, and reconstruction zoom shall yield a target voxel size of 3 – 4 mm in all three dimensions, although not necessarily isotropic and not achieved by re-binning, etc., of the reconstructed images. |
| --- | --- | --- |

2.2 Imaging Procedure

2.2.1 Timing of Image Data Acquisition

FDG uptake into both tumors and other body tissues is a dynamic process that may increase at different rates and peak at various time points dependent upon multiple variables. Therefore, it is extremely important that (1) the time interval between FDG administration and the start of emission scan acquisition is consistent and (2) when repeating a scan on the same subject, it is essential to use the same interval between injection and acquisition in scans performed across different time points.

While the “target” tracer uptake time is 60 minutes, the “acceptable” window is from 55 to 75 minutes to ensure that imaging does not begin prematurely so as to allow adequate tumor uptake of FDG and to account for the practicality of work flow that can result in delays in imaging later than 60 minutes after FDG injection. The exact time of injection must be recorded; the time of injection initiation should be used as the time to be recorded as the radiotracer injection time. The injection and flush should be completed within one minute with the rate of injection appropriate to the quality of the vein accessed for FDG administration so as to avoid compromising the integrity of the injection vein.

When repeating a scan on the same subject, especially in the context of therapy response assessment, it is essential to apply the same time interval with target window of +/- 10 minutes (with an acceptable window of +/- 15 minutes) provided that the scan must not begin prior to 55 minutes after the injection of FDG. If a limited anatomy scan is obtained at follow-up after a whole body scan was performed at baseline, one should consider adjusting the timing of the follow up scan to be congruent with the timing for the same anatomic region as achieved during the baseline study.

If, for scientific reasons, an alternate time (between dose administration and scan acquisition) is specified in a specific protocol, then the rationale for this deviation should be stated; inter-time point consistency must still be followed.

| Tracer Uptake Time | Technologist | The time of FDG injection shall be entered into PET/CT scanner console during the acquisition and will be recorded by the scanner into the appropriate DICOM field according to the DICOM conformance statement |
| --- | --- | --- |
| Tracer Uptake Time: | Technologist | The Technologist shall ensure that the tracer uptake time for baseline is 55 to 75. When repeating a scan on the same subject, especially in the context of therapy response assessment, it is essential to apply the same time interval with target window of +/- 10 minutes (with an acceptable window of +/- 15 minutes) provided that the scan must not begin prior to 55 minutes after the injection of FDG. |

The following sections describe the imaging procedure.

2.2.2 Subject Positioning (UPICT Section 7.2.1)

Consistent positioning avoids unnecessary variance in attenuation, changes in gravity-induced shape and fluid distribution, or changes in anatomical shape due to posture, contortion, etc. During PET-CT, subjects should be positioned in the center of the field of view (FOV), preferably with the subjects’ arms to be positioned over head for whole-body imaging (to minimize beam hardening and FOV truncation artifacts).  In the case of dedicated brain or head/neck scans, the arms should be positioned along.  If the subject is physically unable to maintain arms above head for the entire whole-body examination then the arms can be positioned along the side before the start of the scan, unless the protocol specifically excludes such subjects. Arm positioning in a particular subject should be consistent between the PET emission and CT transmission scans at each time point and should be as consistent as possible across all time points.

| Subject Positioning | Technologist | The Technologist shall position the subject according to the UPICT and/or specific protocol specifications consistently for all scans |
| --- | --- | --- |

| Positioning Non-compliance | Technologist | The Technologist shall document issues regarding subject non-compliance with positioning |
| --- | --- | --- |
|  |  | The Technologist shall document issues regarding subject non-compliance with breathing and positioning using a standardized method for recording such information Add language about common data format mechanism |

2.2.3 Scanning Coverage and Direction (UPICT Section 7.1.1)

For most Oncology indications, anatomic coverage should include from the skull base (external auditory meatus) to the mid-thigh.  If other ranges are used, which may be appropriate for specific clinical trials, then the clinical trial protocol should provide specific instructions with justification.  Scanning direction should be caudiocranial to minimize effects from increasing bladder activity during the scan.  Scanning direction should be specified in the clinical trial protocol.  It is critical that for a given subject, scanning direction on baseline scans be duplicated at follow-up time points.

| Scanning Direction | Technologist | The Technologist shall scan the subject caudocranial for whole body examination unless otherwise specified by the protocol. Scanning direction shall be the same for each subject at all time points. The scanning direction shall be entered into the PET/CT console during the acquisition and will be recorded by the scanner into the appropriate DICOM field. |
| --- | --- | --- |
| Anatomic Coverage | Technologist | The Technologist shall perform the scan such that the anatomic coverage is acquired according to the protocol specifications and the same for all time points |
|  |  | Need to capture the number of bed positions Common data form input needed here |

3. Imaging Data Reconstruction & Post-Processing

3.1 Imaging Data Reconstruction (UPICT Section 7.3)

Reconstructed Image Data: This is the image data exactly as produced by the reconstruction process on the PET or PET/CT scanner, i.e., a stack of DICOM slices/files constituting a PET image volume with no processing other than that occurring during image reconstruction. This is always a stack of DICOM slices/files constituting a PET image volume that can be analyzed on one or more of the following: PET scanner console, PET image display workstation, PACS system, etc.  See Section IV Compliance – Image Reconstruction Software for specifications.

Add normative text

## 3.2 Imaging Data Post-processing (UPICT Section 8)

Post-Processed Image Data: An image that has been transformed in some manner, including but not limited to: smoothing, sharpening, image zoom, rotation/translation, resampling, interpolation, slice averaging, MIP, etc. This is typically a stack of DICOM slices/files constituting a PET image volume that can still be analyzed on one or more of the following: PET scanner console, PET image display workstation, PACS system, etc.

Standard whole-body FDG-PET oncology studies typically include all necessary data corrections and processing within the reconstruction process and do not require additional post-processing other than data de-identification (see section 11.2). More advanced studies such as those including dynamic imaging may require additional post-processing as specified in the individual protocol.

| Post-Processing | Image Reconstruction Software and Display Workstation | Image reconstruction and post-processing processing parameters are entered into the protocol used on the PET/CT scanner console and the protocol should be used consistently for all patients and studies in the trial. The parameters are recorded in the appropriate DICOM fields according to the DICOM conformance statement for the PET/CT scanner. This information should also be recorded into relevant CRFs as indicated by individual trials. Quantaitive analysis should be appllied to unprocessed images (i.e. Image registration/interpolation/zooming) |
| --- | --- | --- |
| Data Archiving |  | The originally reconstructed PET data set should always be archived at the local site and where possible raw data should also be archived. If processed PET datasets are required, they should be saved as separate secondary datasets and archived. |

Concepts presented in UPICT Section 8.2.3 regarding difference between ‘visualized data’ and ‘data used for quantification’. Acceptable level: For visual inspection/interpretation of PET/CT data using the display workstation,  bi-linear or tri-linear interpolation and zooming may be used to display the images in a different matrix size that the original data. In addition, so-called maximum intensity projections (MIP) may be generated as they may facilitate localization and detection of lesions. Additional processing, such as zooming, re-binning, reorientation and filtering may be applied upon user request only. User should be able to manipulate color scale settings (window/level and color table). It should always be possible to revert to the default orientation, zoom and binsize (preferably a ‘revert to default’ button is available). For PET images, only linear color scales should be used to preserve the quantitative representation of SUV units.

3.3 Imaging Data Storage And Transfer

DICOM format should meet the Conformance Statement written by manufacturer of the PET/CT system. PET data shall be encoded in the DICOM PET or Enhanced PET Image Storage SOP Class, and in activity-concentration units (Bq/ml) with additional parameters in public DICOM fields to calculate SUVs (e.g. height, weight, scale factors). CT data should be encoded in CT or Enhanced CT Image Storage SOP Class. DICOM data shall be transferred using the DICOM Part 8 network protocol or as offline DICOM Part 10 files for media storage including CDs and DVDs. They shall be transferred without any form of lossy compression.

| Data Archiving | Imaging Facility |  |
| --- | --- | --- |
| Data Archiving | Workstation |  |

## 4. Image Analysis (UPICT Section 9)

The Image Analyst, through interaction with the Workstation Analysis tools, will be able to perform certain measurements. Image Analysis has qualitative and quantitative tasks.  Both require consistency and high quality images.  Quantitative imaging requires additional system characteristics described further in this Profile.

4.1 Input Data

The output images of Reconstruction and Post processing software activity are considered the input for Image Analysis.  If the Image Analyst alters input data (e.g. zoom) this is considered part of Image Analysis activity.  If this occurs, the original input data will be maintained as separate file, both to be stored. (See Section 3.2)

4.2 Methods to Be Used

Each tissue/organ to be investigated quantitatively (either tumor lesion or normal tissue) is characterized by defining a region-of-interest (ROI) or volume-interest (VOI) and calculating a parameter such as the maximum within the ROI/VOI as a metric of glycolytic activity for that tissue/organ.  The image analyst will use tools (as defined in Section IV Compliance – Software Analysis Tools) to define ROI/VOI’s both in clinical subject datasets and in the digital reference object (DRO).

4.3 Required Characteristics of Resulting Data (UPICT Section 9.3)

The specific trial protocol should prospectively define the SUV parameter that is required for each lesion, which will be used for the imaging endpoint. Some studies may also compare different metrics and will require recording multiple parameters. SUV measures (and the analysis tools used to obtain them) need to be specified for each protocol and need to be used consistently across all subjects and across all sequential lesion measurements. In addition to reporting the SUVmax (maximum voxel value or hottest voxel) for a given lesion, a signal-averaged measurement such as the SUVpeak, can be recorded to avoid potential biases due to noise.

- Automatic vs. Semi-automatic vs. Manual and circumstance qualifiers - Strategy for analysis of necrotic core lesion

- Normalization (if required) (e.g. based on lesion size, glycemic status)  - Strategies around the issue of partial volume effects - Strategies for glucose correction   Analysis to be based on original reconstructed PET image (no additional re-binning, resampling, smoothing by user is allowed Issues relevant to dynamic analysis

## 5. Image Interpretation and Reporting (UPICT Section 10)

No QIBA Profile specification can be provided for image interpretation at this time.  Image Interpretation is considered to be beyond the scope of this document.  Refer to FDG-PET/CT UPICT Protocol (Section 10).

| Image Reporting | Imaging Facility | Populate reports from DICOM header information Use structured reporting |
| --- | --- | --- |

## 6. Quality Control

The following section deals with multiple aspects of quality control in FDG-PET/CT studies.  (See FDG-PET/CT UPICT Protocol Section 12 for additional information). This includes selecting and qualifying a PET/CT imaging facility, imaging personnel and specific PET/CT scanners.  In addition, the use of phantom imaging (prior to study initiation and ongoing) is discussed as well as identifying subjects whose data may need to be censored due to lack of data integrity.  Finally, post-image-acquisition quality assessment is detailed.

  - Image Acquisition Facility

It is essential to implement procedures that ensure quality assurance of the scanning equipment and reliable image acquisition methodology.  These processes must be set-up at the outset, and be followed throughout the duration of the trial. A facility “imaging capability assessment” prior to facility selection is therefore a requirement for any trial using FDG-PET/CT.  This will include assessment of:

- appropriate imaging equipment and quality control processes

- appropriate ancillary measurement equipment and radiotracer quality control processes

- experienced personnel (technologists, radiologists and other imaging experts)

- procedures to assure imaging protocol compliance during the trial

    - FDG-PET/CT Acquisition Scanner (UPICT Sections 12.1.1.5 & 12.1.1.6)

FDG-PET/CT studies as developed in this profile require a dedicated PET/CT scanner.  PET/CT scanners should be identified based on their vendor, model and machine name. Hardware specifications should be documented. Software versions in place at the time of trial initiation and at all upgrades should be documented as well.

The PET/CT scanner must undergo routine quality assurance and quality control processes (including preventive maintenance schedules) appropriate for clinical PET/CT applications.  In order to assure adequate quantitative accuracy and precision of PET/CT imaging results, additional quality assurance measures are required, as discussed below.

If there is more than one PET/CT scanner at a facility, it is beneficial to qualify each of the scanners.  This will ensure that if the primary PET/CT scanner is temporarily unavailable, the FDG-PET/CT study may proceed on a secondary scanner.

| Physical Inspection | Technologist | Shall, on a daily basis, check gantry covers in tunnel and subject handling system |
| --- | --- | --- |
| QA/QC Checks | Technologist | QC/QC procedures shall be performed each day according to vendor recommendations. A table of QA/QC procedures for a subset of specific PET/CT scanners form each vendor is included in Appendix G.2. Daily QC procedures shall be performed prior to any patient scan. |

    - Other devices (e.g. dose calibrator, glucometer, scales. . .)

6.1.2.1 Dose calibrator  (UPICT Section 12.1.1.4)

From ANSI standard N4213, 1978, 1986, 2004. Also IAEA document [Reference]

Add definitions from IAEA, i.e. 'accuracy' = 'traceability'

Everything is based on unit dose and calibration sources are in the 'syringe'.

| Constancy | Technologist | Shall be evaluated daily (or after any dose calibrator event) using a NIST-traceable (or equivalent) simulated F-18, Cs-137, or Co-57 dose calibrator standard and confirmed that net measured activity differs by no greater than ±2.5 % from the expected value. |
| --- | --- | --- |
| Accuracy | Technologist | Shall be evaluated monthly (or after any dose calibrator event) with a NIST-traceable (or equivalent) simulated F-18 dose calibrator standard. Shall confirm that net measured activites differ no greater than ±2.5% from expected value. |
| Linearity | Technologist or Radiation safety officer or Qualified Medical Physicist | Shall be evaluated quarterly (or after any dose calibrator event) using either F-18 or Tc-99m and Should be within ±2.5 % of the true value over an operating range of 1 to 30 mCi and the true value is determined by a linear fit (to the log data) over the same operating range. |

6.1.2.2 Scales and stadiometers (UPICT Section12.1.1.2)

Scales and stadiometers should be inspected and calibrated at installation and annually.

| Scales and stadiometers | Approved personnel | Shall be evaluated annually or after any repair by qualified personnel. Shall be confirmed that error is less than +/- 2.5% from expected values using NIST-traceable or equivalent standards. |
| --- | --- | --- |

6.1.2.3 Blood glucose level measurement device (UPICT Section 12.1.1.3)

Glucose measurements should be made using a CLIA approved, CLIA cleared, or equivalent (outside the US) glucose measurement technique.

| Blood glucose level measurement device | Approved personnel | Shall have QA/QC testing and calibration performed using a CLIA approved, CLIA cleared, or equivalent (outside US) procedure. |
| --- | --- | --- |

6.1.2.4 Clocks and timing devices (UPICT Section 12.1.1.1)

PET/CT scanner computer and all clocks in an Imaging facility used to record dose/injection measurements should be synchronized to standard time reference within +/-1 minute. These include any clocks or timekeeping systems that are connected with a patient's FDG-PET/CT study, in particular those associated with the dose calibrator, the injection room, and the scanner itself. The synchronization of all clocks should be monitored periodically as part of ongoing QA program. In particular, clocks should be inspected immediately after power outages and changes in daylight savings.

| Scanner and site clocks | Approved personnel | Synchronization of all clocks used in the conduct of the FDG-PET/CT study shall be checked weekly and after power outages and changes in daylight savings time settings. PET/CT scanner computer and all clocks in an Imaging facility used to record dose/injection measurements shall be synchronized to standard time reference within +/-1 minute |
| --- | --- | --- |

Display system

    - Personnel

For each of the personnel categories described below, there should be training, credentialing, continuing education and peer review standards defined.  Guidelines for training/credentialing for each resource category are summarized below (UPICT Protocol section 2.1).

| Personnel Roster | Imaging Facility Coordinator | Each site shall have the support of technical, physics, and physician personnel (as further defined below), trained in the use of FDG-PET/CT in the conduct of oncologic clinical trials prior to trial activation and subject accrual. |
| --- | --- | --- |
| Technologist Qualification | Imaging Facility Coordinator | Technologist certification shall be equivalent to the recommendations published by the representatives from the Society of Nuclear Medicine Technologists Section (SNMTS) and the America Society of Radiologic Technologists (ASRT) in J Nucl Med Technol. 2002;30:201-204 and should also be congruous with all local, regional, and national regulatory requirements for the administration of ionizing radiation to patients. |
| Medical Physicist Qualification | Imaging Facility Coordinator | Medical physicists shall be certified in Medical Nuclear Physics or Radiological Physics by the American Board of Radiology (ABR); in Nuclear Medicine Physics by the American Board of Medical Physics (ABSNM); in Nuclear Medicine Physics by the Canadian College of Physicists in Medicine; or in Nuclear Medicine Physics and Instrumentation by the American Board of Science in Nuclear Medicine, or equivalent certification in other countries. |
| Physician Qualification | Imaging Facility Coordinator | Physicians overseeing and interpreting PET/CT scans shall be qualified by the ABR (Diagnostic and/or Nuclear Radiology) or American Board of Nuclear Medicine (ABNM) or equivalent within the United States or an equivalent entity appropriate for the geographic location in which the imaging study(ies) will be interpreted. |

    - Site Qualification/Accreditation Process (UPICT Section 12.1.1)

Imaging facility accreditation is typically considered for clinical practice purposes (e.g., ACR, ICANL) while facility qualification (e.g., SNM-CTN, ACRIN, CQIE and imaging core labs) is the analogous term applied to clinical research/clinical trial utility.  In order to be compliant with this Profile, the imaging facility needs to provide documentation and verification of the process they have followed for their fit for purpose usage.  Appropriate forms, checklists or other process documents should be maintained to facilitate this verification.  If exceptions to any of the performance standards stated below occur and cannot be remediated on site, the site should promptly communicate the issue to the appropriate internal overseer for advice as to how the irregularity should be managed.

In addition to documenting the level of performance required (and the level of performance achieved), the frequency of facility accreditation/qualification also needs to be described.

    - Phantom Imaging

To qualify the PET/CT scanner for clinical practice or for a clinical trial, a phantom imaging process is required.  Specific phantoms might be needed for certain types of cancers or anatomic locations and therefore might vary from trial to trial based on diagnosis, treatment and/or anatomic location.  Options that might be considered on a per-protocol basis include, but are not limited to 1) each site would use the same phantom for the duration of the trial (but the phantoms might not be exactly the same among all sites), 2) all sites would use the same general type of phantom for the duration of the trial, 3) all sites would use phantoms that are precisely specified for the duration of the trial, 4) all sites would share the exact same phantom for the duration of the trial.

The phantom scans and performance evaluation should be performed prior to the start of a trial and repeated during the course of the trial as specified by the individual trial.  Any changes to scanner equipment, including major hardware changes or any software version change, should be immediately reported to trial sponsor and/or imaging CRO and may result in the need for imaging qualification renewal prior to imaging additional trial patients.  In particular, it is strongly recommended that patients undergoing longitudinal study be scanned on the same PET/CT system with the same software version whenever possible.

      - Uniformity, Normalization  and Calibration

Verification of scanner normalization with a phantom uniformity is a minimum requirement for all scanners used in clinical trials including those that only have qualitative endpoints. For trials with quantitative PET measurements, this assessment should also include a comparison against a dose calibrator to ensure quantitative accuracy; that is, a comparison of the absolute activity measured versus the measured amount injected should be performed (5% NL, 10% EANM, ACRIN). This comparison is particularly important after software or hardware upgrades. If the trial requires absolute quantification in baseline images or absolute changes in longitudinal studies, it should be considered to include an image quality and/or contrast recovery QC assessment as part of the routine QC procedures and/or scanner validation process, see Appendix E Of UPICT Protocol. Clinical trials using only relative changes in longitudinal studies may not require contrast recovery assessments provide there is appropriate consideration for the minimum size of target lesions based on the partial volume effect.

An essential requirement for extracting quantitative data from images is that there be known calibration accuracy and precision and/or cross calibration of the PET/CT system against the (locally) used dose calibrator (within 10%).   The QC procedures should utilize the same acquisition/reconstruction protocol, software and settings that are used for the subject scans.

| Uniformity QC | Technologist | Shall assess (at least quarterly and following software upgrades) transverse and axial uniformity across image planes by imaging a uniformly filled object and . . . |
| --- | --- | --- |
| Cross Calibration | Technologist | Shall perform (at least quarterly and after scanner upgrades, new setups and modifications to the dose calibrator) to monitor and identify discrepancies between the PET scanner and dose calibrator |

      - Resolution  (UPICT Section 12.1.1.11)

| Resolution | Nuclear Medicine Physician | Shall perform (on at least an annual basis) and document a qualitative resolution QC test by using the manufacturer’s seetings and demonstrating resolution of normal gross anatomic features within clinical images of the brain, heart and abdomen. |
| --- | --- | --- |
| Resolution | Medical Physicist | Shall perform (on at least an annual (?) basis) and document performance of a quantitative assessment (using phantom with size defined ‘lesions’) for lesion resolution. |

The assessment of adequate resolution should include both a qualitative evaluation (using clinical images) and quantitative assessment (using phantom-defined criteria).  The phantom-defined requirements are more completely described in UPICT protocol Section 12.1.1.11.

      - Noise (UPICT Section 12,1.1.12)

| Noise | Medical Physicist | Shall perform qualitative assessment of image noise in phantom images to be of consistent and acceptable quality |
| --- | --- | --- |

    - Phantom imaging data analysis

Display station to produce correct SUVs from QIBA DRO

  - Quality Control of FDG-PET/CT studies

    - Data Integrity

The integrity of DICOM image headers should be reviewed and confirmed for DICOM standard compliance, regulatory compliance (including privacy protection, such as may be required by such rules as the HIPAA Privacy Rule if applicable), protocol compliance, sufficiency for the intended analysis (e.g., to compute SUV) and consistency with source data such as CRFs. In some cases, internal references such as the liver can be used for quality control to confirm acceptable ranges of SUVs.

    - Determination of Image Quality (?refer to UPICT re: DRO noise/resolution/uniform?)

    - Determination of suitable tumor lesions

    - Determination of subjects unsuitable for FDG-PET/CT analysis

    - Determination of FDG-PET exams unsuitable for FDG-PET/CT analysis

CT images should be reviewed by the Image Analyst for potential artifacts such as beam hardening, metal objects, and motion. PET images should be compared to the CT images for proper image registration and potential attenuation correction artifacts.

1. Quality assessment of exam and actions based on assessment (of what specific parameters??)  Both uncorrected and attenuation corrected images need to be assessed in order to identify any artifacts caused by contrast agents, metal implants and/or subject motion (EU).  CONSENSUS: - Image quality should be noted. For example, movement or misregistration can lead to poor quality quantitative data and invalid numbers. Some images may be too poor in quality to quantify. CONSENSUS: - Statistical quality of images is important to report, but not a full substitute for quality. Liver noise per PERCIST was considered a reasonable start.

  - Quality Control of Interpretation

# Minimum is Intra-reader variability study needed. If a 2-Reader paradigm, then inter-reader variability is needed as well.  Lesion quantitative vs. interpretation; target lesion choice; calling new lesion – additional investigation needed.  Currently unclear which statistics to evaluate and what should it trigger. The following requirements are placed on QC associated with interpretation.

# IV. Compliance

## 7. Image Acquisition Site

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large subject population under consideration.  For imaging it is important to consider the availability of:

Appropriate imaging equipment and quality control processes,

Appropriate ancillary equipment and access to radiotracer and contrast material,

Appropriately trained Radiologists/Nuclear Medicine Physicians for image analysis . . .

Experienced Technologists (CT and PET trained) for the subject handling and imaging procedure

Medical Physics support to ensure appropriate . . .

Processes that assure imaging QIBA Profile-compliant image generation in appropriate time window

A QA program for PET/CT scanners and ancillary devices must be in place to achieve the goals of the clinical trial.  The minimum requirements are specified above. This program shall include (a) elements to verify that imaging facilities are performing correctly and (b) elements to verify that facilities’ PET/CT scanners(s) is(are) performing within specified calibration values.  These may involve additional phantom testing that address issues relating to both radiation dose and image quality (which may include issues relating to water calibration, uniformity, noise, spatial resolution – in the axial plane-, reconstructed slice thickness z-axis resolution, contrast scale, and others.  This phantom testing may be done in addition to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.

Discuss PET-only or MR/PET scanners as later extensions

| PET/CT Scanner | Acquisition Facility | What are minimum (big picture) requirements - e.g. dedicated PET/ single slice CT vs. multi? PET component full-ring (i.e. from UPICT) or performance based on mfr NEMA NU-2 specs. (or user demonstrated specs). Note this is separate from scanner qualification testing. This profile only addresses PET/CT scanners. No statement about PET-only or MR/PET |
| --- | --- | --- |
| CT Scanner Calibration | Technologist | Shall perform daily water equivalent phantom analysis; ensure that output is acceptable and manually enter on form /electronic database |
| PET Scanner Calibration | Technologist | Shall perform daily/weekly/monthly scanner QA; ensure that output values are acceptable and manually enter on form/electronic database |
| Dose Calibrator |  | Calibrated to F-18 using NIST traceable source |

## 8. Acquisition Device

Use DICOM attributes to follow version number of software for: 1 Acquisition, 2 Reconstruction, 3 Post-processing, 4 Display/VOI analysis, 5 Dynamic Analysis. Build list (on console) of dates of all software versions.  The scanner software version should be identified and tracked across time, with updates and changes in scanner software noted during the course of the trial.

Decay reference time should equal (PET) scan acquisition start time and integral model should be used for decay correction.  Scanner should perform all decay corrections (not the operator).

Image data are given in units Bq/ml.

All needed information for fully corrected administered activity (e.g. residual activity, injection time, calibration time) is required.) Note that use of the term administered activity below refers to fully corrected activity.

Baseline level compliance requires that the DICOM image set from the subject’s PET/CT scan and necessary metadata (that is not currently captured by all PET scanner acquisition processes) is captured in trial documentation, e.g. case report forms.  The metadata is required to perform the quantitative analysis and perform quality control on SUV covariates.  This includes for example, post-injection residual dose and subject height.

Clarify who does what when for calibration and point to site QA/QC above to reduce confusion.

| CT Calibration Tracking | Acquisition Device | Daily water equivalent phantom values shall be tracked in the DICOM header |
| --- | --- | --- |
| PET Calibration Tracking | Acquisition Device | Daily/weekly/monthly scanner QA values shall be included in the DICOM header |
| Dose Calibrator Calibration Tracking | Acquisition Device | Calibration factor for an F-18 NIST -traceable (or equivalent) source with identifying information shall be tracked in the DICOM header. |
| PET Scanner calibration | Acquisition Device | Shall be able to be calibrated according following ACRIN type U cyl (F-18 or Ge-68) - ideally same used for calibration Long scan time 60 min or less, ACRIN-type ROI analysis Ave SUV 0.98 to 1.02 Slice-to-slice variability of ± 5%. (not including end slices, as per ACRIN) |
|  |  | In-plane uniformity for above phantom should be less than X % |
| Weight | Acquisition Device | Shall be able to record patient weight in lbs or kg as supplied from the modality worklist and/or operator entry into scanner interface. Shall be stored in Patient Weight field (0010,1030) in the DICOM image header, as per DICOM standard. |
|  |  | Patient weight shall be specifiable with 4 significant digits. Patient weight is transferred directly from measurement device into scanner by electronic, HIS/RIS, or other means, bypassing all operator entry, but still requiring operator verification. |
| Height | Acquisition Device | Shall be able to record patient height in feet/inches or cm/m as supplied from the modality worklist and/or operator entry into scanner interface. Shall be stored in Patient Size field (0010,1020) in the DICOM image header, as per DICOM standard. |
|  |  | Patient height shall be specifiable with 3 significant digits. Patient height is transferred directly from measurement device into scanner by electronic, HIS/RIS, or other means, bypassing all operator entry, but still requiring operator verification. |
| Blood glucose level | Acquisition Device | Shall be able to Record patient blood glucose level, in units of mg/dl, or mMol/l, time of measurement, as supplied by operator entry into the scanner interface. Shall be recorded in a dedicated field in the DICOM image header. Patient blood glucose level is transferred directly from measurement device into scanner by electronic, HIS/RIS, or other means, bypassing all operator entry, but still requiring operator verification. |
| Administered Radionuclide | Acquisition Device | Shall be able to record the radionuclide (i.e. F-18), as supplied by operator entry into the scanner interface. Shall be recorded in Radionuclide Code Sequence field (X,X) in the DICOM image header. |
| Administered Radiotracer | Acquisition Device | Shall be able to record the radiotracer (i.e. FDG), as supplied by operator entry into the scanner interface. Shall be recorded in Radiotracer Code Sequence field (0054,0300) in the DICOM image header. |
| Administered Radiotracer radioactivity | Acquisition Device | Shall be able to record the administered radioactivity , in mCi or MBq, as supplied by operator entry into the scanner interface. Shall be recorded in Radionuclide Total Dose field (0018,1074) in the DICOM image header. |
|  |  | Shall be able to record with separate entry fields on scanner interface: the pre-injection FDG radioactivity time of measurement of pre-injection FDG radioactivity the residual dose after injection time of measurement the residual radioactivity after injection Shall automatically calculate the administered radioactivity and store in the Radionuclide Total Dose field (0018,1074) in the DICOM image header. |
|  |  | Patient Administered Radiotracer radioactivity information shall be transferred directly from measurement device into scanner by electronic, HIS/RIS, or other means, bypassing all operator entry, but still requiring operator verification. |
| Administered Radiotracer Time | Acquisition Device | Shall be able to record the time of the start of dose injection as supplied by operator entry into the scanner interface. Shall be recorded in Radiopharmaceutical Start Date Time field (0018,1078) (preferred) or Radiopharmaceutical Start Time field (0018,1072). |
|  |  | Shall be able to record the time of the start of dose injection as supplied by operator entry into the scanner interface. Shall be recorded in Radiopharmaceutical Start Date Time field (0018,1078). I.e. not Radiopharmaceutical Start Time field (0018,1072). Shall be able to record the time of the stop of dose injection as supplied by operator entry into the scanner interface. Shall be recorded in Radiopharmaceutical Stop Date Time field (0018,1079). |
| Decay Correction Methodology | Acquisition Device | Encoded voxel values with Rescale Slope field (0028,1053) applied shall be decay corrected by the scanner software (not the operator) to a single reference time (regardless of bed position), which is the start time of the first acquisition, which shall be encoded in the Series Time field (0008,0031) for original images. Corrected Image field (0028,0051) shall include the value “DECY” and Decay Correction field (0054,1102) shall be “START”. |
| Scanning Workflow | Acquisition Device | Shall be able to support Profile Protocol (Section III) PET and CT order(s) of acquisition. Shall be able to pre-define and save (by imaging site) a Profile acquisition Protocol for patient acquisition. |
|  |  | Shall be able to interpret previously-reconstructed patient images to regenerate acquisition protocol. Shall be configurable to store (or receive) acquisition parameters as pre-defined protocols (in a proprietary or standard format), to allow re-use of such stored protocols to meet multi-center specifications and to achieve repeatable performance across time points for the same subject |
| CT Technique | Acquisition Device | Shall record all technique (e.g. kVp pitch, exposure (mAs) etc.) in the CT image header. |
| CT based attenuation correction | Acquisition Device | Shall record information in PET image header which CT images were used for corrections (attenuation, scatter, etc.). |
| PET-CT Alignment | Acquisition Device | Shall be able to align PET and CT images within ±2 mm. |
|  |  | Shall be able to align PET and CT images within ±2 mm precision under maximum load over the co-scan length. |
| CT Absorbed Radiation Dose | Acquisition Device | Shall be able to record the absorbed dose (CTDI, DLP) in a DICOM Radiation Dose Structured Report. |
| Activity Concentration in the Reconstructed Images | Acquisition Device | Shall be able to store and record (rescaled) image data in units of Bq/ml and use a value of BQML for Units field (0054,1001). |
| Tracer Uptake Time | Acquisition Device | Shall be derivable from the difference between the Radiopharmaceutical Date Time field (0018,1078) (preferred) or Radiopharmaceutical Start Time field (0018,1072) and the Series Time field (0008,0031) or earliest Acquisition Time field (0008,0032) in the series (i.e., the start of acquisition at the first bed position). |
| PET Voxel size | Acquisition Device | Shall be able to reconstruct PET voxels with a size of 3 to 4 mm in all three dimensions (as recorded in Voxel Spacing field (0028,0030) and computed from the reconstruction interval between Image Position (Patient) (0020,0032) values of successive slices). Voxels shall be square, although they are not required to be isotropic in the z (head-foot) axis. |
|  |  | Shall be able to reconstruct PET voxels with a size of 1-3 mm in all three dimensions (as recorded in Voxel Spacing field (0028,0030) and computed from the reconstruction interval between Image Position (Patient) (0020,0032) values of successive slices). Voxels shall be isotropic. |
| CT Voxel size | Acquisition Device | Shall be no greater than the reconstructed PET voxel size. Voxels shall be square, although are not required to be isotropic in the Z (head-foot) axis. Not required to be the same as the reconstructed PET voxel size. |
| Subject Positioning | Acquisition Device | Shall be able to record the subject position in the Patient Orientation Code Sequence field (0054,0410) and Patient Gantry Relationship Code field Sequence (0054,0414). This refers to head v. feet first and prone v. supine |
| Scanning Direction | Acquisition Device | Shall be able to record the scanning direction (craniocaudal v. caudocranial) |
| Documentation of Exam Specification | Acquisition Device | The Acquisition Device shall record the z-axis FOV which represents the actual distance of scan anatomic coverage (cms) as well as the number of bed positions. Shall be able to record and define the x-y axis FOV. |
|  |  | Shall be able to define the extent of anatomic coverage based on distance from defined landmark site (e.g. vertex, EAM). (both the landmark location (anatomically) and the distance scanned from landmark) would require DICOM tags) Shall be able to be reportable for future scanning session |
| Bed Position Temporal Differences |  | Shall be able to provide and document non uniform scan times for different bed positions dependent upon areas of clinical concern |
| DICOM Writing Compliance | Acquisition Device |  |

| DICOM Editing | Acquisition Device | Shall be able to edit all fields relevant for SUV calculation and blood glucose before image distribution from scanner. |
| --- | --- | --- |

## 9. Reconstruction Software

Reconstruction Software shall propagate the information collected at the prior Subject Handling and Imaging Acquisition stages and extend it with those items noted in the Reconstruction section.

| Metadata Mapping | Reconstruction Software | Shall be able to accurately propagate the information collected at the prior stages and extend it with those items noted in Reconstruction section |
| --- | --- | --- |
|  | Reconstruction SW | Shall provide the options/tools suitable for quantitative imaging as described below. |

Data can be reconstructed including all corrections needed for quantification as well as without scatter and attenuation correction. Analytical and iterative reconstruction methods should be applied.   PET/CT systems are equipped with a harmonized reconstruction protocol (NB not the algorithm or settings are harmonized, but the resulting specified quantitative image characteristics for specified phantoms/experiments) – SNM Image Reconstruction Workgroup.

- Standardization of reconstruction settings is necessary to obtain comparable resolution and SUV recoveries across the same subject and inter-subject across sites < model specific issues to allow multicenter convergence - reference in template to 7.3.1 and Appendix??>

| Data Corrections | Reconstruction Software | PET emission data must be able to be corrected for geometrical response and detector efficiency, system dead time, random coincidences, scatter and attenuation |
| --- | --- | --- |
| Reconstruction Methodology | Reconstruction Software | Shall be able to provide both iterative and analytical (e.g. filtered back projection) reconstruction algorithms |
|  |  | Shall be able to ‘turn off’ resolution recovery, TOF (if capability available) for purposes of reconstruction |
| Reconstruction Methodology / Output | Reconstruction Software | Shall be able to perform reconstructions with and without scatter and/or attenuation correction |
| Data Reconstruction 2D/3D Compatibility | Reconstruction Software | Shall be able to perform reconstruction of data acquired in 3D mode using 3D algorithm or if data re-binned into 2D, can be reconstructed with a 2D algorithm |

## 10. Image Analysis Workstation

The image analysis workstation shall have the ability to receive and propagate the data output (imaging and metadata) collected from the prior activities (Subject Handling, Image Acquisition, Reconstruction and Post-Processing).  With the input data, the analysis workstation (and software analysis tools) will be able to make use of certain attribute values to perform certain measurements and computational analysis.  The analysis workstation and software may be coupled to the PET/CT scanner system or provided by a 3rd-party vendor.

| Metadata Mapping | Image Analysis Workstation | Shall receive and propagate . . . |
| --- | --- | --- |

Input for Image Analysis is considered output of Reconstruction and Post processing software activity.  If the Image Analyst  alters input data (e.g. zoom) this is considered part of Image Analysis activity.  If this occurs, the original input data will be maintained as separate file, both to be stored. (See Section 3.2)

MAKE NORMATIVE

A check should be performed by receiving devices that the Series Time field (0008,0031) is not later than the earliest Acquisition Time field (0008,0032) of all images in the Series in case the images have been derived, in which case the earliest Acquisition Time field (0008,0032) is the reference time for decay correction.

10.1 Voxel and Sphere specification

The use of partial voxel values to secure a 1.2cm diameter (1cc volume) ROI/VOI was suggested to be appropriate and desirable, since standard voxel sizes would not allow selection of a 1 cc volume precisely in most cases.

| Voxel Inclusion Policy | Analysis Tool | Shall describe methodology describing voxel inclusion and weighting policy |
| --- | --- | --- |
|  |  | Use weighting for partial voxels; fully included voxels use weight of 1.0. Weighting should be proportionate to volumes of voxels which are partly included. |
| Sphere Specification | Analysis Tool | Shall describe capabilities and limits of sphere specification and placement. |
|  |  | Ideally, VOI dimensions (sphere) would be specifiable to +/- 1 mm and to volumes to +/- 1%. Methods to determine the highest peak must iterate to most intense 1cc volume VOI in a candidate VOI . Search must, at a minimum, evaluate candidate peak regions centered at all voxels in an analyst selected region. |

10.2 ROI/VOI

The scanner-display-analysis system shall provide a tool for the user to define a 3D volume of interest around the SUVmax test object both in clinical subject datasets and in the QIBA supported digital reference object (DRO). The VOI should fully encompass the entire test object in the DRO in all 3 directions by at least 2 voxel dimensions or 1 cm. ROI/VOI and analysis software is responsible for correct implementation of SUVmax extraction and calculation. The manufacturer should verify that maximum voxel value and/or SUVmax represents the activity concentration of a single voxel having the highest value within a 3D VOI. Moreover the manufacturer is responsible for correct implementation of SUV calculations, with e.g. correct implementation of corrections for decay. Moreover, the manufacturer should implement various versions of SUV normalizations (body weight, body surface area and lean body mass). Finally the manufacturer should verify that the same methodologies for obtaining the maximum voxel value (Bq/ml) and SUVmax from the DRO are available and can be applied to clinical whole body data.

| Quantification Computation | Analysis Tool | Tool specs |
| --- | --- | --- |
| ROI/VOI Computation | Analysis Tool – Image Analyst | Shall provide a tool and user strategy, the combination of which should allow the placement of an ROI/VOI which allows determination of the value and location of a max voxel. Further, shall allow determination of the mean SUV in a "1 cm" diameter VOI manually placed on the hottest area of a selected region to approximate SUV peak. |
|  | Analysis Tool | Ideally, would provide a tool which is capable of automatically determining the location and values of both SUV max and the SUV peak VOI in a selected region. |
| Edge/Volume Detection | Analysis Tool | Threshold methods: fixed value, % of max or peak, gradient. Manufacturer must specify. |

10.3 Calculation of SUV

The workstation and repository shall be able to create, store and retrieve markups (ROI/VOI) of SUV measurements in accordance with a standard definition for ROI/VOI that provides a known balance between precision and accuracy.  For a given voxel cluster . . .

| ROI/VOI saving/retrieve | Analysis/Archival | Shall have the capability to label, save, recall ROI/VOI using DICOM structured sets. |
| --- | --- | --- |
| ROI/VOI saving/retrieve | Analysis/Archival | Ability to duplicate, save VOI and apply for subsequent analyses. Capability to label, save, recall (in a std format allowing visualization and portability across workstation) |

10.4 Required Characteristics of Resulting Data (UPICT Section 9.3)

The scanner, viewing and analysis system combination shall produce an output image data set from which actual maximum activity concentration or voxel value (units Bq/ml) and maximum SUV (bw, lbm) measures can be determined.  The maximum SUV(bw or lbm) represents the activity concentration value of a single voxel from the actual original image that has the highest value within a constrained volume of interest drawn around the object of interest.  ?The latter will avoid additional error in activity concentration values and SUVmax due to viewing/display interpolation and/or minimize round-off errors in SUV calculation within defined tolerances/ROI (or VOI) tool to be utilized to define lesion constraint condition and strategy to define edge detection; region growing versus based on threshold condition (% of max / absolute threshold / gradient sampling/ iterative / fuzzy clustering)

Patient meta-data recorded, e.g. using DICOM attributes

Tumor ROI's reflecting the metabolic volume of the tumor is desirable. Volumes based on a 50% and (?) 70% threshold of the peak tumor SUV should be produced. Alternative is to use the PERCIST threshold at 2SD above mean liver activity as a floor value for the ROI's. (?) These can reflect tumor volume and tumor volume x mean SUV lean - viewed as exploratory.

| Activity Concentration | Acquisition Device | The scanner shall produce an output image data set from which actual activity concentration and the different SUV/SUL can be determined |
| --- | --- | --- |
| Lesion Analysis: Single Voxel | Analysis Tool | Shall provide tools to determine hottest voxel measurement using body weight (SUVmax) and lean body mass (SULmax) |
| Lesion Analysis: Multiple Voxels | Analysis Tool | Shall provide tools to measure and report SUV/SUL mean |
|  |  | Ideally, would provide tools to measure and report SUV/SUL mean and SUV/SULpeak |
| Lesion Analysis Tool: Size | Analysis Tool | Shall provide tools to measure and report size assessment of ROI/VOI(e.g. number of voxels and cc/ml (volume)) |
| ROI/VOI tool and usage | Analysis Tool | ROI (or VOI) shall be provided and shall be utilized to define lesion constraint condition and strategy to define edge detection; region-growing versus based on threshold condition (% of max / absolute threshold / gradient sampling/ iterative / fuzzy clustering |
| Output Statistics | Analysis Tool | Shall provide standard uptake statistics to include single voxel SUV & SUL; and multislice (ROI/VOI) max, mean (with std dev) |
|  |  | Shall be able to provide peak SUV & SUL data |
| Rounding Convention | Analysis Tool | Shall provide capability to display and record standard uptake statistic to 3 significant digits |
| Preservation of the orignal quantitaitvie data | Analysis Workstation | Shall quantify directly from the originally reconstructed voxels (not interpolated and/or zoomed) |

Displays should have ability to show information that affects SUVs (uptake time, etc.)

10.5 Performance Specifications

The digital reference object (DRO), which is a noise-free image with no smoothing, shall be used in order to evaluate compliance to the level of performance of analysis station/display station. Users should use the DRO (as per the DRO user's guide) in order to verify correct implementation of maximum voxel extraction and SUV calculation.

| Performance Evaluation | Analysis Workstation | Shall use the version? Of DRO to achieve threshold of . . . |
| --- | --- | --- |
| Analysis Accuracy | Analysis Workstation/SW | Shall reproduce the known SUV max with a threshold of .01? to expectation to XX number of digits (precision) and other measurements (range of other tests) as outlined in the DRO |
| Other Activities TBD | Analysis Workstation/SW | Specifics to be determined |
| DICOM Compliance Reading Check? |  |  |

[output parameters relative to a DRO have to be within tolerances specified in the details section above]

## 11. Software version tracking

At a minimum, Software Version should be manually recorded during the qualification along with the phantom imaging performance data and the record should be updated for every Software-upgrade duration the trial. This includes the  flagging of the impact on quantification for now; in the future, record all software version numbers in DICOM header..

Parameter	Entity/Actor	Normative text

| Hardware Version tracking | Acquisition Recon Workstation Technologist | A full phantom imaging qualification shall be done after any key Hardware upgrades. QIBA digital reference phantom to verify compliance. Will record scanner used |
| --- | --- | --- |
| Software Version tracking | Acquisition Device | Shall record the software (what about hardware?) version used for subject image acquisition in DICOM tag |
| Workstation Hardware and/or Software upgrade/change | Image Analyst | QIBA digital reference phantom to verify compliance. Will record version of software used for analysis |
| Software Version tracking | Technologist | Shall record the software version used for subject image reconstruction in relevant trial-specific documentation such as CRFs or practice specific documentation. |
|  | Reconstruction SW | Shall record the software version used for subject image reconstruction in DICOM tag (s)--need additional DICOM |
| Software Version tracking Image Analysis | Technologist | Shall systematically record the software version used for subject image analysis in relevant trial-specific documentation such as CRFs or practice specific documentation |
|  | Workstation | Shall record the software version used for subject image reconstruction in DICOM tag (s)--need additional DICOM |
| Software version back-testing compatibility | Workstation | Shall provide mechanism to provide analysis of the image data using updated as well as prior (platform-specific?) versions of analysis software. |

# References

**fill in Extraction performed from Hallet / EANM/ EORTC / Netherlands/NCI/PERCIST/Delbeke

Medicare National Coverage Determinations Manual Chapter 1, Part 4 (Sections 200 – 310.1) Coverage Determinations (Rev. 142, 02-03-12)

http://www.cms.gov/Regulations-and-Guidance/Guidance/Manuals/downloads//ncd103c1_part4.pdf Referenced 19 April 2012.

UPICT Paper

KInahan and Fletcher

Adams et al.

# Appendices

## Appendix A: Acknowledgements and Attributions

This document is proffered by the Radiological Society of North America (RSNA) Quantitative Imaging Biomarker Alliance (QIBA) FDG-PET/CT Technical Committee.  The FDG-PET/CT Technical Committee is composed of physicians, scientists, engineers and statisticians representing the imaging device manufacturers, image analysis software developers, image analysis facilities and laboratories, biopharmaceutical companies, academic institutions, government research organizations, professional societies, and regulatory agencies, among others.  A more detailed description of the QIBA FDG-PET/CT group and its work can be found at the following web link:  Insert link

The FDG-PET/CT Technical Committee (in alphabetical order):

List members here

The FDG-PET/CT Technical Committee is deeply grateful for the support and technical assistance provided by the staff of the Radiological Society of North America.

Appendix B: Background Information

## Appendix C: Conventions and Definitions

Appendix D: Standard template for patient preparation and scan acquisition

- Needed information (in order from least frequently changing to most frequently changing)

- Should also identify which information needs a DICOM field (e.g. blood glucose level)

- Site specific

  - Site information

  - Scanner make and model

  - Hardware Version numbers

- Protocol specific

  - Duration per bed

  - CT technique

  - Reconstruction method

- Scan specific

  - Most recent calibration factors (scanner)

  - Scanner daily check values

  - Software Version numbers

  - most recent clock check

  - most recent scanner QA/QC

- Patient specific

  - Height

  - Weight

  - Blood glucose

  - Pre- and post-injection doses, times

  - Uptake time

  - Were above measurements made directly or estimated? (e.g. if patient cannot be moved from bed). If the latter was the date and source?

## Appendix G: Model-specific Instructions and Parameters

The presence of specific product models/versions in the following tables should not be taken to imply that those products are fully compliant with the QIBA Profile. Compliance with a Profile involves meeting a variety of requirements of which operating by these parameters is just one. To determine if a product (and a specific model/version of that product) is compliant, please refer to the QIBA Conformance Document for that product.

G.1. Image Acquisition Parameters. The following technique tables list acquisition parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

These technique tables may have been prepared by the submitter of this imaging protocol document, the clinical trial organizer, the vendor of the equipment, and/or some other source. (Consequently, a given model/version may appear in more than one table.) The source is listed at the top of each table.

Sites using models listed here are encouraged to consider using these parameters for both simplicity and consistency. Sites using models not listed here may be able to devise their own acquisition parameters that result in data meeting the requirements of Section 7.1 and conform to the considerations in Section 13.

In some cases, parameter sets may be available as an electronic file for direct implementation on the imaging platform.

G.2. Quality Assurance Procedures

Examples of recommend quality assurance procedures are shown for specific GE, Philips, and Siemens PET/CT scanners are listed in the tables below.

# Comments

Update when done

add references and analysis as an appendix

definitions needed

Add references to QIBA / MITA position papers.

Need to clarify with QIBA SC

1. what can we say here

2. how is compliance tested

update

Add References.

ESP: Multiple references were provided by RW and SW post 04May telecon in emails.

who's job to validate? More appropriate in Section IV

update

update

make consistent

Move to overal scan QA/QC & reword

Add normative text

From 12-23 call: In each section where it matters to record DICOM attribute, should insert in the appropriate section throughout section 3.  This should be combined with CDFM and DICOM attributes

Should this be normative text in the Compliance section (done by the scanner and the responsibility of the scanner manufacturer)?

Minimum is intra-reader variability study needed. If a 2 Reader paradigm, then inter-reader variability is needed as well.  Lesion quantitative vs. interpretation; target lesion choice; calling new lesion – additional investigation needed.  Currently unclear which statistics to evaluate and what should it trigger.  The following requirements are placed on QC associated with interpretation.

update with CQIE

this wording was cut/paste from Section 6.1.4 – belongs here.

Ling to check if there is a precise definition

clean up

Eric to check with DC and JW

Is this qualification for the individual protocol, or the Profile?

Who does the qualification ?

I agree this should probably go to section IV (all highlighted subsections

This was cut from here and paste into Section 6.1.1

Delete from this section and paste into the following Section.

This wording was cut/paste from Section 6.1.4

This wording is cut from this section and paste into following section.

This wording is cut/paste from previous section.

Future NIST-traceable cross calibration

Overlap with UPICT can be used

references

Activities that go on regardless of patient acqusitions, as per site acrreditation etc.

test this Jeff, Paul , Tim

Ask Josh schuerman

Need coordination with DICOM  be sure DICOM is using correct unit

look up

make consistant throughout

ask to Correct in DICOM wish list

from SW - does thids mean axial FOV?  What’s the DICOM field/tag number?  Also,need to include the number of bed positions.

This should go on DICOM wish list

DICOM request – this may already be captured.

Inserted on 12-23 tcon. Review esp notes.

May be able to delete this row.

From Anne Smith

Is there more detail on this statement here in this document, and if so, could we reference it?  This point I think is very important for vendors to understand.

Copy wording from above section into here as well

Should we describe use of voxel appropriate shape – e.g. rectangular rather than spherical – is this future groundwork activity?

references

redudant?

redundant

redundant

From 12-23 tcon: Is this duplication of Lesion Analysis: Multiple Pixels/Voxels and therefore can be deleted or ‘output’ rather than tool and therefore should remain?

testing registration(?) and. testing interpretation of DICOM information

Put reference to DRO specific threshold values - DRO user's guide as stand-a-lone  document?? but the specific item as appendix

consider insertion of background region to evaluate fundamental registration issue as possible reason for error as opposed to a DICOM interpretation issue.

There may be a bifurcation between the clinical trials and the clinical practice worlds

These may be elements which could be reported in structured report

separate hardware / software or distinct? model number can be captured in DICOM?

Specifc Model number for each manufacturer and specific software VERSIONS. ?Action item for manufacturers to provide

From Ling - Whether vendor has the software version in the DICOM header or not, the QIBA should develop a phantom imaging qualification sheet that includes places to record the hardware change and software change SINCE the beginning of the trial along with the key imaging performance parameters of this phantom. The key imaging parameters are resolution, biases, variance, uniformity. This sheet should be updated after every phantom imaging.

Would insert manufacturer specific QC tables (from JY)

There is no Section 7.1.