Mapping of Brain Regions using Blood Oxygenation Level Dependent (BOLD) functional MRI as a Presurgical Assessment Tool.

Version 1.0

February 12, 2013

Table of Contents

1. Executive Summary	2

2. Clinical Context and Claims	2

3. Profile Details	2

3.1. Subject Handling	2

3.2 Imaging Data Acquisition	2

3.3 BOLD Image Data Post Acquisition Processing	2

3.4 Parametric Image formation	2

3.5 Parametric Image Analysis	2

4. Compliance	2

4.1 Acquisition - Scanner	2

4.2 Acquisition - Peripheral Equipment	2

4.3 Software Analysis	2

4.4 Performing Site	2

4.	References	2

5.	Appendices	2

Appendix A: Acknowledgements and Attributions	2

Appendix B: Conventions and Definitions	2

Appendix C: Task Paradigm Specifications	2

Appendix D:	2

# 1. Executive Summary

This Profile has been developed to provide a systematic approach for optimizing Blood Oxygen Level Dependent (BOLD) fMRI brain mapping for treatment planning prior to surgery or invasive treatment intervention. Whereas the primary purpose of this Profile development is for individual patient care, application of the best practice guidelines it creates has application to clinical trials as well.

Task-induced BOLD fMRI can be used clinically as a biomarker for functionally eloquent brain tissue that might be at risk of damage from invasive procedures used to treat brain cancer or other focal pathologies (ref).  The clinical utility and professional acceptance of BOLD as a biomarker is dependent on the reproducibility and validity of task-induced BOLD response patterns - the primary measure produced by BOLD exams and from which secondary quantitative measures are derived (ref). Current methodology is quite variable at all stages from exam administration, data acquisition, analysis and report of results, and can best be described by a model of integration across multiple data acquisition systems, MR and data analysis platforms.  To address reproducibility we take into account the degree to which variability in methodological approach (e.g., patient training, data acquisition methods, data analysis approaches and devices employed) impact the accuracy and specificity of readout measures derived. The current priority of the QIBA BOLD fMRI Technical Committee is to characterize the current state of the art and to identify sources of variability in methodology which contribute significantly to variance and negatively affect quantitative measures derived. If we can reduce the variability associated with the methodological approach we can improve reproducibility and the quantitative value of fMRI as a biomarker.

Our initial studies provide quantitative measures of BOLD signal reproducibility that will be used in the statement of claims included as part of QIBA BOLD fMRI Profile 1.0. The same results will be provided to the scientific community at large in order to fill a critical gap in existing knowledge about BOLD fMRI reproducibility as assessed using quantitative measures that are particularly relevant for clinical use in pre-surgical planning.

This QIBA BOLD Profile 1.0 is expected to provide specifications that may be adopted by users as well as equipment developers (hardware and software devices) to meet targeted levels of clinical performance in identified settings. This profile makes claims about the precision with which hemodynamic response in eloquent cortex can be measured and displayed under a set of defined image acquisition, processing, and analysis conditions.

The intended audience of this document is:

- Technical staff of vendors planning to participate in the QIBA initiative

- Practicing clinicians at healthcare institutions considering appropriate specifications for acquiring equipment

- Experts involved in quantitative medical image analysis

- Anyone interested in the technical and clinical aspects of medical imaging

# 2. Clinical Context and Claims

BOLD fMRI is used as a tool for pre-treatment planning in individual patients with brain lesions, including tumors, vascular malformation and epileptogenic foci. The presenting symptoms and location of the affected brain tissue determine the particular region or regions of the brain to be mapped and the behavioral paradigm(s) selected (e.g. motor task, language task) to evoke a change in BOLD response. The change in BOLD signal (relative to a control condition) provides information about the brain region(s) involved in task performance and about the proximity of this eloquent cortex to brain site(s) to be treated. Endpoints that will influence treatment planning include risk assessment (impact of treatment on functioning cortex, e.g. surgical or radiation induced damage) and predictive value estimation (will damage to eloquent tissue result in a deficit). The goal of this profile is to specify the procedures and quantitative parameters under which BOLD fMRI is an accurate and reliable predictor of brain function, that is, as a valid imaging biomarker for medically meaningful changes in brain activity elicited by a particular task.

The clinical context sets out the utilities and endpoints for presurgical mapping cases and then proceeds to identify targeted levels of quality for named measurement read-outs that may be used in the relevant clinical indications. A brief description of the technique and assumptions is provided below:

Assumpton 1: Increased BOLD signal repeatably and appropriately timed with a stimulus is a valid indicator (biomarker) of hemodynamic response to that stimulus.  Hemodynamic response is an indicator of brain function in individual patients.

Assumption 2: Increased BOLD signal within brain area X produced by paradigm Y is a valid indicator of the function of area X (which for presurgical mapping could be further restricted to mean that excision or damage of the area could produce a neurological deficit.)

Repeated clinical experience, though anecdotal, is that these assumptions are valid.* Provide brief summary of evidence and provide appendix of extensive review of literature in support of assumptions.

Claims characterizing reproducibility of BOLD response

- Biomarker measurand: Hemodynamic response to change in brain activity – commonly referred to as BOLD fMRI signal (the biomarker is a measurable physical property)

  - Context of use: Preoperative fMRI mapping of eloquent cortex for treatment planning

    - Cross-sectional measurement: Localization of BOLD signal as index of eloquent cortex (motor, language, and/or visual cortical areas)

      - Index: the center of mass of activation of a focus of interest

- Bias Profile:

- Precision profile:

    - On a test-retest basis, the center of mass of activation of a focus of interest can be determined with a 5mm repeatability coefficient

      - Index: the spatial extent half-maximum border of activation clusters

- Bias Profile:

- Precision profile:

    - On a test-retest basis the spatial extent half-maximum border of activation clusters can be determined with a 10mm repeatability coefficient

      - Index: the relative magnitude of activation in homologous regions across hemispheres

- Bias Profile:

- Precision profile:

    - On a test-retest basis, the relative magnitude of activation in homologous regions across hemispheres can be determined with a 10% repeatability coefficient

- For each index, should also indicate Reproducibility (Intra-class Correlation Coefficient [ICC]; Concordance Correlation Coefficient [CCC], Reproducibility Coefficient [RDC]):

  - Specify conditions, e.g.,

    - Measuring System variability (hardware and software)

    - Site variability

    - Operator variability (Intra- or Inter-reader)

    - Time interval (across days/weeks etc)

    - Longitudinal change measurement (if specified)

      - List Indices: (as above, including sub-parts)

      - (Define reproducibility/repeatiblity terms) as part of an appendix – Erich to add language

Utilities and Endpoints for Clinical Trials

**Describe one or more utilities or endpoints this Imaging Protocol could serve in a Clinical Trial. (e.g. to determine eligibility of potential subjects in the clinical trial; to triage eligible subjects into cohorts based on stage or severity of disease; to assess response to treatment; to establish the presence of progression for determining TTP, PFS, etc.; to monitor for adverse events; to establish a database for the development, optimization, and validation of imaging biomarkers, etc.)

# 3. Profile Details

The Profile is documented in terms of “Actors” performing “Activities”.

Equipment, software, staff or sites may claim conformance to this Profile as one or more of the “Actors” in the following table.  Compliant Actors shall support the listed Activities by meeting all requirements in the referenced Section.  Failing to comply with a “shall” is a protocol deviation.  Although deviations invalidate the Profile Claim, such deviations may be reasonable and unavoidable as discussed below.

Table 1: Actors and Required Activities

| Actor | Activity | Section |
| --- | --- | --- |
| Acquisition Device and its providers | Subject Handling | 3.1. |
|  | Image Data Acquisition | 3.2. |
| Technologist | Subject Handling | 3.1. |
|  | Image Data Acquisition | 3.2. |
|  | Image Data Reconstruction | 3.3. |
| Physician or Scientist | Subject Handling | 3.1. |
|  | Image Analysis | 3.4. |
| Reconstruction Software | Image Data Reconstruction | 3.3. |
| Image Analysis Tool | Image Analysis | 3.4. |

The sequencing of Activities specified in this Profile are shown in Figure 1: Feroze new illustration

The requirements in this Profile do not codify a Standard of Care; they only provide guidance intended to achieve the stated Claims.  Although deviating from the specifications in this Profile may invalidate the Profile Claims, the radiologist or supervising physician is expected to do so when required by the best interest of the patient or research subject.  How study sponsors and others decide to handle deviations for their own purposes is entirely up to them.

The requirements included herein are intended to establish a baseline level of capabilities. Providing higher performance or advanced capabilities is both allowed and encouraged and the profile is not intended to be limiting in any way with respect to capabilities. The intention is not to dictate implementation details.

It is assumed that the patient’s referring physician(s) will determine the appropriateness and utility of an fMRI exam based on the patient’s medical history, symptoms, treatment options, prognosis and other relevant information. It is further assumed that the physicians will anticipate the likelihood that an fMRI exam will provide information that will be useful to the assessment, diagnosis, treatment of the patient’s medical condition*. (In other words, the following procedures are not intended to specify the medical rationale for conducting an fMRI exam of a particular patient.)

## 3.1. Subject Handling

3.1.1 Timing Relative to Treatment

BOLD imaging exams are typically performed prior to interventional procedures such a surgery or radiation treatment.

3.1.2. Subject Selection Criteria related to Imaging

Local policies for contraindications for absolute MRI safety should be followed; definitions of relative and/or absolute contraindications to MRI are not within the scope of this document.

The QIBA fMRI committee acknowledges that there are potential and relative contraindications to MRI in patients suffering from claustrophobia. Methods for minimizing anxiety and/or discomfort are at the discretion of the physician caring for the patient.

3.1.3 Subject Preparation Prior to fMRI exam

3.1.3.1 Discussion

Patient skills and abilities, as well as associated pathology, influence task paradigm selection and performance expectations. These factors greatly influence the ability to quantify task-induced BOLD signal with specificity and sensitivity.  Pathology and resulting cognitive or motor deficits may not permit an accurate characterization of the BOLD signal in eloquent cortex in terms of quantification of amplitude, center of mass and/or spatial extent of signal change. Therefore, consistent use of criteria for patient skill assessment and task selection are required to meet the claims of this Profile.

The choice of paradigm should be such that the task should be simple to perform, and functionally specific. Behavioral responses should be monitored when possible. The task should be able to produce BOLD signal of sufficient amplitude following specifications below. The patient should be trained in advance so that they are familiar and comfortable with the task and performance expectations prior to entering the MR room. A description of the types of paradigms used to map sensorimotor, language, vision and auditory brain regions can be found in Appendix ?.

The following specifications are minimum requirements to meet Profile claims. Ideally, image acquisition procedures should be identical when collected over multiple time periods (i.e. pre and post-surgery) to enhance reproducibility of results, and the use of task paradigms should be consistent for all imaging time-points.

Radiologists and supervising physicians may modify tasks or relax certain criteria when required for the best interest of patients or research subjects, in which case BOLD signal change may still be measured but the measurements will not be subject to the Profile claims.

Recording type of paradigm, performance accuracy and modifications to paradigm by the Acquisition Device is recommended.  This may be by automatic interface with stimulus display devices in combination with text entry fields filled in by the Technologist.  Alternatively, the technologist may enter this information manually on a form that is scanned and included with the image data as a DICOM Secondary Capture image.

3.1.3.2 Specification

| Parameter | Specification |
| --- | --- |
| Creation of the Paradigm | The Physician/Scientist shall determine if the task paradigm is appropriate for the subject. Specifically, tasks should be selected to achieve ___signal change within _____ minutes task to produce an tSNR> ___%. (mean tSNR at a minimum level) |

(Multiple paradigms specific to pathology)

3.1.4 Subject Ability Assessment and Training

3.1.4.1 Discussion

Consistent training and assessment avoids performance anxiety and/or poor performance which negatively affect exam results. It is important to provide information to the patient regarding the flow of the exam (e.g. order of the tasks, what can be expected in terms of time for each paradigm administered). If the patient has never been in the MR, the technologist should review what can be expected in terms of noise, discomfort, etc.

Once positioned in the MR, a quick review of the task is recommended to be sure that the patient is still familiar with what they will see or hear, and what they are asked to do during the task.

Recording patient performance or interruptions in performance in the patient file is helpful for auditing and interpreting characteristics of signal change during image analysis.

3.1.4.2 Specification

| Parameter | Specification |
| --- | --- |
| Subject Training | The Technologist shall train the subject on all task paradigms they are expected to perform during the exam and observe/record performance. If the patient is unable to perform the task successfully, a task that is more appropriately matched to patient ability should be selected and practiced. |

3.1.5 Subject Positioning

3.1.5.1 Discussion

Consistent positioning avoids unnecessary changes in attenuation, changes in gravity induced shape and fluid distribution, or changes in anatomical shape due to posture, contortion, etc.  Significant details of subject positioning include the position of their arms, the anterior-to-posterior curvature of their spines as determined by pillows under their backs or knees, the lateral straightness of their spines. When the patient is supine, the use of positioning wedges under the knees and head is recommended so that the lumbar lordosis is straightened and the scapulae are both in contact with the table. However, the exact size, shape, etc. of the pillows is not expected to significantly impact the Profile Claim.  It is expected that clinical trial documentation or local clinical practice will specify their preferred patient positioning.

Recording the Subject Positioning and Table Heights in the image header is helpful for auditing and repeating baseline characteristics.

3.1.5.2 Specification

| Parameter | Specification |
| --- | --- |
| Subject Positioning | The Technologist shall position the subject consistent with baseline. If baseline positioning is unknown, position the subject Supine if possible, with devices such as positioning wedges placed as described above. |
| Table Height & Centering | The Technologist shall adjust the table height for the mid-axillary plane to pass through the isocenter. The Technologist shall position the patient such that the “sagittal laser line” lies along the sternum (e.g. from the suprasternal notch to the xiphoid process). |

3.1.6 Subject Positioning w/respect to Peripheral Stimulus Delivery Equipment

3.1.6.1 Discussion

Comfortable positioning avoids unnecessary fluctuations in comfort, attention and/or performance.  Significant details of subject positioning include visual acuity correction to assure clarity of visual stimulus, as well as adjustment of volume via auditory stimulus delivery systems.  It is expected that clinical trial documentation or local clinical practice will outline the method for determining that stimulus delivery equipment recommendations outlined by the manufacturer/vendor are followed.

It is critical to properly adjust the fMRI stimulus presentation devices (e.g., goggle device or mirror on or above the head coil) to correctly adjust for visual acuity and ensure the entire visual field is visible this minimizes squinting and movement of the eyes/head. If stimuli are presented aurally, placement and adjustment of headphones is important for establishing appropriate volume control. For monitoring responses additional stimulus response devices such as a MR compatible button boxes, grip devices or trackballs should be positioned such that the patient is able to operate the device easily and without hindrance. It is advisable to use foam padding to reduce head motion, and use foam ear plugs to reduce interference from scanner noise when a proper audio presentation system is not used. Details regarding types of peripheral equipment alternatives are described in section 3.2.

3.1.6.2 Specification

| Parameter | Specification |
| --- | --- |
| Peripheral Equipment Adjustment | The Technologist shall adjust the peripheral equipment such that the patient is able to see stimuli clearly and easily through the visual stimulus delivery equipment and to hear stimuli via audio system device. Response devices should be placed appropriately before the beginning of the exam. |

## 3.2 Imaging Data Acquisition

MRI scans for fMRI analysis will be performed on qualified equipment. It is recommended to use a field strength of 1.5 Tesla and higher with echo planar capabilities. Once the patient is positioned inside the scanner within the head coil it is good practice for the MRI technician to provide brief instruction to the patient about the task(s), as a reminder of what they are expected to do. It is recommended to have appropriate personnel present during the scan to meet CPT code requirements. The MRI scan starts with a localizer followed by T1 or T2 scans to cover the whole brain following the local imaging protocol. Following the anatomical image acquisition, a shim scan is followed by an fMRI BOLD series .The fMRI scan duration is defined by the paradigm design and the MR protocol should be configured and named with the same description used to name the stimulus paradigm.

Precise synchronization of stimulus presentation and image acquisition is highly recommended. The system to control the stimulus presentation can be a standalone workstation or PC with software for presenting stimulus paradigms or it can be software which is integrated within the MR console. As previously mentioned, the visual presentation of the stimulus can be displayed via MR compatible systems such as binocular goggle-based system, or projector-based system (LCD Monitor, or Projector). The audio stimulus can be presented using audio delivery systems provided with the MR or a third party systems which are MR compatible and specifically designed for presenting stimuli in the MR environment. Monitoring task performance (direct observation of eye movement, finger/hand/foot movement etc) as well as recording patient responses (button box or other devices to monitor patient performance) is highly recommended. Frequent communication between the patient and technician between scan series to assess comfort and attention, and to provide intermittent instruction is required.

It is highly recommended that the sites perform some or all of these quality assurances on their devices for improved and consistent data quality. They include routine SNR and fSNR measurements to test for signal and image quality, routine checks on the fMRI specific equipments such as the response buttons, projector, goggles, audio etc prior to the scan. Motion artifacts can significantly impede the quantitative fMRI outcome measures. Hence it is important to use head restrainers such as foam pads and provide reminders to the patient’s before the scan to reduce the motion inside the scanner while performing the test.

3.2.1 Scan Synchronization/Triggering requirements

3.2.1.1 Discussion

It is recommended to perform BOLD fMRI imaging with precise synchronization at the start of the scan and the start of the stimulus presentation using trigger pulses.

3.2.1.2 Specification

Acceptable: Manually trigger

Ideal: is to start scan within 500 ms of paradigm initiation using a synchronization device which captures the first signal from the MR EPI sequence onset.

The following recording requirements are noted: Actual Timing and Triggers shall be recorded in the DICOM Header

3.2.2 Visual Stimulus specifications

3.2.2.1 Discussion

The visual stimulus should be clearly seen and stable and should be of a specific size.

3.2.2.2 Specification

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Minimum Display size of _____________ Ideal Minimum Display size of _____________ |

3.2.3 Visualization/Monitoring of Task performance (SOV; Compliance Item)

3.2.3.1 Discussion

Visual monitoring of the patients during the performance of the task is recommended. This may aid in evaluating compliance of certain fMRI tasks such as motor tasks. It is also recommended to conduct an interview after the scan for patient compliance.  The recommended ways to monitorare described below.

3.2.3.2 Specification

2.7.1 Monitor Task performance

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Ideal |

3.2.4 Monitor Respiration (Decided to have a different category but don’ know where for the moment.

This is a huge source of variability, 3% on 3T.

3.2.5 Anatomical/Structural Images  (SOV; Compliance Item, Standardization needed)

3.2.5.1 Discussion

The BOLD T2* images are reconstructed on the scanner as individual images or as mosaics.  An fMRI series will typically consist of several measurement periods. Each individual measurement period will have a set of images corresponding to the anatomical coverage specified by the user (typically whole brain).  The total imaging time to acquire an fMRI series will depend on the repetition time (TR) and the number of measurement periods acquired throughout the series.

3.2.5.1 Specifications

The following parameters describe what the acquired images shall contain/cover.

| Parameter | Compliance Levels |
| --- | --- |
| Anatomic Coverage | Acceptable Coverage of Area of interest |
| Field of View | Acceptable Coverage of Area of interest Ideal Whole brain |
| Scan Duration Motor Task; Language Task; | Acceptable 2 min Ideal 3 min |
| Scan Plane (Image Orientation) | Acceptable Transverse or Axial Ideal Transverse or Axial |

The following recording requirements are noted:

| Parameter | Compliance Levels |
| --- | --- |
| Image Header | Acceptable Number of Measurement Periods; Actual Anatomic Coverage, Field of View, Scan Duration, and Scan Plane shall be recorded. |

3.2.6 NVU assessment as a screening test, how is it related to our claims?

3.2.6.1 Discussion

Breath hold scans (NVU)

3.2.6.2 Specifications

Data Quality Requirements

3.2.7 MRI Signal specifications (SNR, fSNR) (SOV; Compliance Item) (Leave it in the profile) ((Justification Required)

3.2.7.1 Discussion

3.2.7.2 Specifications

| Parameter |  | Compliance Levels |
| --- | --- | --- |
| 1.Spatial SNR |  | Acceptable Minimal artifact (Potentially corrected prospectively) Ideal No artifact |
| 2. Temporal SNR |  | Acceptable Ideal |
| 3. Subject Motion |  | Acceptable: Should be done immediately after data acquisition Ideal: To monitor in real time. |
| 5. Task dependant SNR |  | % Signal change measured is greater than _________ (It encompasses a number of items listed above) |
| 4. Task performance |  |  |

Spatial SNR measurement:

SNR: Mean signal of the Tissue 1/ Spatial Stdev of noise (air) (measured where?)

CNR: (Mean signal of Tissue 1 – Mean signal of Tissue 2) / Spatial Stdev of noise (air).

Coil Characteristics affect the BOLD signal. (field homogeneity/image bias fields)

Include in the DRO various pulse sequence parameters that might affect the SNR.

(Bandwidth, FOV, slice thickness, number of averages, flip angle)

spatial resolution (spatial frequency specification)

N/2 ghosting in EPI images (to follow ACR definitions and guidelines)

SAR measurements (to follow ACR definitions and guidelines)

geometrical distortions (EPI versus T1) (Susceptibility effects)

phantom – signal in sample/noise (acquire 2 images and calculate difference to get noise)

T1, T2, T2* weighted images

Human –

anatomical (T1, T2)

T2* - Measured on phantom (scanner manufacturers), measured on subject

Temporal SNR measurement:

Mean signal / Temporal Stdev after removing the drift (averaged over non-task regions?) (reference from Dominco) expected value = 70-100.

Measured on phantom (to measure temporal stability on a standard ACR phantom), measured on subject (It can task based data or resting state data, if task based data is used then signal regressed from the task based regions). The data should be motion corrected (rigid body alignment between various time points) and linear detrended before the TSNR measurements.

Subject motion measurement:

Translation, rotation of each image/volume as functions of time (threshold of what is good data or not?) (Another paper on quantitative measurement on motion limits acceptance, Dominoco).

Levels: mean motion, maximum motion, cumulative motion, percent of scan without motion

Task performance measurement:

Performance (graded or binary evaluation) as function of time

Levels: mean performance, percent of scan with acceptable performance (same as mean if binary)

Task-dependent SNR measurement:

Net task-correlated signal change/Temporal Stdev [t-value?] (in putative active regions?)

Temporal contrast to noise (peak to peak task-based signal/signal without task)

In Fourier domain: power at fundamental task freq/power in adjacent 2-3 freq bands

Levels:

3.2.8 Post/During Scan Requirements

3.2.8.1 Patient assessment (Compliance Issue)

| Parameter | Compliance Levels |
| --- | --- |
| Subject Interview (Post) | Acceptable Please check with the subject if they performed the task Ideal |
| Parameter | Compliance Levels |
| Technologist Evaluation | Acceptable Technologist’s evaluation of the subjects performance Ideal |

Real-time results evaluation

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Its is acceptable to evaluate the results offline Ideal Online real-time |

## 3.3 BOLD Image Data Post Acquisition Processing

The Post-acquisition processing and statistical analysis can be performed on the scanner provided software or on a standalone workstation offline. A variety of software’s and algorithms are available for this purpose (See Appendix). This section provides guidelines and recommends the following steps to obtain high quality and reliable color maps of the fMRI data.  The BOLD data is typically corrected for a low frequency signal drift, spatial smoothed to improve SNR, artifacts identified (manual or automatic) and corrected, corrected for slice timing (if event related design is used), motion identified and corrected, and coregistered with a T1 or T2 structural data. The registration parameters are typically saved in a file for later Q/A check to access the patient’s motion. A variety of statistical tools (GLM, non GLM) methods can be used to perform statistical analysis to create color functional maps. These maps are later overlayed onto the structural data or 3D maps created for better visualization by the end users. Individual maps pertaining to different paradigms are created. These maps can be saved in DICOM, generic formats on the scanner or off the scanner.

It is highly recommend to generate a technical report that includes the summary of the imaging procedure, patient performance of the task, qualitative and quantitative summary of the head motion, subjective assessments of artifacts and outliers, assessment of the data alignment (functional vs structural), pre, during and post scan evaluation of the patient, and neurovascular uncoupling of the patient.

The items listed below described below provides the general characteristics and requirements for fMRI post acquisition processing:

    - Anatomical Image Quality (SOV)

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable The scanner functions within its specification and gives good data. Draft some kind of specifications for this item. Ideal High resolution being able to identify gyri and sufficient to segment gray/whitee matter. |

3.3.1.1	Segmentation of Anatomical data

(Keep it in here but less importance at this point)

3.3.1.2 Intensity  nonuniformity correction of Anatomical data

(Keep it in here but less importance at this point)

3.3.5	Field Inhomogeneity Correction/Compensation (SOV) (Actors: Scanner/Software Manufacturer

| Parameter | Compliance Levels |
| --- | --- |
| Drift Specs | Acceptable Have to check for geometric distortion and minimize distortion using manual functions such as nudge and informing clinicians about the distortion. Ideal Field to be uniform and field mapping distortion correction software to be installed on the scanner. |

3.3.2	Patient Motion (SOV) (One of the most important SOV) (To be discussed in a separate meeting).

There are several factors that contribute to motion.

3.3.3	Motion Correction (SOV)

3.3.4	Coregistration (SOV)

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Gyri line up, anatomy between EPI and or T1/T2. Ideal Field map corrections needed, |

Cox paper.

3.3.5	Low Frequency drift correction (SOV; Compliance)

Low freq is less than ~0.016 Hz – need to identify acceptable scanner signal change at low freqs

3.3.5	High Frequency noise (SOV; Compliance)

High freq is greater than ~0.24 Hz – need to identify acceptable scanner signal change at low freqs

    - Slice timing correction – need to be aware

3.3.7.1	Spatial signal non-independence  (SOV, Compliance)

Non-independence of signals in neighboring voxels – how much MR signal spread is a problem?

Some PSD’s result in more spread (e.g. spiral); some scanners do more spatial smoothing

3.3.7.2   Spatial variability in BOLD HRF  (SOV, Compliance)

HRF’s can vary by region (e.g. due to vascular differences, pathology)

Variability can affect timing of HRF (faster/slower) or could affect shape

Could affect sign of HRF (ref Ted)

    - Statistical Model Specification (GLM/Correlation/tTest/??) (SOV; Compliance)

Big topic – includes all signal sources presumed to contribute to observed response, and how to extract brain activation signal of interest

To be pursued

    - Statistical map generation (SOV ??; Compliance, Standardization needed)

This is presumably the step of applying the above model to make a map.

Variability here may be mostly standardization.

    - Statistical map threshold (SOV ??; Compliance, Standardization needed)

Variability here may be mostly standardization

3.3.11	Overlay fMRI data onto Anatomical data (SOV ??)

Variability here may be mostly standardization

3.3.12	Functional map generation (Compliance)

3.3.13 Color Map Specifications (SOV; Compliance ??)

Variability here may be mostly standardization

3.3.14 Registration parameters

## 3.4 Parametric Image formation

## 3.5 Parametric Image Analysis

Color overlay characterization

- Review of color overlays

- Gyral/Sulcal location

- Gyral/Sulcal distance

- Lesion/Cluster margin characterization

- Characterization of Artifacts

- Confidence Index

- Documentation for CPT requirements

Technical Report Specifications

Color Images

Imaging procedure summary

Patient motion summary

Artifact summary

Patient evaluation summary

NVU summary

Clinical Report Structure

-Demographic information

- Indications

- Techniques

- Structural imaging findings

- Functional imaging findings

- Annotated images

- Clinical Impression

5. Storage and Distribution

5.1	Data Storage specification

  - Data sharing specification

# 4. Compliance

## 4.1 Acquisition - Scanner

Compliance to specifications as set out in the Image Acquisition section above.  Additionally, compliant Acquisition Devices shall provide means to record the information identified in the Subject Handling section as means to document compliance of the Performing Site to the specifications noted there.

## 4.2 Acquisition - Peripheral Equipment

Compliance to specifications as set out in the Image Reconstruction section above.  Additionally, compliant Reconstruction Software shall propagate the information collected at the prior Subject Handling and Imaging Acquisition stages and extend it with those items noted in the Reconstruction section.  See the compliance procedure notes associated with Acquisition Devices above for procedural assistance to identify Model Specific Parameters for Reconstruction Software.

## 4.3 Software Analysis

Compliance to specifications as set out in the Image Analysis section above.  Additionally, compliant Software Analysis Tools shall propagate the information collected at the prior Subject Handling, Imaging Acquisition, and Imaging Reconstruction stages and extend it with those items noted in the Analysis section.

4.2 Image processing software specifications (Compliance)

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable Ideal |

## 4.4 Performing Site

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large patient population under consideration.  For imaging it is important to consider the availability of:

- appropriate imaging equipment and quality control processes,

- appropriate peripheral equipment and contrast media,

- experienced MR technologists for the imaging procedure, and

- processes that assure imaging protocol compliant image generation.

A protocol specific calibration and QA program shall be designed consistent with the goals of the clinical trial. This program shall include (a) elements to verify that sites are performing the specified protocol correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address issues relating to both radiation dose and image quality (which may include issues relating to water calibration, uniformity, noise, spatial resolution -in the axial plane-, reconstructed slice thickness z-axis resolution, contrast scale, CT number calibration and others). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.

Staff requirements for fMRI

| Parameter | Compliance Levels |
| --- | --- |
| Specs | Acceptable MRI technologist trained in performing fMRI behavior protocols Ideal Please follow the description in the CPT code described below |

Description of currently approved CPT codes.

70554
Magnetic resonance imaging, brain, functional MRI; including test selection and administration of repetitive body part movement and/or visual stimulation, not requiring physician or psychologist administration   (Approved at 2.11 RVUs for physician work and a total of 2.81 RVUs for professional payment, 13.71 RVUs for technical payment adding up to 16.52 RVUs when billed globally)  70554 is not to be reported in conjunction with 96020 or 70555

70555 
Magnetic resonance imaging, brain, functional MRI; requiring physician or psychologist administration of entire neurofunctional testing   (Approved at 2.54 RVUs for physician work for a total of 3.37 RVUs)  70555 can only be reported when 96020 is performed.

96020 
Neurofunctional testing selection and administration during noninvasive imaging functional brain mapping, with test administered entirely by a physician or psychologist, with review of test results and report.  (Approved at 3.43 RVUs for physician work for a total of 4.46 RVUs.)  Do not report 70554 & 70555 in conjunction with 70551-53 unless a separate diagnostic MRI is performed.

# References

# Appendices

## Appendix A: Acknowledgements and Attributions

This imaging protocol is proffered by the Radiological Society of North America (RSNA) Quantitative Imaging Biomarker Alliance (QIBA) Volumetric Computed Tomography (v-CT) Technical Committee. The v-CT technical committee is composed of scientists representing the imaging device manufacturers, image analysis software developers, image analysis laboratories, biopharmaceutical industry, academia, government research organizations, professional societies, and regulatory agencies, among others. All work is classified as pre-competitive. A more detailed description of the v-CT group and its work can be found at the following web link: http://qibawiki.rsna.org/index.php?title=Volumetric_CT.

The Volumetric CT Technical Committee (in alphabetical order):

- Athelogou, M. Definiens AG

- Avila, R. Kitware, Inc.

- Beaumont, H. Median Technologies

The Volumetric CT Technical Committee is deeply grateful for the support and technical assistance provided by the staff of the Radiological Society of North America.

## Appendix B: Conventions and Definitions

Acquisition vs. Analysis vs. Interpretation: This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be computational steps that transform the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge. (The transformation of knowledge into wisdom is beyond the scope of this document.)

Bulls-eye Compliance Levels Acquisition parameter values and some other requirements in this protocol are specified using a “bulls-eye” approach. Three rings are considered from widest to narrowest with the following semantics:

ACCEPTABLE: failing to meet this specification will result in data that is likely unacceptable for the intended use of this protocol.

TARGET: meeting this specification is considered to be achievable with reasonable effort and equipment and is expected to provide better results than meeting the ACCEPTABLE specification.

IDEAL: meeting this specification may require unusual effort or equipment, but is expected to provide better results than meeting the TARGET.

An ACCEPTABLE value will always be provided for each parameter. When there is no reason to expect better results (e.g. in terms of higher image quality, greater consistency, lower dose, etc.), TARGET and IDEAL values are not provided.

Some protocols may need sites that perform at higher compliance levels do so consistently, so sites may be requested to declare their “level of compliance”. If a site declares they will operate at the TARGET level, they must achieve the TARGET specification whenever it is provided and the ACCEPTABLE specification when a TARGET specification is not provided. Similarly, if they declare IDEAL, they must achieve the IDEAL specification whenever it is provided, the TARGET specification where no IDEAL level is specified, and the ACCEPTABLE level for the rest.

Other Definitions:

Image Analysis, Image Review, and/or Read: Procedures and processes that culminate in the generation of imaging outcome measures, such tumor response criteria. Reviews can be performed for eligibility, safety or efficacy. The review paradigm may be context specific and dependent on the specific aims of a trial, the imaging technologies in play, and the stage of drug development, among other parameters.

Image Header: The Image Header is that part of the file or dataset containing the image other than the pixel data itself

Imaging Phantoms: Devices used for periodic testing and standardization of image acquisition. This testing must be site specific and equipment specific and conducted prior to the beginning of a trial (baseline), periodically during the trial and at the end of the trial.

Intra-Rater Variability is the variability in the interpretation of a set of images by the same reader after an adequate period of time inserted to reduce recall bias.

Inter-Rater Variability is the variability in the interpretation of a set of images by the different readers.

A Time Point is a discrete period during the course of a clinical trial when groups of imaging exams or clinical exams are scheduled as defined in the study protocol.

## Appendix C: Task Paradigm Specifications

## Appendix D:

Table G.1: Acquisition Device Model-specific Parameters Demonstrated to Achieve Compliance

IMPORTANT NOTE with respect to this example table: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  These settings were determined by the team in the 1C study as an example of how it could be done but more strict attention to all parameters identified in the Profile are necessary in order for a company to claim any particular model is compliant.  That said, we appreciate the good will and help that the vendors represented here have provided in this early phase of QIBA.

| Acquisition Device | Product Setting to Achieve Compliance Levels |
| --- | --- |
| GE Discovery HD750 sct3 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.625 Gantry Rotation Time in seconds 1 mA 120 Pitch 0.984 Scan FoV Large Body (500mm) |
| Philips Brilliance 16 IDT mx8000 | kVp 120 Number of Data Channels (N) 16 Width of Each Data Channel (T, in mm) 0.75 Gantry Rotation Time in seconds 0.75 Effective mAs 50 Pitch 1.0 Scan FoV 500 |
| Philips Brilliance 64 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.625 Gantry Rotation Time in seconds 0.5 Effective mAs 70 Pitch 0.798 Scan FoV 500 |
| Siemens Sensation 64 | kVp 120 Collimation (on Operator Console) 64 x 0.6 (Z-flying focal spot) Gantry Rotation Time in seconds 0.5 Effective mAs 100 Pitch 1.0 Scan FoV 500 |
| Toshiba Aquilion 64 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.5 Gantry Rotation Time in seconds 0.5 mA TBD Pitch .828 Scan FoV Medium and Large |

Table G.2: Reconstruction Software Model-specific Parameters Demonstrated to Achieve Compliance

IMPORTANT NOTE: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  These settings were determined by the team in the 1C study as an example of how it could be done but more strict attention to all parameters identified in the Profile are necessary in order for a company to claim any particular model is compliant.  That said, we appreciate the good will and help that the vendors represented here have provided in this early phase of QIBA.

| Reconstruction Software | Product Setting to Achieve Compliance Levels |
| --- | --- |
| GE Discovery HD750 sct3 | Reconstructed Slice Width, mm 1.25 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel STD |
| Philips Brilliance 16 IDT mx8000 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm (contiguous) Display FOV, mm 350 Recon kernel B |
| Philips Brilliance 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm (contiguous) Display FOV, mm 350 Recon kernel B |
| Siemens Sensation 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel B30 |
| Toshiba Aquilion 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm Display FOV, mm TBD Recon kernel FC12 |

Table G.3: Image Analysis Software Model-specific Parameters Demonstrated to Achieve Compliance

IMPORTANT NOTE: The presence of specific product models/versions in the following tables shall not be taken to imply that those products are fully compliant with the QIBA Profile.  In particular, the following example table only has placeholders for these example products which need to be replaced with product model-specific settings in order to claim compliance.

| Image Analysis Software | Product Setting to Achieve Compliance Levels |
| --- | --- |
| Siemens LunCARE | a <settings to achieve…> b <settings to achieve…> c <settings to achieve…> d <settings to achieve…> |
| GE Lung VCAR | e <settings to achieve…> f <settings to achieve…> g <settings to achieve…> h <settings to achieve…> |
| R2 ImageChecker CT Lung System | i <settings to achieve…> j <settings to achieve…> k <settings to achieve…> l <settings to achieve…> |
| Definiens (name specific product) | m <settings to achieve…> n <settings to achieve…> o <settings to achieve…> p <settings to achieve…> |
| Median (name specific product) | q <settings to achieve…> r <settings to achieve…> s <settings to achieve…> t <settings to achieve…> |
| Intio (name specific product) | u <settings to achieve…> v <settings to achieve…> w <settings to achieve…> x <settings to achieve…> |