QIBA Profile.  Computed Tomography: Lung Densitometry

Iowa Version

15-Aug-2011

Table of Contents

I. Executive Summary	1

II. Clinical Context and Claims	2

Utilities and Endpoints for Clinical Trials	2

Claim 1: Longitudinal Stability of Lung-equivalent phantom Densitometry	2

Claim 2: Longitudinal Stability of Lung Densitometry	2

III. Profile Details	3

1. Subject Handling	4-7

2. Image Data Acquisition	8

3. Image Data Reconstruction	9-10

4. Image Analysis	10-11

IV. Compliance	11

Acquisition Device	11

Reconstruction Software	11

Software Analysis Tool	11

Image Acquisition Site	11

References	12

Appendices	12

Acknowledgements and Attributions	12

Background Information	12

Conventions and Definitions	12

# I. Executive Summary

This document establishes the methods to estimate lung densities and lung volumes from CT images obtained during longitudinal studies.  Lung densities and lung volumes will serve as surrogate endpoints for severity of emphysema, degree of air trapping, and degree of hyperinflation in COPD and asthma.

QIBA has constructed a systematic approach for standardizing and qualifying densitometry as a biomarker. This QIBA Profile is expected to provide specifications that may be adopted by users as well as equipment developers to meet targeted levels of clinical performance in identified settings.

The intended audiences include:

- Technical staffs of software developers and device manufacturers who create products for this purpose

- Clinical trial scientists

- Practicing clinicians at healthcare institutions considering appropriate specifications for procuring new equipment

- Experts involved in quantitative medical image analysis

- Anyone interested in the technical and clinical aspects of medical imaging

Note that specifications stated as “requirements” here are only requirements to achieve the claim, not “requirements on standard of care.”   Specifically, meeting the goals of the profile are secondary to properly caring for the patient.

# II. Clinical Context and Claims

## Utilities and Endpoints for Clinical Trials

This imaging protocol will be able to establish three endpoints to serve a clinical trial:

###### 1. It will establish the severity of emphysema.

Define specific metrics below. Emphysema metrics are global measurements based on CT attenuation histogram. Methods used for calculation of extent of emphysema include the density mask technique, where the percent emphysema is calculated based on the percentage of lung voxels with CT attenuation below a given threshold. The threshold used for extent of emphysema has varied from -900 to -970 HU. Correlation with histologic and morphometric measures of emphysema suggests that the optimal threshold at total lung capacity, using thin-section densitometry CT, may be -960 or -970 HU.

The nth percentile method is defined as the threshold at which n percent of all voxels have a lower density. Therefore, the relation between the two methods (density maks and percentile) is defined exactly by the cumulative histogram and thus both parameters are closely correlated.. In most longitudinal studies on emphysema, the 15th percentile is chosen. (caveats- sliding scale, sensitivity vs specificity, scanner variability, slice thickness).

2. It will establish the degree of air trapping in the lungs.

Define specific metrics below. -856 HU at FRC or RV- pros and cons- limited data

3. It will establish inspiratory and expiratory lung volumes.

Lung volumes should be defined consistently during follow-up of a patient, as this significantly influences volume correction. Lung volumes can be defined as the total volume of voxels, including main bronchi and excluding the trachea (the main carina is a very consistent landmark), excluding vessels with a density higher than the lung detection threshold.

Define specific metrics below. TLC. Correlation with physiologic TLC. FRC/RV

(Do we need to discuss regional progression, like progression per lung lobe, or apical/basal progression?)

## Claim 1: Longitudinal Stability of Lung-equivalent phantom Densitometry

Using this profile, we will be able to achieve adequate calibration and longitudinal stability of CT scanners in order to permit measurement of longitudinal change in attenuation of air on CT phantom of ±1 HU. CT phantom measurements should be made at least once a month, and after every recalibration, hardware/software change, or CT service, to ensure long-term stability and precision of CT measurements. CT phantom should be optimized for measurement of air and lung equivalent material (COPDGene II phantom or equivalent phantom as approved by QIBA technical committee). Phantom should be scanned using the same CT parameters as used in the CT study, and at each available scan acquisition data collection diameter. While there are some data on short-term precision, there is little information on long-term precision of CT measurements of emphysema.

## Claim 2: Longitudinal Stability of Lung Densitometry

Claim regarding precision of in vivo patient measurement. Using this protocol, we will be able to measure severity of lung emphysema (measured using density mask technique at -950 HU) with repeatability of X%, or 5 HU shift in 15th percentile of lung attenuation (do we have data for this?) with repeatability of X%.

Great care should be taken to define the baseline measurements at inclusion of a CT scanner in a study. The scanner should be maintained fully, so that it can be guaranteed that this state can be reproduced when a shift or drift has been encountered during quality control in a later stage.

Compliance Levels for Measurement CT phantom measurements of air or lung equivalent material)

| Measurement Result | Performance Levels Achieved under Bull's Eye Conditions |
| --- | --- |
| CT attenuation (air inside phantom) | If Activities are Performed at Target Level -995 to -1005 HU |
| Longitudinal stability: CT attenuation (lung equivalent material) Deviation from baseline measurement | If Activities are Performed at Target Level ± 2 HU |

# III. Profile Details

## 1. Subject Handling

1.1 Contrast Preparation and Administration

Contrast agents are not to be used in the CT densitometric assessment of COPD or Asthma.

1.2 Subject Positioning

Discussion

Consistent positioning avoids unnecessary variance in attenuation, changes in gravity induced shape and fluid distribution, or changes in anatomical shape due to posture, contortion, etc. Significant details of subject positioning include the position of their upper extremities, the anterior-to-posterior curvature of their spines as determined by pillows under their backs or knees, the lateral straightness of their spines, and, if prone, the direction the head is turned.  Positioning the subject Supine/Arms Up/Feet first has the advantage of promoting consistency, and reducing cases where intravenous lines go through the gantry, which could introduce artifacts.

Specification

| Parameter | Specification |
| --- | --- |
| Subject Positioning | The Technologist should place patient in a supine position, arms positioned comfortably above the head in a head-arm rest, lower legs supported. Using the laser positioning lights, the Technologist must line up the patient so the chest is iso-center (in the middle: left-right; up-down) of the CT gantry. |
| Table Height | The Technologist shall adjust the table height to place the mid-axillary line at iso-center. |

1.3 Breathing Instructions during Acquisition

Discussion

Breath holding reduces motion that might degrade the image. Full inspiration at Total Lung Capacity (TLC) inflates the lungs so structures are optimally visualized including airways and lung parenchyma. Full expiration or Residual Volume (RV) decreases lung volume to a point at which air trapping may be measured.  Functional Residual Volume (FRC) is a possible alternative.

Rigorous adherence to instructing the patient how to breathe before the scan is obtained in order to attain TLC or RV volume before each chest CT is performed. These must be done verbally as the breathing should mimic breathing instructions typically provided in Pulmonary Function Laboratories.

Specification

| Parameter | Specification |
| --- | --- |
| Breath hold | One TLC and One RV scan is recommended. |

## 2. Image Data Acquisition

Discussion

CT scans for COPD and Asthma assessment, densitometry & airway measurements, will be performed on equipment that complies with the specifications set out in this profile.  At this stage of development, we continue to recommend that all CT scans for an individual participant be performed on the same platform throughout the trial. In the rare instance of equipment malfunction, follow-up scans on an individual participant can be performed on the same type of platform. All efforts should be made to have the follow-up scans performed with identical parameters as the first.

A set of scout images should be initially obtained. Pitch is chosen so as to allow completion of the scan in a single breath hold.

Faster scans shorten the scan time and reduce the breath hold requirements, thus reducing the likelihood of motion artifacts.

Total Collimation Width (defined as the total nominal beam width) is often not directly visible in the scanner interface. Wider collimation widths can increase coverage and shorten acquisition, but can introduce cone beam artifacts which may degrade image quality.  Slice Width directly affects voxel size along the subject z-axis. Smaller voxels are preferable to reduce partial volume effects and provide higher accuracy due to higher spatial resolution. Currently, the recommendation is to use the widest collimation and smallest slice width.

Scan FOV (SFOV) is a setting that must be determined during the scan acquisition phase only on certain model scanners. Smaller SFOV’s are typically better at reducing scan artifacts; however using a smaller SFOV may clip the lungs on larger subjects.  Therefore, it is recommended that the scan FOV is set to a large focal spot size and is consistent throughout the trial.

X-ray CT uses ionizing radiation.  Exposure to ionizing radiation increases known health risks to the subject.  It is recognized that there are tradeoffs between radiation dose and image quality.  As the radiation dose is reduced, image quality can be degraded.  Because of these tradeoffs, the use of the CT Dose Index Volume (CTDIvol) measurement is recommended to use as exposure control across all scanner models.  FDA and International compliance standards require CTDIvol as a measured output, available on all CT platforms. Therefore this value may be used to standardize exposure across scanners.  Three levels of CTDIvol are recommended for small, medium, and large subjects, based on three small, medium, and large Body Mass Index (BMI) sizes. This mechanism allows for optimal doses to the various size subjects.

Dose modulation is currently not recommended because of varying manufacturer specifications.

Radation Exposure Specification

| Parameter | Specification |
| --- | --- |
| Pitch | Target: 0.9-1.1 |
| Exposure Time | Target: 0.5sec |
| kV | Target: 120 |
| Dose Modulation | Target: OFF |
| Scan FOV (SFOV) | Target: Large |
| Scan Collimation (# of Detector Channels x Slice Width) | Target: ≥64 x 0.625mm |

Radation Exposure (CTDIvol): Matched across scanners within 3% of target

| Parameter | Specification |
| --- | --- |
| CTDIvol Small Range | Target: BMI <20 4.2mGy ±3% |
| CTDIvol Medium Range | Target: BMI 20-30 7.6mGy ±3% |
| CTDIvol Large Range | Target: BMI >30 11.4mGy ±3% |

Scan Coverage

The following is from the corresponding section of vCT:

| Parameter | Specification |
| --- | --- |
| Anatomic Coverage | The Technologist shall perform the scan such that the acquired anatomy is entire lung fields only, bilaterally (lung apices through bases). |

## 3. Image Data Reconstruction

Discussion

Spatial Resolution quantifies the ability to resolve spatial details. Lower spatial resolution can make it difficult to accurately determine airway morphometry, and as a consequence, decreases the precision of airway measurements.  Increased spatial resolution typically comes with an increase in noise. Therefore, the choice of factors that affect spatial resolution typically represent a balance between the need to accurately represent fine spatial details of objects and the noise within the image. Spatial resolution is mostly determined by the scanner geometry (which is not usually under user control) and the reconstruction kernel (which is somewhat under user control as the user usually gets to choose from a limited set of choices of reconstruction kernels provided at the scanner).  It is stated in terms of “the number of line-pairs per cm that can be resolved in a scan of resolution phantom (such as the synthetic model provided by the American College of Radiology and other professional organizations).” –OR– “the full width at half maximum of the line spread function”.

Noise Metrics quantify the magnitude of the random variation in reconstructed CT numbers.  Some properties of the noise can be characterized by the standard deviation of reconstructed CT numbers over a uniform region in phantom.  The standard deviation is limited since it can vary by changing the reconstruction kernel, which will also impact the spatial resolution.  A more comprehensive metric would be the noise-power spectrum which measures the noise correlation at different spatial frequencies.

Reconstruction Field of View affects reconstructed pixel size because the fixed image matrix size of most CT scanners is 512 X 512.  If it is necessary to expand the field of view to encompass more anatomy, the resulting larger pixels may be insufficient to achieve the claim. A targeted reconstruction with a smaller field of view may be necessary, but a reconstruction with that field of view would need to be performed for every time point. Pixel Size directly affects voxel size along the subject x-axis and y-axis. Smaller voxels are preferable to reduce partial volume effects and provide higher measurement precision.  Pixel size in each dimension is not the same as resolution in each dimension; inherent resolution is different than how the data is reconstructed and is strongly affected by the reconstruction kernel.  When one is comparing data fields of different resolution, one should not sacrifice higher resolution data to match the level of lower resolution data.

Reconstruction Interval (a.k.a. Slice spacing) For multidetector row CT (MDCT) scanners, creating overlapping image data sets has NO effect on radiation exposure; this is true because multiple reconstructions having different kernel, slice thickness and intervals can be reconstructed from the same acquisition (raw projection data) and therefore no additional radiation exposure is needed.  <Note that the slice thickness is “nominal” since the thickness is not technically the same at the middle and the edges>

Reconstruction kernel characteristics need to be defined to optimize the analysis while still meeting the requirements for noise and spatial resolution. A softer kernel can reduce noise at the expense of spatial resolution.  An edge enhancing sharp kernel can improve resolving power at the expense of increased noise. Therefore a standard kernel, between a soft and sharp kernel, is recommended.  For lung densitometry soft reconstruction filters have been recommended and for bronchial analysis a sharp reconstruction is being used.

The effects of iterative reconstructions across CT models on quantitative accuracy and reproducibility are currently not fully understood.

Beam hardening correction should be turned on.

Specification

For quantification of airways and lung parenchyma, the reconstruction software produces images that meet the following specifications:

| Reconstruction field of view (DFOV) | Target 1cm of tissue along each lateral border of the lungs |
| --- | --- |
| Matrix | Target 512X512 |
| Slice thickness | Target 0.6 -0.9mm |
| Slice Interval | Target 25% to 50% overlap |
| Reconstruction Kernel | Target Standard: Siemens B35, Philips B, GE STD, Toshiba FC01 |

## 4. Image Analysis

Discussion

Lung and airway segmentation method (protocol specific). Image processing shall use the same or equivalent lung and airway segmentation algorithm throughout each study.

With the advance of 3D volumetric CT the field has moved away from parenchymal analysis within 2D interleaved slices towards a full 3D segmentation of the lung volumes. Given the typically large number of slices in a multi-slice volume (100-500 slices), automated segmentation is a practical necessity. Accurate automated segmentation of the lung structures requires combining several common segmentation algorithms depending on the level of regional analysis required. The most basic and achievable is automated segmentation of the lung volume.  It is readily achievable to segment the right and left lung volumes themselves using thresholding followed by morphological operators and region growing as outlined in [Hu et al IEEE]. This can be subsequently followed by regional analysis to identify “core” of central regions of the lung volume distinct from more distal or peripheral regions of the lung volume, “peel” as described by Hoffman et al. Further refinement of automated segmentation methods have identified  the lung lobes by using feature space and k-means based classifiers to identify the voxels lying along the fissures between lobes and thus segment the lobes independently from the lung volume with further refinement using atlas-based training sets from prior lung data [cite Zhang et al.]

Validated segmentation algorithms of the lung and airways will be used.

Several commercial and prototype commercial packages are becoming available for these analyses [VIDA, MeVIS, Goldin…]. Volume correction should be applied as necessary using the most current acceptable method.

Histogram measurements will include the fraction of lung less than a “reference CT number”, usually -950 HU, 15% method and are to be the same throughout the study.

Once the volumes of interest have been identified, parenchymal voxel intensities can be plotted on their frequency histogram and thresholded using 3 empirical thresholds depending on the measure of interest and lung inflation volume: -950, -910, and -850 HU are typically used. The voxels which fall below the threshold are typically then normalized by the total number of voxels in the segmented volume (e.g. lung volume) to provide a percent measure of “diseased” volume in the parenchyma. For slices thinner than 5 mm, the -950 threshold at the inspiratory, or TLC, lung volume is widely accepted as the cutoff for emphysema, -910 for more moderate disease, and -850 for mild emphysema. In studies of asthma, the most common parenchymal analysis is at the expiratory, or FRC and sometimes RV, lung inflation with the -850, or -856, cutoff to identify regions of air trapping.

Airways

1. Seeded region growing followed by skeletonization.

2. minimum 5-7 generations of the airway tree at inspiratory (TLC) lung inflation.

3. Need to identify airway segments for longitudinal assessment because of dependence on demographics (i.e. height, gender) and airway generation (i.e. wall thickness and area measures vary with position along the tree).

4. Manual selection, or 2D analysis involves subject selection criteria and fail to take advantage of the 3D volumetric data set.

Longitudinal analysis:

Shall be on same scanner, same acquisition parameters, similar breath hold.  If different scanner, adjust for inter-scanner differences, adjust for differences in breath hold.

9.3. Required Characteristics of Resulting Data

**Introduce the section.

Lung Volume

Emphysema metrics

# IV. Compliance

## Acquisition Device

Compliance is certified according to specifications set out in the Image Acquisition section above.  Additionally, compliant Acquisition Devices shall provide means to record the information identified in the Subject Handling section as means to document compliance of the Image Acquisition Site to the specifications noted there.

## Reconstruction Software

Compliance to specifications as set out in the Image Reconstruction section above.  Additionally, compliant Reconstruction Software shall propagate the information collected at the prior Subject Handling and Imaging Acquisition stages and extend it with those items noted in the Reconstruction section.  See the compliance procedure notes associated with Acquisition Devices above for procedural assistance to identify Model Specific Parameters for Reconstruction Software.

## Software Analysis Tool

Compliance to specifications as set out in the Image Analysis section above.  Additionally, compliant Software Analysis Tools shall propagate the information collected at the prior Subject Handling, Imaging Acquisition, and Imaging Reconstruction stages and extend it with those items noted in the Analysis section

## Image Acquisition Site

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large patient population under consideration.  For imaging it is important to consider the availability of:

- appropriate imaging equipment and quality control processes,

- appropriate injector equipment and contrast media,

- experienced CT Technologists for the imaging procedure, and

- processes that assure imaging profile compliant image generation at the correct point in time.

A calibration and QA program shall be designed consistent with the goals of the clinical trial. This program shall include (a) elements to verify that sites are performing correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address issues relating to both radiation dose and image quality (which may include issues relating to water calibration, uniformity, noise, spatial resolution -in the axial plane-, reconstructed slice thickness z-axis resolution, contrast scale, CT number calibration and others). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.

# References

<fill in, mostly from those cited in the Background appendix>

# Appendices

## Acknowledgements and Attributions

This document is proffered by the Radiological Society of North America (RSNA) Quantitative Imaging Biomarker Alliance (QIBA) CT Lung Densitometry Technical Committee, d-CT. This technical committee is composed of scientists representing the imaging device manufacturers, image analysis software developers, image analysis laboratories, biopharmaceutical industry, academia, government research organizations, professional societies, and regulatory agencies, among others. All work is classified as pre-competitive. A more detailed description of the d-CT group and its work can be found at the following web link: http://qibawiki.rsna.org/index.php?title=Densitometry_CT.

The Densitometry CT Technical Committee (in alphabetical order):

- …

The Densitometry CT Technical Committee is deeply grateful for the support and technical assistance provided by the staff of the Radiological Society of North America.

## Breathing Instructions

It is important for the Technologist to ensure that the subject is able to follow the commands appropriately and also that the subject is appropriately coached PRIOR to performing the actual imaging procedures. Practice breathing maneuvers are recommended and built into the breathing instructions to reduce subject anxiety and limit mistakes during actual scan acquisition.

(a) Practice Breathing (TLC)

For the first part of this scan, I am going to ask you to take a couple of deep breaths in and out before we have you hold your breath all the way in.

First let’s practice

Take a deep breath in (watch chest to ensure a deep breath as far in as possible then start scan)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Now breathe all the way IN…IN…IN… (watch chest to ensure a deep breath in as far as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure spine remains on the table, patients is not shaking- watch for these throughout the study!)

Breathe and relax

(b) Scout views: PA & Lateral (TLC)

OK let’s get started

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Now breathe all the way IN... IN... IN... and hold it (watch chest to ensure a deep breath as far in as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure the spine remains on the table, subject is not shaking and then start the scout scan)

Perform scout

At the end of the scout- Breathe and relax

Scout views: Lateral (TLC)

OK let’s get started

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Now breathe all the way IN... IN... IN... and hold it (watch chest to ensure a deep breath as far in as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure the spine remains on the table, subject is not shaking and then start the scout scan)

Perform scout

At the end of the scout- Breathe and relax

(c) Inspiratory CT (TLC)

Now we’re ready again so please

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Now breathe all of the way IN...IN...IN as far as possible and hold it in (watch chest to ensure a deep breath in as far as possible)

Keep holding your breath – DO NOT BREATHE!

At the end of scan- Breathe and relax

(d) Practice Breathing (RV)

For the second part of this scan, I am going to ask you to take a couple of deep breaths in and out before we have you hold your breath all the way out.

First let’s practice

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take another deep breath in (watch chest to ensure a deep breath in)

Now breathe all the way OUT…OUT…OUT…as far as possible and hold it out (watch chest to ensure a deep breath out as far as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure spine remains on the table, patients is not shaking- watch for these throughout the study!)

Breathe and relax

(e) Scout views: PA & Lateral (RV)

Scout views: PA (RV)

OK let’s get started

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take another deep breath in (watch chest to ensure a deep breath in)

Now breathe all the way OUT…OUT…OUT…as far as possible and hold it out (watch chest to ensure a deep breath out as far as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure spine remains on the table, patients is not shaking- watch for these throughout the study!)

Breathe and relax

Scout views: Lateral (RV)

OK let’s get started

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take another deep breath in (watch chest to ensure a deep breath in)

Now breathe all the way OUT…OUT…OUT…as far as possible and hold it out (watch chest to ensure a deep breath out as far as possible)

Keep holding your breath- DO NOT BREATHE!  (watch chest to ensure spine remains on the table, patients is not shaking- watch for these throughout the study!)

Breathe and relax

(f) Expiratory CT (RV)

Now we’re ready again so please

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take a deep breath in (watch chest to ensure a deep breath in)

Let it out (watch chest to ensure air is out)

Take another deep breath in (watch chest to ensure a deep breath in)

Now breathe all of the way OUT...OUT...OUT… as far as possible and hold it out (watch chest to ensure a deep breath out as far as possible)

Keep holding your breath – DO NOT BREATHE!

At the end of scan- Breathe and relax

## Background Information

## Conventions and Definitions

Acquisition vs. Analysis vs. Interpretation: This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be computational steps that transform the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge. (The transformation of knowledge into wisdom is beyond the scope of this document.)

Other Definitions:

Image Analysis, Image Review, and/or Read: Procedures and processes that culminate in the generation of imaging outcome measures, such tumor response criteria. Reviews can be performed for eligibility, safety or efficacy. The review paradigm may be context specific and dependent on the specific aims of a trial, the imaging technologies in play, and the stage of drug development, among other parameters.

QIBA

The Quantitative Imaging Biomarker Alliance (QIBA) is an initiative to promote the use of standards to reduce variability and improve performance of quantitative imaging in medicine. QIBA provides a forum for volunteer committees of care providers, medical physicists, imaging innovators in the device and software industry, pharmaceutical companies, and other stakeholders in several clinical and operational domains to reach consensus on standards-based solutions to critical quantification issues. QIBA publishes the specifications they produce (called QIBA profiles), first to gather public comment and then for field test by vendors and users.

QIBA envisions providing a process for developers to test their implementations of QIBA profiles through a compliance mechanism. After a committee determines that a profile has undergone sufficient successful testing and deployment in real-world care settings, it is released for use.  Purchasers can specify conformance with appropriate QIBA profiles as a requirement in requests for proposal. Vendors who have successfully implemented QIBA profiles in their products can publish conformance statements (called QIBA Compliance Statements) represented as an appendix called “Model-specific Parameters.”  General information about QIBA, including its governance structure, sponsorship, member organizations and work process, is available at http://qibawiki.rsna.org/index.php?title=Main_Page.

CT Densitometry for Lung Disease

Inflammatory parenchymal lung diseases are common and are significant causes of disability and premature death. These diseases are the result of sub-acute/chronic or chronic inflammatory processes, and are linked to cigarette smoking, either as a cause or as a modifying agent. Chronic obstructive pulmonary disease (COPD) is currently the 12th leading cause of disability in the world and is predicted to be 5th by the year 2020 (201). In the United States alone, it has been estimated that the annual cost of morbidity and early mortality due to COPD is approximately 4.7 billion dollars (202). COPD is a complex condition in which environmental factors interact with genetic susceptibility to cause disease. Tobacco smoke is the most important environmental risk factor, and in susceptible individuals it causes an exaggerated inflammatory response that ultimately destroys the lung parenchyma (emphysema) and/or increases airway resistance by remodeling of the airway wall (203). It has long been known that the pathway varies between individuals; some patients have predominant emphysema while others can have similar degrees of airflow obstruction due to severe small airway disease with relatively preserved parenchyma, but the proportion and contribution of each to the pathogenesis of disease is still unknown. (175)

The pathologic events leading to emphysema are insidious and include structural and physiologic alterations that are characterized by inflammatory processes within the peripheral pulmonary parenchyma, thickening of arteriolar walls, and parenchymal destruction. A growing body of literature documents that these changes are likely to be associated with alterations in blood flow dynamics at a regional, microvascular level, and thus may serve as a beacon pointing toward the onset of early emphysema. Regional alterations in blood flow parameters may not only serve as an early marker for inflammatory processes but may also be a major etiologic component of the pathologic process, leading to emphysema in a subset of the smoking population (not all smokers have emphysema). (176)

Measures based on airflow or other measures of global lung function have reached their limits in their ability to provide new insights into the etiology of the disease, or even in leading us to an understanding of how lung volume reduction, in late stages of the disease, provides patient improvements. A number of articles have been written in which attempts are made to explain improvements of physiologic status post-LVRS (9, 16) on the basis of lung mechanics, and we find it difficult to understand how these relate to the observations from the NETT (15) showing that subjects with apical but not basal prevalence of disease receive the greatest benefit from surgery. However, if regional pulmonary perfusion is again brought into consideration, it makes sense that, if one removes apical lung that is not contributing well to gas exchange and blood is shunted to less diseased basal lung, gas exchange will be improved. Furthermore, by removing a diseased portion of the basal lung when the disease is predominantly basal, then it is likely that blood will be preferentially shunted to the contralateral basal lung. Using scintigraphy to assess regional V˙ /Q˙ , Moonen and colleagues (2) have recently concluded that an important mechanism for improvement in functional status post-LVRS relates to the reduction of regional shunt (i.e., blood flow may be directed toward regions of improved ventilation whereas regions receiving blood flow but that have poor ventilation are removed).

A recent international consensus statement on the diagnosis and therapy of COPD—the Global Strategy for the Diagnosis, Management, and Prevention of Chronic Obstructive Lung Disease (GOLD [Global Initiative for Chronic Obstructive Lung Disease])—has established diagnostic criteria that currently do not include CT findings (17). This is not surprising given that the consensus statement has been developed in part for the World Health Organization. It is notable that the summary makes the observation that different inflammatory events occur “in various parts of the lung,” a reference to the marked heterogeneity of the disease which cannot be defined without imaging. Of interest also are the future recommended research directions, which include identifying better defining characteristics of COPD, developing other measures to assess and monitor COPD, and recognizing the increasing need to identify earlier cases of the disease, all potential outcomes of improvements to quantitative lung imaging.

It has been well demonstrated that lung function declines with age (18–21) and perhaps also as a result of inflammation (22). This has confounded research related to the effects of smoking cessation on lung health. There are mixed results as to whether or not smoking cessation halts the progression of emphysematous lung disease (20, 23–27). Work by Bosse and colleagues (28) attempted to take into account the aging process and suggested that the disease process is slowed if one stops smoking. However, tests were not sensitive enough to conclude this definitively. There appear to be important sex differences in the effects of cigarette smoking and cessation (29). No reliable specific biochemical markers of disease presence or progression have been identified (30), in part perhaps because of the lack of a sensitive standard to diagnose and follow the diseases. More recently, CT parameters have been shown to be likely more sensitive to disease progression (31). Furthermore, the long time course of these diseases means that clinical trials using only whole lung function as primary outcome measures require huge numbers of subjects for extremely long periods of time.

Anatomic–Physiologic Correlates of Emphysema

The lack of a direct marker for emphysema has meant that epidemiologic studies have been limited to COPD, and these perhaps give a limited view as to the epidemiology of emphysema, a specific subset of COPD but not identified as such by spirometry. The direct effect of cigarette smoking on lung function has been widely studied, with differences in relative changes in FEV1 and the effects of smoking noted (28, 32–35). Although some studies show an increased rate of loss of FEV1 for current smokers, there is a less significant decrease in FEV1 for reformed Hoffman, Simon, and McLennan: CT-Based Lung Structure and Function 521 smokers (28, 32).However, more recent studies have found similar FEV1 declines with age in both smokers and never-smokers (34, 35). It must be emphasized that these changes are likely related to bronchial hyper-reactivity (36, 37) rather than to emphysema, highlighting again the need for objective measurement tools to assess emphysema. As indicated, chronic airflow limitation (COPD) is commonly seen in emphysema, but it is not essential. Measurements of lung physiology are not always able to distinguish the abnormalities that result from emphysema from those which result from the other causes of COPD, such as chronic bronchitis or asthma (38). The strongest positive association between an index of airflow limitation, FEV1 (% predicted) and a pathologically derived emphysema score comes from the National Institutes of Health Intermittent Positive-pressure Breathing Trial (39). There were only 48 subjects in this study, as autopsies were required for the pathologic assessment to be performed. Pulmonary function tests were performed every 3 months during the study, and were therefore available at some point before death. However, these subjects were highly selected; to enter the study, they were required to have very significant airflow obstruction, and could not be severely hypoxic; and to complete the study, they had to die during the observation period. In contrast, a study examining pathologic lung specimens taken during surgery, and appropriately fixed, showed no relationship between the pathologic emphysema rating and indices of airflow (40). Furthermore, an autopsy study enrolling 242 subjects over 6 years demonstrated that, although those subjects with greater pulmonary disability tended to have a greater degree of pathologic emphysema, 17 subjects with greater than 30% pathologic emphysema had no evidence for clinical COPD (41). Other pulmonary function tests—namely, diffusing capacity for carbon monoxide (DlCO) and the exponential description of the deflation pressure/volume curve (K)—have been used to identify, and to obtain a measure of, severity for pulmonary emphysema. A number of studies have found that measurement of DlCO has a very weak correlation with the pathologic assessment of emphysema (42–44). Measurements of elastic recoil pressure curves in life compared with pathologic assessment of emphysema at subsequent lung resection or postmortem have yielded conflicting results on the value of static compliance and K as a measure of emphysema (45–49). More recent studies show a weak but significant correlation between K and macroscopic emphysema (r _ 0.49) (47, 48), with K believed to be a measure of alveolar distensibility. This background highlights the continuing search for a marker for emphysema presence and severity.

Even though recent research has advanced our understanding of COPD pathogenesis, leading to the identification of potential targets and pathways for drug development, there are still major difficulties in conducting clinical trials designed to evaluate the benefits of new drug treatment for several reasons. These reasons include:

(1) the lack of validated short- and intermediate-term endpoints (or surrogates) that are predictive of future hard clinical outcomes, and

(2) the lack of a method to provide precise phenotypes suitable for large-scale studies.

It is for these reasons that computed tomography (CT) has become such an important tool in COPD research. CT provides a noninvasive method to obtain images of the lung that look similar to anatomic assessment, and CT images themselves are densitometry maps of the lung. Therefore any change in the structure of the lung will change the densitometry of the lung and, therefore, the image. Virtually every clinical center in all regions of the world has access to a CT scanner, so it is thought that CT images should be quite easy to obtain and it should be easy to conduct large, meaningful clinical studies.

# Comments

This is open for debate, of course. The CT lung volume is very difficult to define because of the hilar structures., and one has to stop excluding structures at some distinct point in any case Correlations with TLC will always be limited by their differences in definition and differences in patient positioning.

See Proc Am Thorac Soc Vol 5. pp 929–933, 2008: Stable scanners may vary only +/- 2HU in foam over a period of 2 years. Others can slowly drift 4 HU over this period; just before an X-ray tube breaks down, a sudden shift of 8HU can be observed. Etc.

Compare with DCEMRI claim

Insert data from Iowa, COPDGene, SARP, UCLA, ? Europe (Stoel/Wigstrom)

Corrected for lung volumes

This is what we have seen a couple of times. At baseline the CT scanner was stable and accepted, but it was not calibrated correctly. Once the scanner is recalibrated in a later stage, one can never go back to the initial state.

995-1000

We scan our patients caudocraneally so that artifacts are reduced if a patient cannot hold their breath long enough. (Although with current acquisition speeds, this is less of an issue)

In our trials we record the table heights for each patient and keep them constant over time.

Or: one could reconstruct left and right lung separately with high resolution, where the pixel size is  completely standardized over all patients.

Need to resolve attributions