QIBA Profile.  Computed Tomography: Change Measurements in the Volumes of Solid Tumors

Version 2.1

DD Month 2012

Table of Contents

Open Issues:	2

Closed Issues:	2

I. Executive Summary	5

II. Clinical Context and Claims	5

Utilities and Endpoints for Clinical Trials	5

Claim:  Measure Change in Tumor Volume	6

III. Profile Details	7

1. Subject Handling	9

2. Image Data Acquisition	12

3. Image Data Reconstruction	14

4. Image Analysis	16

IV. Compliance	18

1. Performance Assessment: Tumor Volume Change Variability	19

1. Performance Assessment:	19

Image Acquisition Site	20

References	20

Appendices	23

A: Acknowledgements and Attributions	23

B: Background Information	24

C: Conventions and Definitions	28

D: Model-specific Instructions and Parameters	29

# Open Issues:

The following issues have not been resolved to the satisfaction of the technical committee.  An open issue may be a short question prompting a proposed resolution or discussion. The issues and answers below may represent some of the directions the Committee is currently leaning. Feedback on these issues is encouraged, particularly during the Public Comment period for the profile.

| 1 | Q. Is the claim appropriate/supported by the profile details, published literature, and QIBA groundwork? Is it stated in clear and statistically appropriate terms? A. |
| --- | --- |
| 2 | Q. What kind of additional study (if any is needed) would best prove the profile claim? A. |
| 3 | Q. How do we balance specifying what to accomplish vs how to accomplish it? A. E.g. if the requirement is that the scan be performed the same way, do we need to specify that the system or the Technologist record how each scan is performed? If we don’t, how will the requirement to “do it the same” be met? |
| 4 | Q. Should there be a “patient appropriateness” or “subject selection” section? A. The protocol template includes such a section to describe characteristics of appropriate (and/or inappropriate) subjects. E.g. a requirement that the patient be able to hold their breath for 15 seconds. We could also discuss what constitutes an “assessable lesion” (the claim introduces this term) |
| 8 | Q. Are there sufficient DICOM fields for all of what we need to record in the image header, and what are they specifically? A. For those that exist, we need to name them explicitly. For those that may not currently exist, we need to work with the appropriate committees to have them added. |
| 9 | Q. Have we worked out the details for how we establish compliance to these specifications? A. We are continuing to work on how this is to be accomplished but felt that it was helpful to start the review process for the specifications in parallel with working on the compliance process. |

# Closed Issues:

The following issues have been considered closed by the technical committee.  They are provided here to forestall discussion of issues that have already been raised and resolved, and to provide a record of the rationale behind the resolution.

| 5 | Q. Does 4cm/sec “scan speed” preclude too many sites? A. No. Most 16-slice (and greater) scanners would be able to achieve this (although due to an idiosyncracy of the available scan modes, the total collimation needs to be dropped to 16mm rather than 20mm) Some examples that would meet this include: (a) 16 x 1mm collimation with 0.5 second rotation time and pitch ³ 1.25 OR (b) 16 x 1mm collimation with 0.4 second rotation time and pitch ³ 1 OR (c) 16 x 1.25 mm collimation with 0.5 second rotation time and pitch ³ 1 OR (d) 16 x 1.5mm collimation with 0.5 second rotation time and pitch ³ .833 Keep in mind that 16 x 0.75 mm collimation would require (i) pitch > 1.67 at 0.5 second rotation time (which breaks the Pitch< 1.5 requirement OR (ii) pitch > 1.33 at 0.4 second rotation time (which is fine) A 4cm/sec threshold is needed since it would likely alleviate potential breath hold issues. Because the reconstructed image thickness allowed here was > 2 mm, all of the above collimation settings would be able to meet both the breath hold requirements as well as the reconstructed image thickness requirements. |
| --- | --- |
| 6 | Q. What do we mean by noise and how do we measure it? A. Noise means standard deviation of a region of interest as measured in a homogeneous water phantom. FDA has starting looking at Noise Power Spectrum in light of recent developments in iterative reconstruction and an interest in evaluating what that does to the image quality/characteristics. QIBA should follow what comes out of those discussions, but since FDA is not mandating it and since few systems or sites toda are in a position to measure or make effective use of it, this profile will not mandate it either. It has promise though and would be worth considering for future profile work. |
| 7 | Q. Is 5HU StdDev a reasonable noise value for all organs? A. No. Will change to 18HU. Not sure where the 5 HU standard deviation came from. The 1C project used a standard deviation of 18HU. At UCLA, our Siemens Sensation 64 will yield a standard deviation of 17 HU for: a. 120kVp, 50 eff. mAs, 1 mm thickness, B30F filter To get this down to 5 HU would require: a. Increasing the eff. mAs to 550, OR b. Increasing the slice thickness to 2 mm AND increasing eff. mAs to 275 |
| 10 | Q. What is the basis of the specification of 15% for the variability in lesion volume assessment within the Image Analysis section, and is it inclusive or exclusive of reader performance? A. For the basis, see the paragraph below the table in Section B.2. It includes reader performance. Allocation of variability across the pipeline (shown in Figure 1) is fraught with difficulty and accounting for reader performance is difficult in the presence of different levels of training and competence among readers. Input on these points to help with this is appreciated (as is also the case for all aspects of this Profile). |
| 11 | Q. Should we specify all three levels (Acceptable, Target, Ideal) for each parameter? A. No. As much as possible, provide just the Acceptable value. The Acceptable values should be selected such that the profile claim will be satisfied. |
| 12 | Q. What is the basis for our claim, and is it only aspirational? A. Our claim is informed by an extensive literature review of results achieved under a variety of conditions. From this perspective it may be said to be well founded; however, we acknowledge that the various studies have all used differing approaches and conditions that may be closer or farther from the specification outlined in this document. In fact the purpose of this document is to fill this community need. Until field tested, the claim may be said to be “consensus.” Commentary to this effect has been added in the Claims section, and the Background Information appendix has been augmented with the table summarizing our literature sources. |
| 13 | Q. What about dose? A. A discussion has been added in Section 2 to address dose issues. |

# I. Executive Summary

X-ray computed tomography provides an effective imaging technique for assessing treatment response in subjects with cancer. Size quantification is helpful to evaluate tumor changes over the course of illness. Currently most size measurements are uni-dimensional estimates of longest diameters (LDs) on axial slices, as specified by RECIST (Response Evaluation Criteria In Solid Tumors).  Since its introduction, limitations of RECIST have been reported. Many investigators have suggested that quantifying whole tumor volumes could solve some of the limitations of diameter measures, and may have a major impact on subject management [1-2]. An increasing number of studies have shown that volumetry has value [3-12].

This QIBA Profile makes claims about the precision with which changes in tumor volumes can be measured under a set of defined image acquisition, processing, and analysis conditions, and provides specifications that may be adopted by users and equipment developers to meet targeted levels of clinical performance in identified settings.

The claims are based on several studies of varying scope now underway to provide comparison between the effectiveness of volumetry and uni-dimensional longest diameters as the basis for RECIST in multi-site, multi-scanner-vendor settings.

The intended audiences include:

- Technical staff of software developers and device manufacturers who create products for this purpose

- Biopharmaceutical companies, oncologists, and clinical trial scientists designing trials with imaging endpoints

- Clinical trialists

- Radiologists, technologists, and administrators at healthcare institutions considering specifications for procuring new CT equipment

- Radiologists, technologists, and physicists designing CT acquisition protocols

- Radiologists and other physicians making quantitative measurements on CT images

- Regulators, oncologists, and others making decisions based on quantitative image measurements

Note that specifications stated as “requirements” here are only requirements to achieve the claim, not “requirements on standard of care.”   Specifically, meeting the goals of the profile are secondary to properly caring for the patient.

# II. Clinical Context and Claims

## Utilities and Endpoints for Clinical Trials

These specifications are appropriate for quantifying the volumes of malignant tumors and measuring tumor longitudinal changes within subjects. The primary objective is to evaluate their growth or regression with serially acquired CT scans and image processing techniques.

Compliance with this profile by relevant staff and equipment supports the following claim(s):

## Claim:  Measure Change in Tumor Volume

Measured increases or decreases of more than 30% in the volume of a tumor are likely to be associated with a true change in the volume of the tumor.

or

For a given tumor, the probability of a true volume increase (or decrease), given a measured volume increase (or decrease) of more than 30%, is greater than 95%.

i.e. P(true volume change > 0% | measured volume change >30%) > 95%.

or

Measured increases or decreases of more than 30% in the volume of a tumor are likely to be associated with a true (physical? biological change) (therapeutic response?).

or

Measured increases or decreases of more than 30% in the volumes of tumors are not likely to be artifacts of the acquisition or measurement processes, and therefore are likely to be biomarkers of tumor progression or therapeutic response.

or

Measured increases or decreases of more than 30% in tumor volume are more likely to indicate a true change in the status of the tumor (progression or therapeutic response) than to be artifacts or random variation associated with the acquisition or measurement processes.

or

A measured increase of greater than 30% in the volume of a tumor indicates a greater than __% probability that there has been a true increase in tumor volume, and a measured decrease of greater than 30% in the volume of a tumor indicates a greater than __% probability that there has been a true decrease in tumor volume.

or

In formula terms P(TVC>=v | MVC>=t) >= c, where

TVC = a true volume change

v     = a clinically significant/biologically meaningful volume change

MVC = a measured volume change

t      = target threshold

c      = acceptable confidence

1. Currently for us v=0%, t=30%, c=95%?

2. We could do v and t in terms of absolute change, percentage change or proportional change.

3.

Note this is a binary detection paradigm rather than characterizing the measurement tool.

We could also choose to consider a curve instead of the single point above.

<<This is sort of like an acceptance/conformance criteria for a worst-case scenario within the constraints we’ve defined.  Going to a curve and relating to the actual size and bias, would add a lot of complexity.  This is a helpful starting point and the simplicity is appealing at this point in time.  As we mature we can look at more characterization/dimensions>>

This claim holds when the given tumor is measurable (i.e., tumor margins are sufficiently conspicuous and geometrically simple enough to be recognized on all images in both scans), and the longest in-plane diameter of the tumor is 10 mm or greater.  The volume changes refer to proportional changes, where the percentage change is the difference in the two volume measurements divided by the average of the two measurements.  Using the average instead of one of the measurements as the denominator, asymmetries in percentage change values are avoided.

For details on the derivation and implications of the Claim, refer to Appendix B.

While the claim has been informed by an extensive review of the literature, it is currently a consensus claim that has not yet been fully substantiated by studies that strictly conform to the specifications given here.  To date there has not existed a standard utilized by a sufficient number of studies.  The expectation is that during field test, data on the actual field performance will be collected and changes made to the claim or the details accordingly.  At that point, this caveat may be removed or re-stated.

# III. Profile Details

The sequencing of the Activities specified in this Profile are shown in Figure 1:

Figure 1: <Profile Name> - Activity Sequence

The method for measuring change in tumor volume may be described as a pipeline.  Subjects are prepared for scanning, raw image data is acquired, images are reconstructed and possibly post-processed.  Such images are obtained at two (or more) time points.  Image analysis assesses the degree of change between two time points for each evaluable target lesion by calculating absolute volume at each time point and subtracting.  Volume change is expressed as a percentage (delta volume between the two time points divided by the average of the volume at time point 1 and time point t).

The change may be interpreted according to a variety of different response criteria.  These response criteria are beyond the scope of this document.  Detection and classification of lesions as target is also beyond the scope of this document.

The Profile does not intend to discourage innovation.  The above pipeline provides a reference model.  Algorithms which achieve the same result as the reference model but use different methods are permitted, for example by directly measuring the change between two image sets rather than measuring the absolute volumes separately.

Equipment, software, staff or sites may claim conformance to this Profile as one or more of the “Actors” in the following table.  Compliant Actors shall meet all requirements described in the corresponding Activities shown in the table.  Failing to comply with a “shall” is a protocol deviation.  Although deviations invalidate the Profile Claim, such deviations may be reasonable and unavoidable as discussed below.

Table 1: Actors and Required Activities

| Actor | Activity | Section |
| --- | --- | --- |
| Acquisition Device | Subject Handling | 1. |
|  | Image Data Acquisition | 2. |
| Technologist | Subject Handling | 1. |
|  | Image Data Acquisition | 2. |
|  | Image Data Reconstruction | 3. |
| Radiologist | Subject Handling | 1. |
| Reconstruction Software | Image Data Reconstruction | 3. |
| Image Analysis Tool | Image Analysis | 4. |

The requirements included herein are intended to establish a baseline level of capabilities. Providing higher performance or advanced capabilities is both allowed and encouraged.  The profile does not intend to limit how equipment suppliers meet these requirements.

This Profile is “lesion-oriented”.  The profile requires that images of a given tumor be acquired and processed the same way each time.  It does not require that images of tumor A be acquired and processed the same way as images of tumor B; for example, tumors in different anatomic regions may be imaged or processed differently, or some tumors might be examined at one contrast phase and other tumors at another phase.

The requirements in this Profile do not codify a Standard of Care; they only provide guidance intended to achieve the stated Claim.  Although deviating from the specifications in this Profile may invalidate the Profile Claims, the radiologist or supervising physician is expected to do so when required by the best interest of the patient or research subject.  How study sponsors and others decide to handle deviations for their own purposes is entirely up to them.

Since much of this Profile emphasizes performing subsequent scans consistent with the baseline scan of the subject, the parameter values chosen for the baseline scan are particularly significant and should be considered carefully.

In some scenarios, the “baseline” might be defined as a reference point that is not necessarily the first scan of the patient.

## 1. Subject Handling

## This profile will refer primarily to “subjects”, keeping in mind that the recommendations apply to patients in general, and subjects are often patients too.

1.1 Timing Relative to Index Intervention Activity

When the Profile is being used in the context of a clinical trial, refer to relevant clinical trial protocol for further guidance or requirements on timing relative to index intervention activity.

1.2 Timing Relative to Confounding Activities

This document does not presume any other timing relative to other activities.

Fasting prior to a contemporaneous FDG PET scan or the administration of oral contrast for abdominal CT are not expected to have any adverse impact on this profile.

1.3 Contrast Preparation and Administration

1.3.1 Discussion

Contrast characteristics influence the appearance, conspicuity, and quantification of tumor volumes. 
Non-contrast CT may not permit an accurate characterization of the malignant visceral/nodal/soft-tissue lesions and assessment of tumor boundaries. Therefore, consistent use of intravenous contrast is required to meet the claims of this Profile.

However, the use of contrast material (intravenous or oral) may be not be medically indicated in defined clinical settings or may be contra-indicated for some subjects.  Radiologists and supervising physicians may omit intravenous contrast or vary administration parameters when required by the best interest of patients or research subjects, in which case lesions may still be measured but the measurements will not be subject to the Profile claims.

The following specifications are minimum requirements to meet Profile claims. Ideally, intravenous contrast type, volume, injection rate, use or lack of a "saline chase," and time between contrast administration and image acquisition should be identical for all time points, and the use of oral contrast material should be consistent for all abdominal imaging timepoints.

The Acquisition Device shall record the use and type of contrast, actual dose administered, injection rate, and delay in the image header. <Make advisable not required>  This may be by automatic interface with contrast administration devices in combination with text entry fields filled in by the Technologist.  Alternatively, the technologist may enter this information manually on a form that is scanned and included with the image data as a DICOM Secondary Capture image.

1.3.2 Specification

| Parameter | Specification |
| --- | --- |
| Use of intravenous or oral contrast | The Radiologist shall determine if the contrast protocol is appropriate for the subject. The Technologist shall use intravenous contrast parameters consistent with baseline. Specifically, the total amount of contrast administered (grams of iodine) shall not vary by more than 25% between scans; contrast injection rate shall be at least 2ml/sec and shall not vary by more than 2ml/sec for arterial phase imaging, and images to be compared are to be obtained at the same phase (i.e. arterial, venous, or delayed). |
| Image Header |  |

1.4 Subject Positioning

1.4.1 Discussion

Consistent positioning avoids unnecessary changes in attenuation, changes in gravity induced shape and fluid distribution, or changes in anatomical shape due to posture, contortion, etc.  Significant details of subject positioning include the position of their arms, the anterior-to-posterior curvature of their spines as determined by pillows under their backs or knees, the lateral straightness of their spines. Prone positioning is not recommended.  Positioning the subject Supine/Arms Up/Feet First has the advantage of promoting consistency, and reducing cases where intravenous lines go through the gantry, which could introduce artifacts.

When the patient is supine, the use of positioning wedges under the knees and head is recommended so that the lumbar lordosis is straightened and the scapulae are both in contact with the table. However, the exact size, shape, etc. of the pillows is not expected to significantly impact the profile claim.  It is expected that clinical trial documentation or local clinical practice will specify their preferred patient positioning.

Recording the Subject Positioning and Table Heights in the image header is helpful for auditing and repeating baseline characteristics.

Consistent centering of the patient avoids unnecessary variation in the behavior of dose modulation algorithms during scan.

1.4.2 Specification

| Parameter | Specification |
| --- | --- |
| Subject Positioning | The Technologist shall position the subject consistent with baseline. If baseline positioning is unknown, position the subject Supine if possible, with devices such as positioning wedges placed as described above. |
| Table Height & Centering | The Technologist shall adjust the table height to place the mid-axillary plane to pass through the isocenter. The Technologist shall position the patient such that the “sagittal laser line” lies along the sternum (e.g. from the suprasternal notch to the xiphoid process). |

1.5 Instructions to Subject During Acquisition

1.5.1 Discussion

Breath holding reduces motion that might degrade the image. Full inspiration inflates the lungs, which separates structures and makes tumors more conspicuous.

Since some motion may occur due to diaphragmatic relaxation in the first few seconds following full inspiration, a proper breath hold will include instructions like "Lie still, breathe in fully, hold your breath, and relax”, allowing 5 seconds after achieving full inspiration before initiating the acquisition.

Although performing the acquisition in several segments (each of which has an appropriate breath hold state) is possible, performing the acquisition in a single breath hold is likely to be more easily repeatable and does not depend on the Technologist knowing where the tumors are located.

1.5.2 Specification

| Parameter | Specification |
| --- | --- |
| Breath hold | The Technologist shall instruct the subject in proper breath-hold and start image acquisition shortly after full inspiration, taking into account the lag time between full inspiration and diaphragmatic relaxation. The Technologist shall ensure that for each tumor the breath hold state is consistent with baseline. |
| Image Header | The Technologist shall record factors that adversely influence subject positioning or limit their ability to cooperate (e.g., breath hold, remaining motionless, agitation in subjects with decreased levels of consciousness, subjects with chronic pain syndromes, etc.). The Acquisition Device shall provide corresponding data entry fields. |

1.6 Timing/Triggers

1.6.1 Discussion

The amount and distribution of contrast at the time of acquisition can affect the appearance and conspicuity of tumors.

1.6.2 Specification

| Parameter | Specification |
| --- | --- |
| Timing / Triggers | The Technologist shall ensure that the time-interval between the administration of intravenous contrast (or the detection of bolus arrival) and the start of the image acquisition is consistent with baseline. |
| Image Header | The Acquisition Device shall record actual Timing and Triggers in the image header. |

## 2. Image Data Acquisition

2.1 Discussion

CT scans for tumor volumetric analysis can be performed on any equipment that complies with the specifications set out in this profile.  However, we strongly encourage performing all CT scans for an individual subject on the same platform (manufacturer, model and version) which we expect will further reduce variation.

Many scan parameters can have direct or indirect effects on identifying, segmenting and measuring lesions.  To reduce this potential source of variance, all efforts should be made to have as many of the scan parameters as possible consistent with the baseline.

<Scan Protocols don’t map directly between scanner makes (or models).  You’ll have to “interpret” to get something “consistent”. – Mike HW  Also address manual and auto-exposure techniques.>

The purpose of the minimum scan speed requirement is to permit acquisition of an anatomic region in a single breath-hold, thereby preventing respiratory motion artifacts or anatomic gaps between breath-holds. This requirement is applicable to scanning of the chest and upper abdomen, the regions subject to these artifacts, and is not required for imaging of the head, neck, pelvis, spine, or extremities.

Coverage of additional required anatomic regions (e.g. to monitor for metastases in areas of likely disease) depends on the requirements of the clinical trial or local clinical practice.  In baseline scans, the tumor locations are unknown and may result in a tumor not being fully within a single breath-hold, making it “unmeasurable” in the sense of this Profile.

Pitch is chosen so as to allow completion of the scan in a single breath hold.

For subjects needing two or more breath-holds to fully cover an anatomic region, different tumors may be acquired on different breath-holds.  It is still necessary that each tumor be fully included in images acquired within a single breath-hold to avoid discontinuities or gaps that would affect the measurement.

Scan Plane (transaxial is preferred) may differ between subjects due to the need to position for physical deformities or external hardware.  For an individual subject, a consistent scan plane will reduce unnecessary differences in the appearance of the tumor.

Total Collimation Width (defined as the total nominal beam width, NxT, for example 64x1.25mm) is often not directly visible in the scanner interface.  Manufacturer reference materials should explain how to determine this for a particular scanner make, model and operating mode.  Wider collimation widths can increase coverage and shorten acquisition, but can introduce cone beam artifacts which may degrade image quality.  Imaging protocols will seek to strike a balance to preserve image quality while providing sufficient coverage to keep acquisition times short.

Nominal Tomographic Section Thickness (T), the term preferred by the IEC, is sometimes also called the Single Collimation Width.  It affects the spatial resolution along the subject z-axis.

<and point to recon for these other details … Smaller voxels are preferable to reduce partial volume effects and provide higher accuracy due to higher spatial resolution. Kevin Fix>

X-ray CT uses ionizing radiation.  Exposure to radiation can pose risks; however as the radiation dose is reduced, image quality can be degraded.  It is expected that health care professionals will balance the need for good image quality with the risks of radiation exposure on a case-by-case basis.  It is not within the scope of this document to describe how these trade-offs should be resolved.

Anatomic Coverage recording by the Acquisition Device may or may not require the attention of the Technologist.

2.2 Specification

The Acquisition Device shall be capable of performing scans with all the parameters set as described in the following table.  The Technologist shall set up the scan to achieve the requirements in the following table.

| Parameter | Specification |
| --- | --- |
| Scan Duration for Thorax | Achieve a table speed of at least 4cm per second, if table motion is necessary to cover the required anatomy. |
| Anatomic Coverage | Tumors to be measured and additional required anatomic regions shall be fully covered. If multiple breath-holds are required, the technologist shall obtain image sets with sufficient overlap to avoid gaps within the required anatomic region(s), and shall ensure that each tumor lies wholly within a single breath-hold. |
| Scan Plane (Image Orientation) | Consistent with baseline. |
| Total Collimation Width | Greater than or equal to 16mm. |
| IEC Pitch | Less than 1.5. |
| Tube Potential (kVp) | Consistent with baseline (i.e. the same kVp setting if available, otherwise as similar as possible). |
| Nominal Tomographic Section Thickness (T) | Less than or equal to 1.5mm. |
| Acquisition Field of View (FOV) | Consistent with baseline. |
| Image Header | The Acquisition Device shall record actual Field of View, Scan Duration, Scan Plane, Total Collimation Width, Single Collimation Width, Scan Pitch, Tube Potential, Tube Current, Rotation Time, Exposure and Slice Width in the DICOM image header. |

## 3. Image Data Reconstruction

3.1 Discussion

Image reconstruction is modeled as a separate Activity in the QIBA Profile.  Although it is closely related to image acquisition, and is usually performed on the Acquisition Device, reconstruction may be performed, or re-performed, separate from the acquisition.  Many reconstruction parameters will be influenced or constrained by related acquisition parameters.  This specification is the result of discussions to allow a degree of separation in their consideration without suggesting they are totally independent.

Many reconstruction parameters can have direct or indirect effects on identifying, segmenting and measuring lesions.  To reduce this potential source of variance, all efforts should be made to have as many of the parameters as possible consistent with the baseline.

Spatial Resolution quantifies the ability to resolve spatial details. Lower spatial resolution can make it difficult to accurately determine the borders of tumors, and as a consequence, decreases the precision of volume measurements.  Increased spatial resolution typically comes with an increase in noise which may degrade segmentation and quantification of tumors. Therefore, the choice of factors that affect spatial resolution typically represent a balance between the need to accurately represent fine spatial details of objects (such as the boundaries of tumors) and the noise within the image. Maximum spatial resolution is mostly determined by the scanner geometry (which is not usually under user control) and the reconstruction kernel (over which the user has some choice).  Resolution is stated in terms of “the number of line-pairs per cm that can be resolved in a scan of resolution phantom (such as the synthetic model provided by the American College of Radiology and other professional organizations)”. If a followup scan has a significantly different resolution than the baseline, it is likely that the exposure characteristics will change which can affect repeatability.  The impact of partial volume effects can also change, so reasonable consistency of resolution within a given patient is desirable.

Noise Metrics quantify the magnitude of the random variation in reconstructed CT numbers.  Increased levels of noise can make it difficult to identify the boundary of tumors by humans and automated algorithms.

Some properties of the noise can be characterized by the standard deviation of reconstructed CT numbers over a uniform region in phantom.  Voxel Noise (pixel standard deviation in a region of interest) can be reduced by reconstructing images with greater thickness for a given mAs. A constant value for the noise metric might be achieved by increasing mAs for thinner reconstructed images and reducing mAs for thicker reconstructed images.  The use of a standard deviation metric has limitations since it can vary with different reconstruction kernels, which will also impact the spatial resolution.  While the Noise-Power Spectrum would be a more comprehensive metric, it is not practical to calculate (and interpret) at this time.  Therefore, the standard deviation metric is the preferred measure for Voxel Noise.  It is not expected that the Voxel Noise be measured for each subject scan, but rather the Acquisition Device and Reconstruction Software be qualified for the expected acquisition and reconstruction parameters.

The voxel noise specification is defined in terms of a 20cm water phantom.  Achieving equivalent performance in alternate sized phantoms is permitted; however, the group utilizing the alternate phantom is obliged to demonstrate that the performance is equivalent to this specification.

Reconstruction Field of View affects reconstructed pixel size because the fixed image matrix size of most reconstruction algorithms is 512 X 512.  If it is necessary to expand the field of view to encompass more anatomy, the resulting larger pixels may be insufficient to achieve the claim. A targeted reconstruction with a smaller field of view may be necessary, but a reconstruction with that field of view would need to be performed for every time point. Pixel Size directly affects voxel size along the subject x-axis and y-axis. Smaller voxels are preferable to reduce partial volume effects and provide higher measurement precision.  Pixel size in each dimension is not the same as spatial resolution in each dimension. The spatial resolution of the reconstructed image depends on a number of additional factors including a strong dependence on the reconstruction kernel.

Reconstruction Interval (a.k.a. Slice spacing) that results in discontiguous data is unacceptable as it may “truncate” the spatial extent of the tumor, degrade the identification of tumor boundaries, confound the precision of measurement for total tumor volumes, etc.  Decisions about overlap (having an interval that is less than the nominal reconstructed slice thickness) need to consider the technical requirements of the clinical trial, including effects on measurement, throughput, image analysis time, and storage requirements.

Reconstructing datasets with overlap will increase the number of images and may slow down throughput, increase reading time and increase storage requirements.  For multidetector row CT (MDCT) scanners, creating overlapping image data sets has NO effect on radiation exposure; this is true because multiple reconstructions having different kernel, slice thickness and intervals can be reconstructed from the same acquisition (raw projection data) and therefore no additional radiation exposure is needed.

Slice thickness is “nominal” since the thickness is not technically the same at the middle and at the edges.

Reconstruction Kernel Characteristics influence the appearance of tumors in the reconstructed images which may influence measurements.  A softer kernel can reduce noise at the expense of spatial resolution. An enhancing kernel can improve resolving power at the expense of increased noise.  The characteristics of different tissues (e.g. lung) may call for the use of different kernels, and implementers are encouraged to use kernels suitable for the anatomic region and tissue imaged.  The use of multiple kernels in a single study is not prohibited by the specification below, but any given tumor must be measured on images reconstructed using consistent kernels at each time point.

The effects of iterative reconstructions on quantitative accuracy and reproducibility are not fully understood as of this writing of this profile version.

<<Consider a section on Window Level / HU “stability” and how it could affect the measurements.  And perhaps a specification that would help address.  Greg has raised this issue and submitted text.>>

3.2 Specification

The Reconstruction Software shall be capable of producing images that meet the following specifications. The Technologist shall set up or configure the reconstruction to achieve the requirements in the following table.

| Parameter | Specification |
| --- | --- |
| In-plane Spatial Resolution | Greater than or equal to 6 lp/cm and consistent with baseline (i.e. within 1 lp/cm). |
| Voxel Noise | Standard deviation of < 18HU measured near the center of a 20cm water phantom. |
| Reconstruction Field of View | Spanning at least the full extent of the thoracic and abdominal cavity, but not significantly greater than required to show the entire body and consistent with baseline. |
| Slice Thickness | Less than or equal to 2.5 mm and consistent with baseline (i.e. within 0.5mm). |
| Reconstruction Interval | Less than or equal to 2.5 mm and consistent with baseline. |
| Reconstruction Overlap | Greater than or equal to 0 (i.e. no gap, and may have some overlap) and consistent with baseline. |
| Reconstruction Algorithm Type | Consistent with baseline (i.e. do not mix Filtered Back-Projection and Iterative Reconstruction) |
| Reconstruction Kernel Characteristics | Consistent with baseline (i.e. the same kernel if available, otherwise the kernel most closely matching the kernel response of the baseline). This should reduce the chance the spatial resolution/noise specifications are met but the kernels produce very different textured images. |
| Image Header | The Reconstruction Software shall record actual Spatial Resolution, Noise, Pixel Spacing, Reconstruction Interval, Reconstruction Overlap, Reconstruction Kernel Characteristics, as well as the model-specific Reconstruction Software parameters utilized to achieve compliance with these metrics in the image header. |

## 4. Image Analysis

4.1 Discussion

This profile characterizes each designated tumor by its volume change relative to prior image sets.

This is typically done by determining the boundary of the tumor (referred to as segmentation), computing the volume of the segmented tumor and calculating the difference of the tumor volume <for the current and baseline><from one scan to the next?>.

Volume Calculation from a segmentation may or may not correspond to the total of all the segmented voxels.  The algorithm may consider partial volumes, do surface smoothing, tumor or organ modeling, or interpolation of user sculpting of the volume.  The algorithm may also pre-process the images prior to segmentation.

Segmentation may be performed automatically by a software algorithm, manually by a human observer, or semi-automatically by an algorithm with human guidance/intervention, for example to identify a starting seed point, stroke, or region, or to edit boundaries.

If a human observer participates in the segmentation, either by determining while looking at the images the proper settings for an automated process, or by manually editing boundaries, the settings for conversion of density into pixel values (window and level) should either be fixed during the segmentation process or documented so that observers can apply consistent display settings at future scans (or a different observer for the same scan, if multiple readers will read each scan, as for a clinical trial).

Methods that calculate volume changes directly without calculating volumes at individual time points are acceptable so long as the results are compliant with the specifications set out by this profile.

The Image Analysis Tool should be prepared to process both the current data and previous data at the same time and support matching up the appearance of each tumor in both data sets in order to derive volume change values.  Although it is conceivable that they could be processed separately and the results of prior processing could be imported and a method of automated tagging and matching of the tumors could be implemented, such interoperability mechanisms are not defined or mandated here and cannot be depended on to be present or used.

Storing segmentations and measurement results that can be loaded by an Image Analysis Tool analyzing data collected at a later date is certainly a useful practice as it can save time and cost.  For this to happen reliably, the stored format must be compatible and the data must be stored and conveyed.  Although DICOM Segmentation objects are appropriate to store tumor segmentations, and DICOM SR objects are appropriate to store measurement results, these standards are not yet widely enough deployed to make support for them mandatory in this Profile.  Similarly, conveying the segmentations and measurements from baseline (and other time points prior to the current exam) is not done consistently enough to mandate that it happen and to require their import into the Image Analysis Tool.  Managing and forwarding the data files may exceed the practical capabilities of the participating sites.

Image analysis can be performed on any equipment that complies with the specifications set out in this profile.  However, we strongly encourage performing all analysis for an individual subject on the same platform (manufacturer, model and version) which we expect will further reduce variation.

Medical Devices such as the Image Analysis Tool are typically made up of multiple components (the hardware, the operating system, the application software, and various function libraries within those).  Changes in any of the components can affect the behavior of the device.  In this specification, a “device version” reflects any changes to the components.  This may be different than the product release version that appears in vendor documentation.

In addition to the required Recording details, it can be helpful to record the identity of the operator.

<Revisit issue of requirements on user training/performance><And link to the 15% spec below.  The 15% includes the operator without extraordinary training><Easier if the SW is automated>

Determination of which tumors should be measured is out of scope for this Profile.  Such determination may be specified within a protocol or specified by formal response criteria standards, or may be determined by clinical requirements. They may be designated by the oncologist or clinical investigator, by a radiologist at a clinical site, by a reader at a central reading facility, or they may be designated automatically by a software analysis tool.

Comment-selection of the imaging and contrast protocol is typically at imaging site and does not belong in this section. I think we have covered this elsewhere.

4.2 Specification

| Parameter | Specification |
| --- | --- |
| Common Tumor Selection | If multiple readers will evaluate a single set of images, tThe Image Analysis Tool shall allow all tumors selected for volume measurement to be unambiguously labeled, so that all readers can assess the same tumors. |
| Multiple Tumors | The Image Analysis Tool shall allow multiple tumors to be measured, and each measured tumor to be associated with a human-readable identifier that can be used for correlation across time points. |
| Tumor Volume Change | The Image Analysis Tool shall measure tumor volume change (according to Figure 1) with a 95% confidence interval of less than +/- 15%. (with a 1% or less likelihood that a measurement is off by more than 15%?)[Set a bound for the bias and a bound for the variance] [precision?][95% confidence interval?][prediction limits] [asymmetric aspect of this] <<Add a requirement that the operator demonstrate the same performance target in their site (see method below) to claim conformance? Alternatively state thatfor the claim to be realized? Principle of tending toward specifying the result variable of interest rather than contributing factors that might not be contributing factors in some innovative designs. Don’t want to specify exact experience or training of the operator.>> |
| Result Verification | <<Have a human responsible for the result?>> |
| Recording | The Image Analysis Tool shall record the percentage volume change relative to baseline for each tumor, the device version and the actual model-specific Analysis Software set-up and configuration parameters utilized to achieve compliance with these metrics. The Image Analysis Tool shall be capable of recording the tumor segmentation as a DICOM Segmentation. The Image Analysis Tool shall record the identity of each individual making and/or approving a tumor measurement using the software. |

# IV. Compliance

To comply with this profile, participating staff and equipment shall support each of the activities assigned to them in Table 1.  Section III documents each activity states compliance requirements (“shall language”) for each Actor.

This section elabourates on the meaning of performance-oriented requirements in terms of how they are intended to be correctly assessed.

## 1. Performance Assessment: Tumor Volume Change Variability

## Compliance with the Tumor Volume Change Variability performance requirement can be assessed with the following procedure:

## Obtain the designated test image set from (???).  The test set is made up of images of (h) human subjects and includes a total of (n > 5) tumors.  (Note the test tumors have been selected to be within the performance range of the method, i.e. “larger than 10mm diameter”, and have a range of volumes)

## - Would basically have to be done on a phantom in order to have the TV (and delta TV); next step would be to insert digital tumors in real patient data.  At some

## Measure the volume of each of the (n) tumors at both time points (r) times, for a total of (n x r x 2) measurements.  The measurement may be performed by a single fully automated or (y) multiple human semi-automated or manual readers.

## Calculate the joint-distribution of Delta TV and Delta MV

## Also useful to take a set of human coffee break images and look at the variance between the measurement at the two time points.

## If you have the volume at T1 and T2 you not only need the variance of the individual measurements but you need the (sigma of the delta)**2 = 2(1-rho)sigma**2  i.e. we need to know the variance and the correlation between the two measurements.

## (Correlation across visits will be dominated by using a different device?)

## Stage one testing would be to see nicely narrow variance on coffee break data.

## Could also express as reporting what your confidence is that a change of size X is really more than Y.  ie. in the P(A|B)>C we can fix or “vectorize” any of the three variables?  Note that the values might be different for clinical trials vs patient management (so maybe two claims or maybe vector/curve)

## Might want to consider being able to drop a couple tumors if they Blow Up – but have to report the error

## (An issue is whether the algorithm or operator would clearly recognize the blow up or if it is just slightly off)

## Dan Barboriak has done work on related issues.

## Visual Analog Scale (VAS Score) might be a tool for ranking (drag a slider instead of pick a number) tumors and maybe have different variance targets for each category, or weight the results by the VAS.

## Should really include a characterization of the individual volume measurement performance since that is an input to a variety of models (and would be useful for patient enrichment in trials)

## For each tumor, average the (r) measurements and enumerate the number of measurements that are within 30% of the average.  If the total number of enumerated measurements across all tumors is greater than or equal to (z) (95% of n x r), the performance specification has been met.

## This procedure can be used to evaluate the performance of Image Analysis Tools, or Image Analysis Tools together with the Radiologist/Operator/Reader/Individual/Human Observer.

## <Alternatively, as suggested by TSB in comment #164, we might want the performance to be relative to a specified (expert consensus derived) “truth” value, but that is not really what the claim is (the claim is repeatability, not accuracy).>

## The designated test image set has been acquired according to this profile and contains designated tumors located in multiple body parts (i.e., not just lung but also mediastinal, liver, adrenal, neck, axillary, mesenteric, retroperitioneal, pelvic, etc.) and representing different solid tumor types (lung, breast, colon). <Such a public dataset does not currently exist as far as I know.>

## 1. Performance Assessment: Image Acquisition Site

<If we describe how to locally generate a reasonable test image set, the above procedure could be extended to assess the performance of the Image Acquisition Site>

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large patient population under consideration.  For imaging it is important to consider the availability of:

- appropriate imaging equipment and quality control processes,

- appropriate injector equipment and contrast media,

- experienced CT Technologists for the imaging procedure, and

- processes that assure imaging profile compliant image generation at the correct point in time.

A calibration and QA program shall be designed consistent with the goals of the clinical trial. This program shall include (a) elements to verify that sites are performing correctly, and (b) elements to verify that sites’ CT scanner(s) is (are) performing within specified calibration values. These may involve additional phantom testing that address issues relating to both radiation dose and image quality (which may include issues relating to water calibration, uniformity, noise, spatial resolution -in the axial plane-, reconstructed slice thickness z-axis resolution, contrast scale, CT number calibration and others). This phantom testing may be done in additional to the QA program defined by the device manufacturer as it evaluates performance that is specific to the goals of the clinical trial.

# References

[] Moertel CG, Hanley JA. The effect of measuring error on the results of therapeutic trials in advanced disease. Disease 1976; 38: 388-394.

[2] Quivey JM, Castro JR, Chen GT, Moss A, Marks WM. Computerized tomography in the quantitative assessment of tumour response. Br J Disease Suppl 1980; 4:30-34.

[3] Munzenrider JE, Pilepich M, Rene-Ferrero JB, Tchakarova I, Carter BL. Use of body scanner in radiotherapy treatment planning. Disease 1977; 40:170-179.

[4] Wormanns, D., Kohl, G., Klotz, E., Marheine, A., Beyer, F., Heindel, W., and Diederich, S. Volumetric measurements of pulmonary nodules at multi-row detector CT: In vivo reproducibility. Eur Radiol 14: 86–92, 2004.

[5] Kostis WJ, Yankelevitz DF, Reeves AP, Fluture SC, Henschke CI, Small Pulmonary Nodules: Reproducibility of Three-dimensional Volumetric Measurement and Estimation of Time to Follow-up CT, Radiology, Volume 231 Number 2, 2004.

[6] Revel M-P, Lefort C, Bissery A, Bienvenu M, Aycard L, Chatellier G, Frija G, Pulmonary Nodules: Preliminary Experience with Three-dimensional Evaluation, Radiology  May 2004.

[7] Marten K, Auer F, Schmidt S, Kohl G, Rummeny EJ, Engelke C, Inadequacy of manual measurements compared to automated CT volumetry in assessment of treatment response of pulmonary metastases using RECIST criteria, Eur Radiol (2006) 16: 781–790.

[8] Goodman, L.R., Gulsun, M., Washington, L., Nagy, P.G., and Piacsek, K.L. Inherent variability of CT lung nodule measurements in vivo using semiautomated volumetric measurements. AJR Am J Roentgenol 186: 989–994, 2006.

[9] Gietema HA, Schaefer-Prokop CM, Mali W, Groenewegen G, Prokop M, Pulmonary Nodules: InterscanVariability of Semiautomated Volume Measurements with Multisection CT— Influence of Inspiration Level, Nodule Size, and Segmentation Performance, Radiology: Volume 245: Number 3 December 2007.

[10] Wang Y, van Klaveren RJ, van der Zaag–Loonen HJ, de Bock GH, Gietema HA, Xu DM, Leusveld ALM, de Koning HJ, Scholten ET, Verschakelen J, Prokop M, Oudkerk M, Effect of Nodule Characteristics on Variability of Semiautomated Volume Measurements in Pulmonary Nodules Detected in a Lung Cancer Screening Program, Radiology: Volume 248: Number 2—August 2008.

[11] Zhao, B., James, L.P., Moskowitz, C.S., Guo, P., Ginsberg, M.S., Lefkowitz, R.A., Qin, Y., Riely, G.J., Kris, M.G., Schwartz, L.H. Evaluating variability in tumor measurements from same-day repeat CT scans of patients with non-small cell lung cancer. Radiology 252: 263–72, 2009.

[12] Hein, P.A., Romano, V.C., Rogalla, P., Klessen, C., Lembcke, A., Dicken, V., Bornemann, L., and Bauknecht, H.C. Linear and volume measurements of pulmonary nodules at different CT dose levels: Intrascan and interscan analysis. Rofo 181: 24–31, 2009.

[13] Mozley PD, Schwartz LH, Bendtsen C, Zhao B, Petrick N, Buckler AJ. Change in lung tumor volume as a biomarker of treatment response: A critical review of the evidence. Annals Oncology; doi:10.1093/annonc/mdq051, March 2010.

[14] Petrou M, Quint LE, Nan B, Baker LH. Pulmonary nodule volumetric measurement variability as a function of CT slice thickness and nodule morphology. Am J Radiol 2007; 188:306-312.

[15] Bogot NR, Kazerooni EA, Kelly AM, Quint LE, Desjardins B, Nan B. Interobserver and intraobserver variability in the assessment of pulmonary nodule size on CT using film and computer display methods. Acad Radiol 2005; 12:948–956.

[16] Erasmus JJ, Gladish GW, Broemeling L, et al. Interobserver and intraobserver variability in measurement of non-small-cell carcinoma lung lesions: Implications for assessment of tumor response. J Clin Oncol 2003; 21:2574–2582.

[17] Winer-Muram HT, Jennings SG, Meyer CA, et al. Effect of varying CT section width on volumetric measurement of lung tumors and application of compensatory equations. Radiology 2003; 229:184-194.

[18] Buckler AJ, Mozley PD, Schwartz L, et al. Volumetric CT in lung disease: An example for the qualification of imaging as a biomarker. Acad Radiol 2010; 17:107-115.

[19] AMERICAN COLLEGE OF RADIOLOGY IMAGING NETWORK, ACRIN 6678, FDG-PET/CT as a Predictive Marker of Tumor Response and Patient Outcome: Prospective Validation in Non-small Cell Lung Cancer, August 13, 2010.

[20] Miller AB, Hoogstraten B, Staquet M, Winkler A. Reporting results of cancer treatment. Cancer 1981;47:207-214.

[21] Eisenhauer EA, Therasse P, Bogaerts J, et al. New response evaluation criteria in solid tumors: Revised RECIST guideline (version 1.1). Eur J Cancer 2009;45:228-247.

[22] McNitt-Gray MF. AAPM/RSNA Physics Tutorial for Residents: Topics in CT. Radiation dose in CT. Radiographics 2002;22:1541-1553.

[23] Xie L, O'Sullivan J, Williamson J, Politte D, Whiting B,  TU‐FF‐A4‐02: Impact of Sinogram Modeling Inaccuracies On Image Quality in X‐Ray CT Imaging Using the Alternating Minimization Algorithm, Med. Phys. 34, 2571 (2007); doi:10.1118/1.2761438.

[24] Moertel CG, Hanley JA. The effect of measuring error on the results of therapeutic trials in advanced cancer. Cancer 38:388-94, 1976.

[25] Lavin PT, Flowerdew G: Studies in variation associated with the measurement of solid tumors. Cancer  46:1286-1290, 1980.

[26] Eisenhauera EA, Therasseb P, Bogaertsc J, et a. New response evaluation criteria in solid tumours: Revised RECIST guideline (version 1.1).  Eur J Cancer 2009; 45: 228-247.

[27] Boll, D.T., Gilkeson, R.C., Fleiter, T.R., Blackham, K.A., Duerk, J.L., and Lewin, J.S. Volumetric assessment of pulmonary nodules with ECG-gated MDCT. AJR Am J Roentgenol 183: 1217–1223, 2004.

[28] Meyer, C.R., Johnson, T.D., McLennan, G., Aberle, D.R., Kazerooni, E.A., Macmahon, H., Mullan, B.F., Yankelevitz, D.F., van Beek, E.J., Armato, S.G., 3rd, McNitt-Gray, M.F., Reeves, A.P., Gur, D., Henschke, C.I., Hoffman, E.A., Bland, P.H., Laderach, G., Pais, R., Qing, D., Piker, C., Guo, J., Starkey, A., Max, D., Croft, B.Y., and Clarke, L.P. Evaluation of lung MDCT nodule annotation across radiologists and methods. Acad Radiol 13: 1254–1265, 2006.

[29] Zhao, B., Schwartz, L.H., Moskowitz, C.S., Ginsberg, M.S., Rizvi, N.A., and Kris, M.G. Lung cancer: computerized quantification of tumor response--initial results. Radiology 241: 892–898, 2006.

[30] Zhao, B., Oxnard, G.R., Moskowitz, C.S., Kris, M.G., Pao, W., Guo, P., Rusch, V.W., Ladanyi, M., Rizvi, N.A., and Schwartz, L.H. A pilot study of volume measurement as a method of tumor response evaluation to aid biomarker development. Clin Cancer Res 16: 4647–4653, 2010.

[31] Schwartz, L.H., Curran, S., Trocola, R., Randazzo, J., Ilson, D., Kelsen, D., and Shah, M. Volumetric 3D CT analysis - an early predictor of response to therapy. J Clin Oncol 25: abstr 4576, 2007.

[32] Altorki, N., Lane, M.E., Bauer, T., Lee, P.C., Guarino, M.J., Pass, H., Felip, E., Peylan-Ramu, N., Gurpide, A., Grannis, F.W., Mitchell, J.D., Tachdjian, S., Swann, R.S., Huff, A., Roychowdhury, D.F., Reeves, A., Ottesen, L.H., and Yankelevitz, D.F. Phase II proof-of-concept study of pazopanib monotherapy in treatment-naive patients with stage I/II resectable non-small-cell lung cancer. J Clin Oncol 28: 3131–3137, 2010.

# Appendices

## Appendix A: Acknowledgements and Attributions

This document is proffered by the Radiological Society of North America (RSNA) Quantitative Imaging Biomarker Alliance (QIBA) Volumetric Computed Tomography (v-CT) Technical Committee. The v-CT technical committee is composed of scientists representing the imaging device manufacturers, image analysis software developers, image analysis laboratories, biopharmaceutical industry, academia, government research organizations, professional societies, and regulatory agencies, among others. All work is classified as pre-competitive.

A more detailed description of the v-CT group and its work can be found at the following web link: http://qibawiki.rsna.org/index.php?title=Volumetric_CT.

The Volumetric CT Technical Committee (in alphabetical order):

- Athelogou, M. Definiens AG

- Avila, R. Kitware, Inc.

- Beaumont, H. Median Technologies

- Borradaile, K. Core Lab Partners

- Buckler, A. BBMSC

- Clunie, D. Core Lab Partners

- Cole, P. Imagepace

- Dorfman, G. Weill Cornell Medical College

- Fenimore, C. Nat Inst Standards & Technology

- Ford, R. Princeton Radiology Associates.

- Garg, K. University of Colorado

- Garrett, P. Smith Consulting, LLC

- Gottlieb, R. University of Arizona

- Gustafson, D. Intio, Inc.

- Hayes, W. Bristol Myers Squibb

- Hillman, B. Metrix, Inc.

- Judy, P. Brigham and Women’s Hospital

- Kim, HG. University of California Los Angeles

- Kohl, G. Siemens AG

- Lehner, O. Definiens AG

- Lu, J. Nat Inst Standards & Technology

- McNitt-Gray, M. University California Los Angeles

- Mozley, PD. Merck & Co Inc.

- Mulshine, JL. Rush

- Nicholson, D. Definiens AG

- O'Donnell, K. Toshiba Medical Research Institute - USA

- O'Neal, M. Core Lab Partners

- Petrick, N. US Food and Drug Administration

- Reeves, A. Cornell University

- Richard, S. Duke University

- Rong, Y. Perceptive Informatics, Inc.

- Schwartz, LH. Columbia University

- Saiprasad, G. University of Maryland

- Samei, E. Duke University

- Siegel, E. University of Maryland

- Sullivan, DC. RSNA Science Advisor and Duke University

- Tang, Y. CCS Associates

- Thorn, M. Siemens AG

- Yankellivitz, D. Mt. Sinai School of Medicine

- Yoshida, H. Harvard MGH

- Zhao, B. Columbia University

The Volumetric CT Technical Committee is deeply grateful for the support and technical assistance provided by the staff of the Radiological Society of North America.

## Appendix B: Background Information

B.1 QIBA

The Quantitative Imaging Biomarker Alliance (QIBA) is an initiative to promote the use of standards to reduce variability and improve performance of quantitative imaging in medicine. QIBA provides a forum for volunteer committees of care providers, medical physicists, imaging innovators in the device and software industry, pharmaceutical companies, and other stakeholders in several clinical and operational domains to reach consensus on standards-based solutions to critical quantification issues. QIBA publishes the specifications they produce (called QIBA profiles), first to gather public comment and then for field test by vendors and users.

QIBA envisions providing a process for developers to test their implementations of QIBA profiles through a compliance mechanism. Purchasers can specify conformance with appropriate QIBA profiles as a requirement in requests for proposal. Vendors who have successfully implemented QIBA profiles in their products can publish QIBA Conformance Statements.  <Vendor Conformance Statement states…. and shall also provide at least one set of “Model-specific Parameters” (as shown in Appendix D) describing how to configure their products for alignment with the Profile.

General information about QIBA, including its governance structure, sponsorship, member organizations and work process, is available at http://qibawiki.rsna.org/index.php?title=Main_Page.

QIBA has constructed a systematic approach for standardizing and qualifying volumetry as a biomarker of response to treatments for a variety of medical conditions, including cancers in the lung (either primary cancers or cancers that metastasize to the lung [18]).

B.2 CT Volumetry for Cancer Response Assessment

Anatomic imaging using computed tomography (CT) has been historically used to assess tumor burden and to determine tumor response to treatment (or progression) based on uni-dimensional or bi-dimensional measurements. The original WHO response criteria were based on bi-dimensional measurements of the tumor and defined response as a decrease of the sum of the product of the longest perpendicular diameters of measured tumors by at least 50%. The rationale for using a 50% threshold value for definition of response was based on data evaluating the reproducibility of measurements of tumor size by palpation and on planar chest x-rays [24][25]. The more recent RECIST criteria introduced by the National Cancer Institute (NCI) and the European Organisation for Research and Treatment of Cancer (EORTC) standardized imaging techniques for anatomic response assessment by specifying minimum size thresholds for measurable tumors and considered other imaging modalities beyond CT. As well, the RECIST criteria replace longest bi-directional diameters with longest uni-dimensional diameter as the representation of a measured tumor [26]. RECIST defines response as a 30% decrease of the largest diameter of the tumor. For a spherical tumor, this is equivalent to a 50% decrease of the product of two diameters. Current response criteria were designed to ensure a standardized classification of tumor shrinkage after completion of therapy. They have not been developed on the basis of clinical trials correlating tumor shrinkage with patient outcome.

Technological advances in signal processing and the engineering of multi-detector row computed tomography (MDCT) devices have resulted in the ability to acquire high-resolution images rapidly, resulting in volumetric scanning of anatomic regions in a single breath-hold. Volume measurements may be a more sensitive technique for detecting longitudinal changes in tumor masses than linear tumor diameters as defined by RECIST. Comparative analyses in the context of clinical trial data have found volume measurements to be more reliable, and often more sensitive to longitudinal changes in size and thus to treatment response, than the use of a uni-dimensional diameter in RECIST. As a result of this increased detection sensitivity and reliability, volume measurements may improve the predictability of clinical outcomes during therapy compared with RECIST. Volume measurements could also benefit patients who need alternative treatments when their disease stops responding to their current regimens [29-32].

The rationale for volumetric approaches to assessing longitudinal changes in tumor burden is multi-factorial. First, most cancers may grow and regress irregularly in three dimensions. Measurements obtained in the transverse plane fail to account for growth or regression in the longitudinal axis, whereas volumetric measurements incorporate changes in all dimensions. Secondly, changes in volume are believed to be less subject to either reader error or inter-scan variations. For example, partial response using the RECIST criteria requires a greater than 30% decrease in tumor diameter, which corresponds to greater than 50% decrease in tumor volume. If one assumes a 21 mm diameter spherical tumor (of 4.8 cc volume), partial response would require that the tumor shrink to a diameter of less than 15 mm, which would correspond to a decrease in volume all the way down to 1.7 cc. The much greater absolute magnitude of volumetric changes is potentially less prone to measurement error than changes in diameter, particularly if the tumors are irregularly shaped or spiculated. As a result of the observed increased sensitivity and reproducibility, volume measurements may be more suited than uni-dimensional measurements to identify early changes in patients undergoing treatment.

Table B.1 Summarizing Precision/reproducibility of volumetric measurements from clinical studies reported in the literature

| Scan | Reader | # of Readers | # of Patients | # of Nodules | Tumor Size, Mean (range) | Organ System | Volumetry, 95% CI of Measurement Difference | Volumetry, Measurement Difference % | 1D Measurement, 95% CI of Measurement Difference | 1D, Mean Measurement Difference % | Slice Thickness /Recon Interval, mm | Author, Year |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| repeat scans | intra-reader | 1 | 20 | 218 | 9.85 mm | lung, mets | -21.2 to 23.8% | 1.30% |  |  | 1.0/0.7 | Gietama et al. 2007 [9] |
| repeat scans | intra-reader | 3 | 32 | 32 | 38 mm (11–93 mm) | lung, NSCLC | -12 to 13.4% | 0.70% | -7.3% to 6.2% | -0.60% | 1.25/1.25 | Zhao et al. 2009 [11] |
| same scan | intra-reader | 1 | 10 | 50 | 6.9 mm (2.2–20.5 mm) | lung, mets | -3.9 to 5.7% | 0.90% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [4] |
| same scan | inter-reader | 2 | 10 | 50 | 6.9 mm (2.2–20.5 mm) | lung, mets | -5.5 to 6.6% | 0.50% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [4] |
| repeat scans | not specified | not specified | 10 | 151 | 7.4 (2.2–20.5 mm) | lung, mets | -20.4 to 21.9% | 1.50% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [4] |
| repeat scans | not specified | not specified | 10 | 105 | <10 mm | lung, mets | -19.3 to 20.4% | 1.70% | not reported | not reported | 1.25/0.8 | Wormanns et al. 2004 [4] |
| same scan (5 sets, 1 set/phase) | intra-reader ? (consensus by 2 readers), 3 x reading | 2 | 30 | 73 | ~1–9 mm [25.3 (0.2–399 mm3)] | lung, noncalcified nodules | coefficient of variance as large as 34.5% (95% CI not reported) | not reported | not reported | not reported | 0.75/0.6 | Boll et al. 2004 [27] |
| same scan | inter-reader | 2 | 33 | 229 | 10.8 mm (2.8–43.6 mm), median 8.2 mm | lung, primary or mets | -9.4 to 8.0% | 0.70% | -31.0 to 27% | -2.00% | 1.0/0.8 | Hein et al. 2009 [12] |
| same scan | inter-reader, inter-algorithms (6 readers x 3 algorithms) | 6 | 16 | 23 | not reported | lung, nodules | 55% (upper limit) | not reported | not reported | not reported | 1.25/0.625 | Meyer et al. 2006 [28] |
| same scan | intra-reader | 2 | 50 | 202 | 3.16–5195 mm3, median 182.22 mm3 | lung, mets | % not reported | 0.15 to 0.22% | % not reported | 2.34–3.73% (p<0.05 1D vs 3D) | 0.75/0.70 | Marten et al. 2006 [7] |
| same scan | inter-reader | 2 | 50 | 202 | 3.16–5195 mm3, median 182.22 mm3 | lung, mets | % not reported | 0.22 to 0.29% | % not reported | 3.53–3.76% (p<0.05 1D vs 3D) | 0.75/0.70 | Marten et al. 2006 [7] |
| same scan | inter-reader | 2 | 2239 | 4225 | 15–500 mm3 (effective diameter 3.1–9.8 mm) | lung, nodules | -13.4 to 14.5% | 0.50% | not reported | not reported | 1.0/0.7 | Wang et al. 2008 [10] |
| same scan | intra-reader | 2 | 24 | 52 | 8.5 mm (<5 to 18 mm) | lung, noncalcified nodules | 8.9 % (upper limit) | not reported | not reported | not reported | 1.25 or 2.5/not specified | Revel et al.[6] |
| same scan | inter-reader (3 readers x 3 measurements) | 3 | 24 | 52 | 8.5 mm (< 18 mm) | lung, noncalcified nodules | 6.38 % (upper limit) | not reported | not reported | not reported | 1.25 or 2.5/not specified | Revel et al. [6] |

Abbreviations: 1D = unidimensional; mets = metastasis; CI = confidence interval

The above table provides a basis for the 30% value in the Profile Claim.  The range between the minimum and maximum values in the 95% CI of the measurement difference column is mostly within +/- 15%. Considering a large study from Wang et al using 2239 patients [10], the 95% confidence interval ranged [-13.4%,  14.5%]. Thus, 30% is a conservative threshold of measurement variation. For example, the 30% change in the claim is the outside of 95% confidence interval of 15% of measurement variability when sample size is 40 or more.

B.3 Measurement Variability

This section discusses sources of measurement variability and the related terms used in this document. 
<Or should we just add definitions to Appendix C without a surrounding narrative.>

<Here’s some strawmen pieces>

<Tumor Definition Variability – refers to the clarity of the tumor boundary in the images.  It originates from the biological characteristics of the tumor, characteristics of the imaging process, and perhaps on the perception, expertise and education of the operator.  It seems to me that this is out of the scope of the profile, even though imaging biomarkers have probably different performances regarding this variability>

<Technical variability - originates only from the ability in drawing unequivocal objects. In other words, the perception of tumors definition is supposed absolutely clear and similar for any given operators when attempting to assess “Technical” variability.>

<Global Variability - partitioned as the variability in the tumor definition plus the “Technical” variability.>

Intra-Rater Variability is the variability in the interpretation of a set of images by the same reader after an adequate period of time inserted to reduce recall bias.

Inter-Rater Variability is the variability in the interpretation of a set of images by the different readers.

<Repeatability – considering multiple measurements taken under the same conditions (equipment, parameters, reader, algorithm, etc) but different subjects>

<Reproducability – considering multiple measurements taken where one or more conditions has changed>

## Appendix C: Conventions and Definitions

Acquisition vs. Analysis vs. Interpretation: This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be computational steps that transform the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge. (The transformation of knowledge into wisdom is beyond the scope of this document.)

Image Analysis, Image Review, and/or Read: Procedures and processes that culminate in the generation of imaging outcome measures, such tumor response criteria. Reviews can be performed for eligibility, safety or efficacy. The review paradigm may be context specific and dependent on the specific aims of a trial, the imaging technologies in play, and the stage of drug development, among other parameters.

Image Header: The Image Header is that part of the file or dataset containing the image other than the pixel data itself

Imaging Phantoms: Devices used for periodic testing and standardization of image acquisition. This testing must be site specific and equipment specific and conducted prior to the beginning of a trial (baseline), periodically during the trial and at the end of the trial.

Intra-Rater Variability is the variability in the interpretation of a set of images by the same reader after an adequate period of time inserted to reduce recall bias.

Inter-Rater Variability is the variability in the interpretation of a set of images by the different readers.

A Time Point is a discrete period during the course of a clinical trial when groups of imaging exams or clinical exams are scheduled.

## Appendix D: Model-specific Instructions and Parameters

For acquisition modalities, reconstruction software and software analysis tools, profile compliance requires meeting the activity specifications above; e.g. in Sections 2, 3 and 4.

This Appendix provides, as an informative tool, some specific acquisition parameters, reconstruction parameters and analysis software parameters that are expected to be compatible with meeting the profile requirements.   Just using these parameters without meeting the requirements specified in the profile is not sufficient to achieve compliance.  Conversely, it is possible to use different compatible parameters and still achieve compliance.

These settings were determined to be reasonable by the QIBA CT 1C groundwork study team.

Sites using models listed here are encouraged to consider using these parameters for both simplicity and consistency. Sites using models not listed here may be able to devise their own settings that result in data meeting the requirements.

Table D.1 Model-specific Parameters for Acquisition Devices

IMPORTANT NOTE: The presence of a product model/version in the table does not imply it has demonstrated compliance with the QIBA Profile.  Refer to the QIBA Conformance Statement for the product.

| Acquisition Device | Settings Compatible with Compliance |
| --- | --- |
| GE Discovery HD750 sct3 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.625 Gantry Rotation Time in seconds 1 mA 120 Pitch 0.984 Scan FoV Large Body (500mm) |
| Philips Brilliance 16 IDT mx8000 | kVp 120 Number of Data Channels (N) 16 Width of Each Data Channel (T, in mm) 0.75 Gantry Rotation Time in seconds 0.75 Effective mAs 50 Pitch 1.0 Scan FoV 500 |
| Philips Brilliance 64 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.625 Gantry Rotation Time in seconds 0.5 Effective mAs 70 Pitch 0.798 Scan FoV 500 |
| Siemens Sensation 64 | kVp 120 Collimation (on Operator Console) 64 x 0.6 (Z-flying focal spot) Gantry Rotation Time in seconds 0.5 Effective mAs 100 Pitch 1.0 Scan FoV 500 |
| Toshiba Aquilion 64 | kVp 120 Number of Data Channels (N) 64 Width of Each Data Channel (T, in mm) 0.5 Gantry Rotation Time in seconds 0.5 mA 25 Pitch .828 Scan FoV Medium and Large |

Table D.2 Model-specific Parameters for Reconstruction Software

IMPORTANT NOTE: The presence of a product model/version in the table does not imply it has demonstrated compliance with the QIBA Profile.  Refer to the QIBA Conformance Statement for the product.

| Reconstruction Software | Settings Compatible with Compliance |
| --- | --- |
| GE Discovery HD750 sct3 | Reconstructed Slice Width, mm 1.25 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel STD |
| Philips Brilliance 16 IDT mx8000 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm (contiguous) Display FOV, mm 350 Recon kernel B |
| Philips Brilliance 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm (contiguous) Display FOV, mm 350 Recon kernel B |
| Siemens Sensation 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel B30 |
| Toshiba Aquilion 64 | Reconstructed Slice Width, mm 1.00 Reconstruction Interval 1.0mm Display FOV, mm 350 Recon kernel FC11 |

Table D.3 Model-specific Parameters for Image Analysis Software

IMPORTANT NOTE: The presence of a product model/version in the table does not imply it has demonstrated compliance with the QIBA Profile.  Refer to the QIBA Conformance Statement for the product.

| Image Analysis Software | Settings Compatible with Compliance |
| --- | --- |
| Siemens LunCARE | a <settings to achieve…> b <settings to achieve…> c <settings to achieve…> d <settings to achieve…> |
| GE Lung VCAR | e <settings to achieve…> f <settings to achieve…> g <settings to achieve…> h <settings to achieve…> |
| R2 ImageChecker CT Lung System | i <settings to achieve…> j <settings to achieve…> k <settings to achieve…> l <settings to achieve…> |
| Definiens (name specific product) | m <settings to achieve…> n <settings to achieve…> o <settings to achieve…> p <settings to achieve…> |
| Median (name specific product) | q <settings to achieve…> r <settings to achieve…> s <settings to achieve…> t <settings to achieve…> |
| Intio (name specific product) | u <settings to achieve…> v <settings to achieve…> w <settings to achieve…> x <settings to achieve…> |

# Comments

95%? 2x StdDev of CoV

Can we replace this sentence with a note: 
“Refer to relevant clinical trial protocol for further guidance or requirements on timing relative to index intervention activity.”

Do we have evidence that repeatability of these is a critical factor? Most of our references likely have consistent use due to being single site.  Delay may be more significant than rate, but neither are well recorded or conveyed.

Could keep them as requirements and let people do “loose” implementation of the profile at their own discretion and maybe relax the profile based on data in the future.

COMMENT FROM NEIL

3. SO LONG AS THE CONTRAST ADMINISTRATION DATA IS RECORDED, WE REALLY SHOULDN'T CARE WHETHER IT IS IN THE DICOM HEADER OR PROVIDED ON A FORM VIA SCANNING AND SECONDARY CAPTURE.

What did we decide about dose modulation? Currently the spec allows it to be on for some scans and off for others.

Investigate and if these are simply not widely available, consider dropping them or just making them recommended.  They might not be that important since we have the reconstructed slice thickness and

Should this specify two timepoints so we are talking volume change not volume.

Do we require each repeat is by a different operator if there is operator interaction?  Do we require the same number of repeats if the measurement is fully automated?

Need to add people:

Gregory Goldmacher

Mike Silver,

…

Review phrasing with PET and MRI group (and Metrology Cmte)