## Sheet1
Reference | Setting | Reader | # Subjects | # Nodules | Nodule sizes | CT Scanner | Scan number/time | kVp | effective mAs | mAs | mA | Scan time (sec) | Pitch | Detector collimation (mm) | Section thickness (mm) | Section Increment (mm) | Kernel | FOV (mm) | Software | Segmentation | Scan-rescan volume difference | 95% CI | Comments
Wormanns D et al. Eur. Radiol. 2004; 14:86-92 | Metastases | Intra-reader | 10 | 151 | 7.4±4.5 mm (2.2-20.5 mm) 105 nodules < 10 mm | Siemens VZ 4 | 2 within 10 minutes after getting up | 120 | 20 |  |  | 0.5 sec | 1.75 | 1 | 1.25 | 0.8 | "Standard lung reconstruction kernel" | 274-383 | Siemens LungCare | Automated | 0.7±10.6% (-41-29%) | -20.4 to 21.9% | Limits of agreement for 105 nodules ≤10 mm=-19.3% (95% CI -22.6 to -16.0%) to 20.4% (95% CI 17.1-23.7%) | Limits of agreement are variation from mean volume of the two time points (Bland-Altman)
Goodman L et al. AJR 2006; 186:989-994 | Known and unsuspected pulmonary nodules | Avg 3 readers | 29 | 43 | < 20 mm; 35 (81%) ≤10.5 mm; Volume 345.5±361.1 mm^3 | GE Lightspeed 8 or 16 | 3 separated by 10-20 min pause | 120 | X |  | 200-400 | 0.5-0.8 | 1.35-1.375 | X | 1.25 | X | Bone | X | GE Advantage Windows workstation | Automated | X | ±25.6% | IV contrast in 8 for initial scan; <6mm(n=13) ±27% 6 to <9 mm (n=16) ±31.4% 9-19 mm (n=14) ±17.9% | Nice graph nodule size vs confidence limits; larger CL for contrast-enhanced nodules | CL are variation from mean volume of three time points (Bland-Altman)
Bolte H. Invest Radiol 2006;41:28-35 | Ex vivo pig lung phantom | Consensus 2 readers | x | 54 synthetic | Mean 7.8±2.1 mm; range 1.2-13.3 mm(50% 6.7-8.9 mm) | Siemens Sensation 16 | 5 repeated scans after deflation and reinflation | 120 | 50 | x | x | 0.5 | 1.25 | 0.75 | 1 | 0.7 | "intermediate reconstruction kernel" | 350 | Siemens LungCare | Automated without and with manual correction | SD=9.2% Automated, 3.7% corrected | Up to 36.8% automated (-18 to+18%), 14.8% corrected | "mean variation coefficient of 9.2% means that 2 measurements may vary about 37% from each other" | Authors propose 50% volume change threshold for definite growth | Only well-demarcated, rounded, solid nodules included (others discarded)
Gietema H et al. Radiology 2007; 245:888-894 | Metastases | Intra-reader | 20 | 218 | Volume 123.0±101.9 mm^3 (16.4-472.7) | Philips MX8000 16 | 2 scans after getting up from table | 120 (≤80 kg) or 140 (>80 kg) | X | 30 | X | X | X | X | 1 | 0.7 | B ("moderately soft") | Outer rib margins | Siemens LungCare | Automated | 2.4mm^3 (1.3%) (-53-120.8 mm^3) | -21.2 to 23.8% | 106 completely segmented: -11.9 to 12.4% Incompletely segmented: -26.8 to 30.0% | CIs are variation from mean volume of the two time points (Bland-Altman)
deHoop B. Eur Radiol 2009; 19:800-808 | Metastases | Intra-reader | 20 | 214 | Mean 10.9 mm (range 3.3-30.0) | Philips MX8000 16 | 2 within 10 minutes after getting up | 120 or 140 depending on wt. | X | 30 | X | X | X | X | 1 | 0.7 | "moderately soft kernel" | Outer rib margins | Compare 6 programs: GE, Philips, Siemens LungCare I and II, Mevis, Vital | Automated, some with manual correction | X | Upper limit of 95% CI for 89 nodules adequately segmented by all 6 programs: 19.2, 16.4, 22.3, 17.8, 20.0, 19.6% | Upper 95% CI for all nodules with some manual corrections: 19.2, 17.8, 23.8, 17.7, 20.0, 21.3% | CIs are variation from mean volume of the two time points (Bland-Altman) | No correlation between lung and nodule volume difference but trend toward smaller nodule measurement with larger lung volume (lung volume differences were only 10-15%)
Marchiano A. Radiology 2009; 251:919-925 | Lung Screening | Intra-reader | 101 | 233 | 99.1±127.5 mm^3 (median 67, range 5-839) | Siemens Sensation 16 | 3 (baseline, 3 and 12 months) | 120 | 30 | X | X | 0.5 | 1.5 | 0.75 | 1 | 1 | B30f and B50f | 300-400 | Siemens LungCare | Automated | <10% in 70.4% (3 mos) and 61.8% (12 mos) | ±27% | Lack of growth confirmed at 2 yrs. Kernel used for segmentation not specified. | Repeatability coeff 0.27 | CIs are variation from mean volume of the two time points (Bland-Altman)
Rampinelli C. AJR 2009; 192:1657-1661 | Known indeterminate pulmonary nodules | Intra-reader | 66 | 83 | 5-10 mm | GE Lightspeed 16 | 2 without getting up | 140 | X (18.2 calc) | 32 | 40 | 0.8 | 1.75 | X | 2.5 | X | "lung reconstruction kernel" | 320-380 | GE ALA Single | Automated | X | -38 to 60% | Excluded nodules incompletely segmented or attached to vessels or pleura | Coeff of reproducibility=1.63 | CIs are 'percentage variation' which is not defined in methods | No influence of nodule shape
Rampinelli C. AJR 2009; 192:1657-1661 | Known indeterminate pulmonary nodules | Intra-reader | 66 | 83 | 5-10 mm | GE Lightspeed 16 | 2 without getting up | 120 |  | X | 100-440 (ACM) | X | 0.938 | X | 0.625 | X | "lung reconstruction kernel" | 320-380 | GE ALA Single | Automated | X | -27 to 40% | Excluded nodules incompletely segmented or attached to vessels or pleura | Coeff of reproducibility=1.40 | CIs are 'percentage variation' which is not defined in methods
Gavriledes MA. Acad Radiol 2013;20:1364-1370 | Synthetic phantom | Not applicable | x | 16 synthetic | 5, 8, 9, 10 mm diameter x 4 shapes (sphere, ellipse, lobulated, spiculated) | Philips MX8000 16 | 10 repeat | 120 |  | 20.8, 83.3, 166.7 | x | x | 1.2 | .75, 1.5 | .8, 1.5, 3 (0.75 collimation) and 2, 3, 5 (1.5 collimation) | x | Detail (b40f) (?-for Philips scanner?)? | x | Custom | Matched filter-based estimation |  | Min detectable increase (95% confidence) 17 - 19 - 15%; min detectable decrease 18 - 15 - 15% | 5 - 8 - 9 mm thin/thick combined: Min detectable increase (95% confidence) 45 - 21 - 16%; min detectable decrease 20 - 16 - 15% | No effect of shapes