## 28Apr2017
| Name of reviewer: | Combined with Form Site 09Mar2017
| Document correlate: | QIBA_Amyloid_PET_28Apr2017
| Section/Line | Type of Revision | Suggested Revision | ACTION TAKEN/SUGGESTED | Next Step
| CLAIM Disclaimers (149) | REVIEW CLAIM & Disclaimers | We will need to review this in great detail as the statements are not clearly understood and may not be accurate for all radiotracers | Need Radiotracer manufacturer Review to identify any open issues | Re-request feedback during Public Comment phase
AS | 195-196 | Addition of reference | Are we certain that no TOF or PSF recons were done in these studies? "...such as PSF (point spread function) based reconstruction or TOF (time of flight) imaging that were not utilized in published test-retest studies. " | Refer to Rathan - was literature reviewed using TOF and/or PSF or not. May need to revise text depending upon result. Separate there is poster indicating that did not affect it much. | Reassess for closure during Public Comment phase
| Sec 3.2.1.4 Line 472 | Addition of text | Added kVp > 80 recommendation informative; ?Add to normative? | Added infomative text recommending 80 kVp minimum; NOT added to normative text. | RESOLVED
| Sec 3.3.1 Line 509 Table |  | Is PET matrix voxel size of less than or equal to 2.5 mm for x and y axeis and less than or equal to 3 mm for z-axis acceptable for all radiotracer manufacturer's | Check with radiotracer mfr. | Re-request feedback during Public Comment phase
AS | 3.3.2.2.1 | Minor (< 3 sentence) rewrite | Add recommended thresholds for translation and rotations (the x and y in the table) when available. | Inserted acceptable threshold of less than or equal to 4 mm or 4 degrees translation/rotation otherwise need to remove data as per results of Dr. Matthew's NIBIB Rd 5 report | RESOLVED
AAL | 3.4.3.1 Line 728 | Minor (< 3 sentence) rewrite | The profile is very specific on many details, but grey matter segmentation is mentioned as a standard tool. However, it is well known that segmentation results are dependent on the MRI sequence and even the same sequence produces different results on different MRI scanners. | Inserted informative text to acknowledge this. No normative text required at this time. Is there an OPEN Issue here? | Reassess for closure during Public Comment phase
AS | line 945 | Minor (< 3 sentence) rewrite | Replace x% when value is known. | Is % needed or just delete? Consider checking paper or refer to one of the authors if needed. | Re-request feedback during Public Comment phase
AS | 3.6.4.1 Uniformity QC | Major (> 3 sentence) rewrite | In normative text, do we need to give more details about HOW to derive the transaxial and axial uniformity measurements? In other words, specify the type (size, shape), placement and formula for the uniformity calculation? | Ask PET physics subgroup to address this point. | If needed, then may belong in Section 4
| Line 1120 & Appendix I (Line 1996) |  | What should the precision be for the SUVR measurement on a Hoffman phantom scan and which regions (target and reference should be used to do so?) | Use 10% currently; to be re-evaluated when Conformance testing methodology developed | RESOLVED
AS | 4.1 PET Scanner Calib. | Major (> 3 sentence) rewrite | PET Scanner Calibration row: give more details about how to calculate the slice-to-slice variability (perhaps in normative text)? Better definition of end planes as well: just one end plane on each side, or up to 5 end planes on each side? | Ask PET physics subgroup to address this point. | If needed, then may belong in Section 4
| Sec 4.3 Line 1287 Normative -Data Recon 2D/3D |  | If 3D mode data can be re-binned into 2D mode, shall be able to perform reconstruction of data acquired in 3D mode using 2D image | Text inserted in normative | RESOLVED
AS | Section 4.4 | Major (> 3 sentence) rewrite | ROI saving/retrieve row, gray portion: remove the reference to tumor, and rewrite so it makes sense for PET amyloid imaging. | Section 4.4 rewritten in expectation of DRO | RESOLVED
NO | section 4.4 Table Analysis Accuracy | Addition of text | The Profile says, "For each of the specified ROIs in the DRO (Appendix F) the correct SUV values shall be replicated by the Analysis Workstation." It looks like an actor needs to get the exact SUVR value on the ROIs in order to pass the conformance test. Is this correct? Is that too stringent? Other groups are requiring more of a test sample where the actor's overall performance on the sample meets the claim statement (i.e. get a RC <12%). Also, there is some discussion that the conformance test should not be based on the performance of just one reader/technician, but that multiple readers using the actor's software can get the right answer. | Refer to Image Analysis Conformance group - undergoing complete rewrite. | OPEN ITEM: To be addressed with development of Conformance Methodology
| Section 4 | major | I don't see anything in here yet about testing actors' performance to see if they meet the claim. For example, we need to describe a study that actors need to perform to test that 1. their wCV is =<0.043, 2. that the wCV is constant over the range of SUVR, and 3. that linearity with a slope of one is a reasonable assumption. The first two should be tested with clinical data and linearity is usually best assessed in a phantom study, or equivalent. I am happy to help with writing the statistical tests for this section. | For group discussion. Likely a future deliverable. | OPEN ITEM: To be addressed with next version of Profile
AS | References 1301 | Major (> 3 sentence) rewrite | References need to be edited, updated and referenced properly in the main body of the Profile. | Agreed; to be done - REFERENCES ADD Radiotracer label links to REFERENCES | ACTION: REFERENCES as an OPEN ISSUE for Public Comment Version
AS | Appendix A | Major (> 3 sentence) rewrite | Add QIBA PET Amyloid BC member list | Done by RSNA Staff | COMPLETED
AS | Appendix B | Major (> 3 sentence) rewrite | Appendix B: Background Information for Claim needs to be filled in. | Text created - taken from Dr. Subramaniam's meta-analysis | RESOLVED
AS | Appendix C | Major (> 3 sentence) rewrite | Appendx C, Definitions: need to be sure all acronyms used in the body of the Profile are referenced here. Many are captured in the comments above. In particular, remove the language that is specific for the PET FDG Tumor Profile. | Done by RSNA Staff | COMPLETED
| Appendix E |  | NEEDS revision specific to Amyloid rather than FDG-PET | Revision performed | RESOLVED
| Appendix F Line 1785 |  | We need to add the following conformance criteria to the table: 1. The CI for beta_2 must contain zero. 2. The CI for A_1 must be contained in the following interval: 0.95 – 1. | Revision performed | COMPLETED
| Line 52 Open Issues |  | If and how to address concern that large (>8% change in SUVR across time) may be accounted for by biologic change unrelated to amyloid deposition difference alone. | If I understand correctly, you believe that the 8% change is too small to account for natural biologic changes with aging. One possible solution would be to write the claim with different criteria, depending on the amount of time between the two measurements. For example, if the two measurements are <6 months apart, then a change of 8% or larger represents true change, but if the elapsed time is >6 months, then maybe a change >10% represents true change. | OPEN ITEM: To be addressed with next version of Profile
| Section 4 Line 1251 |  | There are still some original lines that were debated at length during the development of the FDG-PET Profile. I think the intent is the same as the FDG-PET BC for this section. | Use the soon-to-be-finalized wording from the corresponding section in the FDG-PET Profile. | Perform this review during Public Comment Phase
SD | GENERAL |  | Critically review and revise wording to distinguish the approved clinical 'subjective' analysis for purposes of clinical practice distinct from the focus of this Profile (not validated by regulatory agency's). See line 133. | DONE - Lines 229/230 | COMPLETED
| GENERAL | Major (> 3 sentence) rewrite | SUV measurement and it's dependencies - Review for need for SUV measurement throughout document |  | COMPLETED
DM | SHUFFLE | Rearrangement of paragraphs | moving the PVE Correction and Smoothing paragraphs into the Post-Processing section, just prior to the Analysis section. | Need to look at this. | Perform this review during Public Comment Phase
Absolute Quantification
Atrophy considerations

## Sheet2
Grammar/Typo
Deletion of text
Addition of text
Minor (< 3 sentence) rewrite
Major (> 3 sentence) rewrite
Rearrangement of paragraphs
Rearrangement of sections
Reference change
Addition of reference