## Instructions
IMPORTANT: Do not edit this template spreadsheet.
FIRST: Use File->Make a copy... to copy this template into a new file for your comments
UPDATE the file name (in the upper left corner of your new sheet window) which says "Copy of ..." to include your Profile Name and remove Copy of
NOW you can edit your copied document.
See http://qibawiki.rsna.org/index.php/Public_Comment_Process for more guidance on the comment resolution process
If you are a commenter submitting comments, fill in columns A-F.
If you are a Biomarker Committee (BC) resolving comments, update columns G-J. (You can tidy up columns A-F as needed.)

## Results and Resolutions
# | Commenter | Email | Section | Line # | Priority | Issue | Proposal | Committee Discussion | Owner (opt.) | Resolution (w Rationale if rejected) | Status
[10 empty] | 0 | TBD | To be discussed
[10 empty] | 1 | OK | No action requested
| IMPORTANT: Read the Instructions Tab FIRST | [8 empty] | 0 | Discuss | Need to decide resolution
[10 empty] | 0 | TODO | Resolution decided
[10 empty] | 13 | Done | Profile update completed
|  |  |  | Use line# of PC PDF, they're stable |  | Describe your issue. Don’t write a book. Do include enough so it's clear what you see as a problem. | Suggest new wording or describe a way to address your issue. The committee may simply accept your suggested text. Even if they don’t, it gives a good sense of what you’re looking for. Leaving this blank means you can’t imagine how to resolve the issue | Cmte Notes as needed. | BC member who will lead resolution of this issue <Delete this column if not used> | How the comment was/will be resolved. BC may simply accept & change as proposed, may accept issue & resolve differently, or may reject issue (with a rationale for why).
1 | Nancy Obuchowski | obuchon@ccf.org | 1.2.1 | 184-188 | Medium | The paragraph starting with "Based on the groundwork studies..." seems to be describing a longitudinal scenario whereby the "comparison flow measurement" could be a follow-up measurement. If the "comparison flow measurement" is made without error, then the confidence statement is ok. If, however, the "comparison flow measurement" also has measurement error associated with it, then the confidence range of +/-20% is inadequate. | If there are two measurements being taken, then multiple the wCV by 2.77 (instead of 1.96). | Nancy's proposed change : If there are two measurements being taken, then multiply the wCV by 2.77 (instead of 1.96). BC decided on proposed revised wording for this section that was shared with Nancy O. for feedback. Wording finalized from consultation with Nancy O. BC reviewed and approved this wording. |  | Wording changed to read “Based on the groundwork studies and literature review carried out by the QIBA USVBF Biomarker Committee, consider the following clinical scenario: Given initial and subsequent mean flow estimates, each with a confidence of +/- 20% of the mean value (coefficient of variation), the subsequent mean will be considered different from the initial mean based on 95% confidence limits if it lies outside of the repeatability range (i.e. if subsequent mean < (initial mean ± (0.20*initial mean)*2.77) < subsequent mean). Either pulsatile or constant flow can be used for the assessment of whether each type of flow measurement meets QIBA claims.” In addition, deleted the text “To put the above Claims in perspective, consider a site that is not conforming to the requirements in the QIBA Profile or making similar special image acquisition efforts.” | Done
2 | David T. Fetzer, MD | david.fetzer@utsouthwestern.edu | 3.3 | 295 | Medium | Can a site be conformant without phantom testing (Section A.2 and A.4) | Specify that either these requirements or option, or have CLEAR pathway for obtaining phantoms for sites that might not have resources to build or purchase their own | Proposed response: At several places in the profile, e.g. 4.3.1. and 4.4.1, the following statement is made. “A phantom may be available to a site for this purpose through a loan if resources are limited. “ This point will be reiterated elsewhere in the profile. Need to identify any additional locations? BC reviewed the proposed response, amended it to the final wording and determined the current locations were sufficient. |  | At several places in the profile, e.g. 4.3.1. and 4.4.1, the following statement is now made. “A phantom may be available to a site for this purpose through a loan or rental program or as part of a QC contract if resources are limited. “ | Done
3 | David T. Fetzer, MD | david.fetzer@utsouthwestern.edu | 3.3 | 295 | Low | In Periodic QA, must Periodic QA (Section A.5) be by AIUM or IEC guidelines, or can another accrediting body (ACR; IAC) be used? | Refrain from calling out specific organizations | Proposed response: Text modified as follows with the text in red indicating new text “Shall perform standard ultrasound pulse echo imaging QA on the Ultrasound Scanner as specified by applicable guidelines of accrediting organizations (e.g., ACR, AIUM, IAC, IEC).” Need to determine if the list here is adequate. At the BC discussion the only additional organization suggested was AAPM. There was discussion of international organizations in addition to the IEC but listing those for every region would have resulted in a protracted list. Since these were just examples then the list was adjusted to the final wording and approved. In addition, the names were added to the abbreviations list. |  | Profile text modified to read “Shall perform standard ultrasound pulse echo imaging QA on the Ultrasound Scanner as specified by applicable guidelines of accrediting organizations (e.g., AAPM, ACR, AIUM, IAC, IEC).” Names were added to the abbreviations list. | Done
4 | David T. Fetzer, MD | david.fetzer@utsouthwestern.edu | 4.2.2 | 465 | Low | Is this referring to phantom-level or patient-level QA? What is "significant change"? | Clarify | As this was a low priority comment the response was composed by the editor in consultation with the Profile Task Group as follows with the text in red indicating new text. Section 4.2.2 Periodic QA is testing performance in a phantom. "Significant change" is described in Lines 471-473 stating that “Each measurement would be compared to Claim 1 that specifies the 95% confidence limit for profile conformance. Additional testing may be required if expected performance does not match Claim 1 and would be considered a significant change in system performance.” To further clarify this, the term “significant change” does appear in the table in A.5.2 and 3.3 so the following text has been provided “If a significant change (as defined in Section 4.2.2 Periodic QA) is seen in results with a given US system, it is necessary to obtain independent verification of whether the phantom or the system has changed.“ |  | "Significant change" is described in Lines 471-473 now stating that “Each measurement would be compared to Claim 1 that specifies the 95% confidence limit for profile conformance. Additional testing may be required if expected performance does not match Claim 1 and would be considered a significant change in system performance.” To further clarify this, the term “significant change” now appears in the table in A.5.2 and 3.3 so the following text has been provided “If a significant change (as defined in Section 4.2.2 Periodic QA) is seen in results with a given US system, it is necessary to obtain independent verification of whether the phantom or the system has changed.“ | Done
5 | David T. Fetzer, MD | david.fetzer@utsouthwestern.edu | 3.4 | 300 | Low | A "DO" is also considered a physician. Your current language says "Physician... or other qualitifed medical personal, such as a PA or DO..." | Remove "DO". Include "PA, NP, or other delegated physician extender" | As this was a low priority comment the text was revised by the editor in consultation with the Profile Task Group to read “The role of the Sonographer/Technologist may also be taken by the Physician or other qualified medical personnel, such as PA, NP or other delegated physician extender, if performing the scanning." |  | The text was revised to read “The role of the Sonographer/Technologist may also be taken by the Physician or other qualified medical personnel, such as PA, NP or other delegated physician extender, if performing the scanning." | Done
6 | David T. Fetzer, MD | david.fetzer@utsouthwestern.edu | A.9.2 | 910-915 | Low/Medium | What artifacts? | Specify what artifacts we should be looking for. Any specific to VBF? Or typical US artifacts? | Proposed response: Text of A.9.2 Specification table now reads “Shall confirm the absence of color flow artifacts (e.g. motion artifacts such as flash, smear or blur) that could affect the volume acquisitions”. What do we add? In addition, at the Profile TG discussion it was noted that aliasing might not be considered an artifact but rather an acquisition issue. Consider moving the following text to “Acquisition Protocol” Shall monitor or review the acquired images to ensure that velocities in the target vessel remain(ed) within the color flow scale range (no aliasing) during the 3D sweep, unless otherwise instructed by manufacturer guidelines. In consultation with others, aliasing has been left in the category of artifact as this is commonly referenced as such in textbooks, e.g. Diagnostic Ultrasound by Rumack and Levine. Also wording was changed elsewhere to clarify artifacts as color flow artifacts, excluding blur that was considered confusing. |  | Text of A.9.2 Specification table now reads “Shall confirm the absence of color flow artifacts (e.g. motion artifacts such as flash or smear) that could affect the volume acquisitions”. In addition, we clarified that the following. "Shall confirm the absence of flow velocity aliasing, or that the aliasing is within the manufacturer-specified level of acceptable aliasing for the volume flow application." | Done
7 | David T. Fetzer, MD | david.fetzer@utsouthwestern.edu | 4.3.1 | 505 | Medium | "... each transducer on a system model used in the clinic..." Keep in mind that if there is a large clinic or large healthcare system, the number of transducer and system (and sonographer) combinations could be insurmountable. | Consider a representative transducer, vendor model, and trained sonographer... if transducers and systems are otherwise passing routine QA, then a representative "system" combination should be suficient. | As this was a low priority comment the response was composed by the editor in consultation with the Profile Task Group as follows with the text in red indicating new text. This was indeed a misunderstanding, and the intent was a representative transducer of each type used on a representative model scanner. Also, this section (4.3.1) is for a clinical trainer so it is not meant to be repeated for every sonographer but rather just for a sonographer who would be training other sonographers. Wording change to “For a representative of each type of transducer capable of 3D volume flow on a given system/model used in the clinic,…”. This should be the wording at Lines 494-495 and 502-503 such that these match. BC reviewed this proposed response and it was accepted. |  | Wording has been changed to “For a representative of each type of transducer capable of 3D volume flow on a given system/model used in the clinic,…”. This is the New wording at the location of Lines 494-495 and 502-503 of the public comment version such that these will match. | Done
8 | David T. Fetzer, MD | david.fetzer@utsouthwestern.edu | 4.4.1 | 521 | Low | "should" or "shall"? |  | As this was a low priority comment the response was composed by the editor in consultation with the Profile Task Group as follows with the text in red indicating new text. Text now reads “Measurements shall be performed according to the guidance of 4.0.1 in a calibrated phantom with the additional following considerations:” Similarly this change has been made in section 4.4.2. This comment on "should" or "shall" resulted in some review of the terms and modifications elsewhere in the Profile for consistency. |  | Text now reads “Measurements shall be performed according to the guidance of 4.0.1 in a calibrated phantom with the additional following considerations:” Similarly this change has been made in section 4.4.2. Other instances of "should" were changed to "shall" for consistency. | Done
9 | David T. Fetzer, MD | david.fetzer@utsouthwestern.edu | A.8.1 | 832 | Low/Medium | Are we not assuming that patient-related factors such as fasting/fluid status, positioning, and breathhold/Valsalva have no impact on these measurements? | May be good to specify as an Assumption. | Needed to Review Section A.8. Solicited comments from several physicians involved in the VBF profile development. Comments were received and Brian F. drafted a modification to reflect these comments. |  | Text was modified (changes in red) to state "For the two applications discussed herein, no special issues related to subject handling exist that are known to uniquely affect VBF as defined in this profile. However, there are clearly physiologic conditions and other factors that affect blood flow and interpretation of a measurement of any parameter depends on the circumstances under which the measurement is made. For VBF, one can expect changes in blood flow for different physiological and pathological states such as resting and active, medication or other factors affecting heart rate and blood pressure, dehydration, Valsalva maneuver, patient’s position (upright vs prone). For dialysis patients, the VBF measurement may be affected by the timing with respect to their dialysis schedule although no specific examination of this has been made. In all cases, a standard protocol may help to minimize the variation in data acquisition for estimating volume blood flow. | Done
10 | David T. Fetzer, MD | david.fetzer@utsouthwestern.edu | A.13.1 | 978 | Low | "echo images" | Do you mean grayscale / b-mode images? | As this was a low priority comment the response was composed by the editor as follows with the text in red indicating new text. Yes. Text changed to read “color and grayscale portion of images are of good quality and depict lumen boundary.” |  | Text changed to read “color and grayscale portion of images are of good quality and depict lumen boundary.” | Done
11 | David T. Fetzer, MD | david.fetzer@utsouthwestern.edu | A.13.1 | 979 | Low/Medium | What artifacts? |  | Currently reads “absence of velocity aliasing and other artifacts in target vessel and immediately surrounding background.” Do we need some more details? Profile text has been updated to clarify that these are color flow artifacts that need to be avoided. This is also related to comment 6 above where additional corrections were made to the text. |  | See corrections made in response to Comment 6. | Done
12 | Ioannis Sechopoulos | ioannis.sechopoulos@radboudumc.nl | A.9.1 | 904 | Low | on page 32, line 904, “…the recommendation is still to acquire with as small a Doppler angle as possible to avoid reduced velocity and power sensitivity, and any wall filter effects on low velocity flow.” The “avoid reduced velocity” popped out to me initially, and I thought this sentence was saying small Doppler angle could avoid reduced velocity which would not be correct. But I re-read it and guess the authors just wanted to say reduced sensitivity. The sentence could be rephrased to improve clarify. | The sentence could be rephrased to improve clarify. | The text originally read “…the recommendation is still to acquire with as small a Doppler angle as possible to avoid reduced velocity and power sensitivity, and any wall filter effects on low velocity flow.” Text should be changed to read "...the recommendation is still to acquire with as small a Doppler angle as possible to avoid reduced sensitivity in both velocity and power, and any wall filter effects on low velocity flow. |  |  | Done
13 |  |  | All | All | Medium | This technology is not commercially implemented yet (or at least as far as I know), so there is not much data or wide experience out there. | The quality index or some kind of feedback is absolutely a good idea, but there is no detail on this in the profile at all – I guess the authors purposefully left it that way, which may be ok at this stage. | Proposed response: Both of these statements are true. This technology is not commercially available at the time of this writing. However, there are companies exploring implementation at different levels and this profile could be applied when commercial release occurs and even before in testing before release. Secondly, specifics on a quality index would be a development by the manufacturer and could either be standardized across platforms or proprietary to a given manufacturer. In either case, this could be used to aid in acquiring image data that would allow measurements of volume blood flow that reach the profile claims. This response was reviewed by the BC and accepted. No modification to the profile was considered necessary. |  | Final Response: Both of these statements are true. This technology is not commercially available at the time of this writing. However, there are companies exploring implementation at different levels and this profile could be applied when commercial release occurs and even before in testing before release. Secondly, specifics on a quality index would be a development by the manufacturer and could either be standardized across platforms or proprietary to a given manufacturer. In either case, this could be used to aid in acquiring image data that would allow measurements of volume blood flow that reach the profile claims. NO MODIFICATION TO THE PROFILE. | Done
14 |  |  | All | All | Low | “As a participant in this and numerous other profile creation effort, I judge this one is done very well for the stage at which this technique is implemented commercially. As more experience is gained with commercial systems in clinical or clinical trials settings, it may be possible to refine, even simplify, the requirements for testing and training.“ | None | Proposed response: This point is well taken and appreciated. It is true that substantial experience will likely be gained once there is an implementation on the scanner such that there is no longer a requirement to process image data offline. This will make testing at sites easier, and more sites would be available to assess performance. It was the intent to position this profile such that it could be applied in the development of 3D VBF so that claims could be reached across platforms from the beginning. |  | This point is well taken and appreciated. It is true that substantial experience will likely be gained once there is an implementation on the scanner such that there is no longer a requirement to process image data offline. This will make testing at sites easier, and more sites would be available to assess performance. It was the intent to position this profile such that it could be applied in the development of 3D VBF so that claims could be reached across platforms from the beginning. No modication to the Profile was requested. | OK