Table of Contents

1. Introduction	2

1.1. Purpose & Scope	3

2. Endpoints and Investigations	3

2.1. Primary and secondary endpoints	3

2.2. Primary Investigations	3

2.3. Secondary investigations (future)	4

3. Study Design	4

3.1. Data	4

3.2. Seeding and Truthing	5

3.3. Statistics	5

3.3.1. Characterizing Performance of Absolute Volume Estimation where Ground Truth is Known	5

3.3.2. Characterizing Performance of Change Estimation in the Absence of Biological Change	7

4. Implementation of the challenge studies	9

4.1. Flow of event for each study	9

4.2. Results	9

5. Definitions	10

6. References	12

# Introduction

X-ray computed tomography (CT) is often an effective imaging technique for assessing therapy response. In clinical practice, qualitative impressions based on nothing more than visual inspection of the images are frequently sufficient for making some clinical management decisions. Quantification becomes helpful when tumor masses change slowly over the course of illness. Many investigators have suggested that quantifying whole tumor volumes could solve many of the limitations of RECIST’s current dependence of uni-dimensional diameters on axial slices, and have a major impact on patient management., A few studies have shown that volumetry has value. Some reports about the precision,, and accuracy of measurement have led to concerns about the risks of confusing variability with medically meaningful changes.

QIBA has constructed a systematic "process map" to qualifying volumetry as a biomarker of response to treatments for a variety of medical conditions, including lung disease. Several trials are now underway to provide a head-to-head comparison between volumetry and RECIST in multi-site, multi-scanner-vendor settings. The QIBA Profile is expected to provide specifications that may be adopted by users as well as equipment developers to meet targeted levels of accuracy and clinical performance in identified settings, both as a correlation to clinical outcomes as well as a comparison to the accepted measure of uni-dimensional diameters.

One approach to encouraging innovation that has proven productive in many fields is for an organization to announce and administer a public “challenge” whereby a problem statement is given and solutions are solicited from interested parties that “compete” for how well they address the problem statement.  The development of image processing algorithms has benefitted from this approach with many organized activities from a number of groups.  Some of these groups are organized by industry (e.g., Medical Image Computing and Computer Assisted Intervention or MICCAI), academia (e.g., at Cornell University), or government agencies (e.g., NIST).  This workflow is intended to support such challenges.

It is important to note that one of the reasons for doing this 3A study is to meet the need that a biomarker is defined in part by the “class” of tests available for it.  That is, it is not defined by a single test or candidate implementation but rather by an aggregated understanding of the results of such tests.  As such, it is necessary through this or other means to organize activities to determine how the class performs, irrespective of any candidate that purports to be a member of the class.  The corresponding workflow is related to the “Compliance / Proficiency Testing of Candidate Implementations” workflow and it may be that an organization such as NIST can both host challenges as well as serve in the trusted broker role using common infrastructure for these separate but related functions.

In summary, 3A is motivated by the following:

- Changes in malignant nodule volume is important for diagnosis, therapy planning, therapy response evaluation

- Measuring volume changes requires high accuracy in measurement of absolute volume

- Volumes of synthetic nodules may be measured with high accuracy

- Therefore it make sense to use such phantom data (as ground truth) in order to calculate accuracy measurement of algorithms

- The study results could be combined with the QIBA 1A and 1B Group work. This combination will improve the QIBA volumetric CT Profile development.

We will proceed to add reference clinical data sets, e.g., from Volcano, LIDC and other studies, moving forward wherein at least volume change can be measured by these algorithms, even if ground truth is not available (e.g., no pathologic specimen to compare to)

## Purpose & Scope

The primary and first aim of the study is to estimate inter- and intra-algorithm variability by the volume estimation of synthetic nodules from CT scans of an anthropomorphic phantom (according to the work of the QIBA 1A Group. An inter-algorithm study, in the same way QIBA has been working on inter-reader, inter-scanner, and inter-site.  We will also connect the output of this study to the analysis section of QIBA Profile.  The aim of the study is not a measure of which of several algorithms provides the best image analysis.  Rather, the aim of the study is to gain knowledge to improve QIBA Volumetric CT Profiles and to provide context in which multiple parties have incentives to participate, while avoiding competition and supporting cooperation with a conjoint approach.

Participants include academic and commercial algorithm developers.  Industrial vendors may include, for example possible vendors, according to the Volcano 2009 challenge could be: Siemens, Philips, MeVis, Kitware, Definiens, Intio, VIA CAD etc…)

Scope of the study includes the following types of approaches:

- An automatic segmentation algorithm does not require any user intervention (include detection).

- A semi- automatic algorithm needs minimal amount of input from user, e.g., a seed point to initialize the segmentation, then

  - user allowed to edit

  - user does not edit

Description/Classification of the algorithms: according to the grade of user intervention is needed (for example Volcano’09, A. P. Reeves et al):

- Totally automatic using seed points (no editing beyond setting initial seed)

- Limited parameter adjustment (on less than 15% of the cases)

- Moderate parameter adjustment (on less than 50% of the cases)

- Extensive parameter adjustment (more than 50% of the cases)

- Limited image/boundary modification (on less than 15% of the cases)

- Moderate image/boundary modification (on less than 50% of the cases)

- Extensive image/boundary modification (more than 15% of the cases)

# Endpoints and Investigations

## Primary and secondary endpoints

The primary endpoint depends on the study data utilized.  In such cases where there is a single time point per case, the primary endpoint is volume in mm3.  In such cases where there are two timepoints, the primary endpoint is percent change in volume from the first time point.

There are no secondary endpoints.

## Primary Investigations

Study objectives (in priority and sequence order):

- Characterize Performance of Absolute Volume Estimation where Ground Truth is Known: Results on phantom data, e.g., accuracy and variability, on scans of an anthropomorphic phantom (according to the work of the QIBA 1A Group. (see Dr. Petrick‘s paper, SPIE 2011).  Results on these data utilize the absolute volume primary endpoint.

- After an initial pilot study (1 above) we will proceed to Characterize Performance of Change Estimation in the Absence of Biological Change for results on clinical data, e.g., minimum detectable change and reproducibility.  The study will utilize the same data set as used in the 1B study.  Results on these data utilize the percent change in volume primary endpoint.

The primary investigations are accomplished using participant-supervised reads (they are able to train and otherwise prepare readers for the optimal use of the algorithm, as well as have full access to the study data sets and results).

The QIBA Profile is used to establish targeted levels of performance with means for formal data selection that allows a batch process to be run on data test by a trusted broker that is requested by commercial entities that wish to obtain a certificate of compliance (formal) or simply an assessment of proficiency as measured with respect to the Profile.

## Secondary investigations (future)

Extended study for additional measures of algorithm effectiveness:

- MICCAI-like objective criteria of algorithm performance

- Usability or workflow-effectiveness evaluations (e.g., how many corrections, how fast the algorithm ran (time), etc.

Extended study to explore trusted broker scenarios:

- Sequester data and have a black-box wrapped full automated algorithm produce results;

- Support impartial readers to use participant-defined interfaces to their algorithms (allowing whatever training they recommend beforehand but otherwise not have access to the raw imagery prior to the test).

These secondary investigations are expected to utilize the informatics infrastructure but are otherwise not being actively developed at this time.

Extended study to characterize performance of change estimation in the presence of biological change is generally understood to be out of scope for 3A at this time, covered in QIBA’s 3B effort.  This does not preclude use of the infrastructure for such studies, however, and in fact it is understood as desirable for this to be supported in the future.

# Study Design

## Data

Reference Data Sets will be established and made available to participants.  In the primary investigations, these data comprise the set on which the test is to be run, with the stipulation that these data NOT be used by participants for algorithm training or optimization.

Extended studies to explore trusted broker scenarios may subsequently be pursued wherein so-called “training” or “development” sets may be made available for the convenience of participants, but the definitive test is run on sequestered data not visible to the participants.  These extended studies are out of scope presently but aspects of the investigational infrastructure are intended to be used for them.

## Seeding and Truthing

Reference Data Sets will be accompanied by seed points defined in the context of an indexing scheme.  The purpose of this is so as to achieve consistency across participants such that each lesion is identified in the same way.  The first step is to constrain participants to the identified seed point, extended studies may be subsequently pursued wherein participants may be allowed to select different seed point, provided they utilize the lesion identification scheme provided.

Truth is known in both of the primary investigations.  Phantoms are utilized for the absolute volume estimation study and external measurements are made.  The volume change study is performed on scans separated by approximately 15 minutes so as to establish that there is no actual volume change, hence truth is known to be zero.  In practice (and hence the point of the exercise) the measurements made will not be zero hence the interest of the study.

## Statistics

Statistical measures calculated in these studies include Uncertainty (specifically Bias) and Variability (specifically Variance).  Interestingly, we don’t include

Accuracy among these, as the term is formally understood to be a qualitiative one that reflects whether a result is “right” or “wrong” rather than a numeric quantity.  As such, we do estimate Bias, but reserve Accuracy for the clinical interpretation of these measurements rather than applying at the level of the measurements themselves.

Likewise, note that Precision and Reliability (comprised of Repeatability and Reproducibility) is not assessed in these studies.  The reason for this is we presently don’t include repeated measurements and/or use of alternative seed points.  It is possible to pursue extended studies that estimate these performance characteristics as well but the present scope precludes them.

### Characterizing Performance of Absolute Volume Estimation where Ground Truth is Known

This part of the study utilizes a reference data set comprised of the same cases as used in the 1A study, and extends the statistical analysis that was conducted for that study but in the standard terms as defined above.  Specifically, the following parameters are assessed:

- Uncertainty

  - Bias: mean of measured volume minus the physical measurement of the anthropomorphic phantom object.  Expressed as percent of actual.

where    is the percent difference in volume (i.e. (measured –phantom size) /phantom size*100) in ith phantom and measured by jth algorithm,    is the mean of the percent difference across phantoms and algorithms, N (= nk) is the number of observation in the sample set.

- Variability

  - Variance: estimate overall variance in the difference of measured volume from known physical measure or in the difference of two calculated measured volumes in the same tumor in two images (e.g. different factor levels; slice thickness).

where    is the percent difference (i.e. (measured –phantom size) /phantom size*100) in ith phantom and measured by jth algorithm,   is the mean of the percent difference across phantoms and algorithms,    is the mean of relative bias across algorithms, N (= nk) is the number of observation, n is the number of phantom, and k is the number of algorithm in the sample set.

The above is assessed at two levels. First, the group of tests that collectively comprise the so-called acceptable assay methods for the biomarker. Second, the performance of individual test, in terms of how the individual results compare with the dispersion evident in the group.

- Perform the methods on the reference data:

  - Analyze statistical variability across the following factors: 1. algorithm type, 2. Image formation factors (exposure, pitch, collimation, slice thickness, reconstruction kernel), 3. Anthropomorphic features (size, morphology, density, and attachment).

    - Overall: estimate bias and variance using mean, SD, box-plot (as a more flexible representation than BA) in the difference of measured volume from the physical volume of phantom

    - Similar analysis for each factor

    - (group discussion needed on whether degree of automation is in scope)

  - Additionally, perform  ANOVA or regression analysis to test the variability among algorithms and degree of automations

  - Identify  outliers whose bias are greater than 30% and report a summary in characteristics of tumor

- Assess the performance of each descriptive statistic (in our case, Bias and Variance but more generally to include others such as Accuracy, Precision, Repeatability and Reproducibility) and describe them in a box plot similar to the following example:

Figure 1: Box plots showing dispersion of participant results for each of the descriptive statisitics selected for the study (Bias and Variance to be used for the first work here, but this example shown extended to include other descriptive statistics also).

- Select a “group value” for each of the descriptive statistics, e.g., as the mean plus 2 std.

- For each participant, report their results back to them in the following form:

Figure 2: Radar plot showing the “group value” and how one of the individuals compares with it (Bias and Variance to be used for the first work here, but this example shown extended to include other descriptive statistics also).

In a later phase, this analysis may be conducted over the set of available FDA acquisitions, e.g., the 1187 cases presently loaded as a larger reference data set.

### Characterizing Performance of Change Estimation in the Absence of Biological Change

This part of the study utilizes a reference data set comprised of the same cases as used in the 1B study, and extends the statistical analysis that was conducted for that study but in the standard terms as defined above.  Specifically: the following parameters are assessed:

- Uncertainty

  - Bias: mean of changes in two measured volumes.  Expressed as percent of actual.

, where    is a volumetric measurement in ith phantom and measured by jth algorithm at the time point 1,   is percent changes in ith phantom and measured by jth algorithm at the time point 2,    is set to zero as the mean in the absence of biological changes, N=nk is the number of observation, n is the number of case, and k is the number of algorithm in the sample set.

- Variability

  - Variance: estimate overall variance in the difference of measured volume from known physical measure or in the difference of two calculated measured volumes in the same tumor in two images.

where    is a volumetric measurement in ith phantom and measured by jth algorithm at the time point 1,   is a percent change in ith phantom and measured by jth algorithm at the time point 2,   is set to zero as the mean in the absence of biological changes,    is the mean of percent difference in across algorithms, N (= nk) is the number of observation, n is the number of case, and k is the number of algorithm in the sample set.

The above is assessed at two levels. First, the group of test that collectively comprises the acceptable assay methods for the biomarker. Second, the performance of individual test, in terms of their membership.

- The following comparisons of markups can be made:

  - Analyze statistical variability across the following factors: 1. Algorithm type, 2. Tumor features (size, morphology, measurable/measurable, attachment, contrast of boundary).

    - Overall: using mean, SD, box-plot using the difference of two time points where there is no biological differences in tumor

    - Similar analysis for each factor

  - Calculate the precision of the changes in estimated volumetric measures in thresholds of 5%, 10%, 15%, 20%, 25%, and 30%.

    - Overall: using mean, SD, graph of the precision (%) vs. the thresholds, water-fall plot

    - Similar analysis for each factor

  - Additionally, perform  ANOVA or regression analysis to test the variability among algorithms and degree of automations

  - Identify  outliers whose precision are greater than 30% and report a summary in characteristics of tumor

Subsequently, follow steps 2-4 as defined above using these descriptive statistics.

# Implementation of the challenge studies

The following outlines the procedure to be taken by participants:

- Submit an application to participate to the trusted registrar (non-competing organization) and sign the Participation Agreement

- Download and read the 3A Challenge Protocol as posted to the 3A Wiki.

- Download the 3A Challenge data as described in the Protocol.  This data will be inclusive of a defined development (e.g., training) set for algorithm adjustment and a test set on which the results would be measured.  Data will include images and one seed point per target lesion defined by a non-participant.

- Once the development set is used by the algorithm to do any parameter tuning, these tuning parameters should be used without further modification on the test set (similar to MICCAI liver challenge in 2008). (Note: in the pilot phase, individual participant integrity is relied on to enforce this policy.)

- Report your results in the required formats, signed by your team leader, to 3A registrar.

- 3A registrar will analyze the reported results as per the Analysis section of this document. 3A registrar will provide Participants with individual analysis of their results. We will publish the results of the evaluation, without publicly identifying individual scores by Participant.

## Flow of events for each challenge study

In this case there are two primary actors: the participant, and the honest broker:

- Individual participant:

  - Algorithms included in the imaging test for data and results interpretation must be pre-specified before the study data is analyzed. Participants will be provided a development set for any algorithm tuning, such development set to be comparable to the test set, but without any repeated use of the same data. Lung data is very different from liver, for example. Alteration of the algorithm to better fit the data is generally not acceptable and may invalidate a study.

  - The individual participant or organization needs to receive back performance data and supporting documentation capable of being incorporated into regulatory filings at its discretion.

- 3A registrar:

  - The honest broker needs means to archive data sets that may be selectively accessed according to specific clinical indications and that may be mapped to image quality standards that have been described as so-called “acceptable”, “target”, and “ideal”

  - It needs to produce documentation regarding results inclusive of a charge-back mechanism to recover operational costs.

- The development set will continued to be available but should be stable whereas    test sets may be refreshed with new cases for direct access by interested investigators for testing of new imaging software algorithms or clinical hypotheses. Future investigators will have access to the development set and test sets for additional studies.

- Define services whereby the test set is indirectly accessible via the trusted broker.

## Results

Each participant has to be informed (only the annonymized group results with an indication of which member they are).  Likewise, the team will produce a publication of the results (to all participants), with authorship representing initial participants.

At this point, it is possible to apply the study infrastructure to new participants as desired, reporting as follows:

# Definitions

- Uncertainty(2)*: A value, associated with the result of a measurement, that characterizes the dispersion of the values that could reasonably be attributed to the measurement, composed of uncertainty from both random and systematic error.  Random error contributes to reliability, whereas systematic error contributes to validity (http://www.ncbi.nlm.nih.gov/pmc/articles/PMC1250265).

  - Bias: A quantitative term describing the difference between the average of measurements made on the same object and its true value. In particular, for a measurement laboratory, bias is the difference (generally unknown) between a laboratory's average value (over time) for a test item and the average that would be achieved by the reference laboratory if it undertook the same measurements on the same test item (http://www.itl.nist.gov/div898/handbook/mpc/section1/mpc113.htm).

- Precision:  Closeness of agreement between indications or measured quantity values obtained by replicate measurements on the same or similar objects under specified conditions (http://www.bipm.org/utils/common/documents/jcgm/JCGM_200_2008.pdf).

- Reliability:  The extent to which an experiment, test, or measuring procedure yields the same results on repeated trials (http://www.ncbi.nlm.nih.gov/pmc/articles/PMC1250265/).

  - Repeatability(2)*: Closeness of the agreement between the results of successive measurements of the same measure and carried out under the same conditions of measurement (http://physics.nist.gov/Pubs/guidelines/appd.1.html).

  - Reproducibility(2)*: Closeness of the agreement between the results of measurements of the same measurand carried out under changed conditions of measurement (http://physics.nist.gov/Pubs/guidelines/appd.1.html).

- Variability: The tendency of the measurement process to produce slightly different measurements on the same test item, where conditions of measurement are either stable or vary over time, temperature, operators, etc. (http://www.itl.nist.gov/div898/handbook/index.htm).

  - Variance: the quantity defined as

where   is the mean of the data,   is number of observations in the sample set. (http://www.itl.nist.gov/div898/handbook/eda/section3/eda356.htm).

  - Bias: see above.

# References

# Footnotes

It should be noted that RECIST has not been shown to reliably achieve an accurate and precise measurement with a 20% SLD measurement.

Deng X and Du G. 3D segmentation in the clinic: A grand challenge II – liver tumor segmentation 2008. http://lts08.bigr.nl

For our purposes, since we are gaining experience with this concept, we arbitrarily consider the list of participants to comprise this group.  As we gain experience with this method and/or replace it with better ones, we will tighten how we define “acceptable methods” in line with the QIBA Profile specifications.

# Comments

Isn’t this section now wrong?  It seems inconsistent with our stated scope if we are not allowing manual editing of detected borders, volume surfaces.

We have been using the term pilot phase, pilot or pivotal here – either is OK, but for consistency I suggest pilot.

Here are extra parenthesis and an extra k at numerator?

May be redundant with the text below « Bias » formula.

Not finisished

define

Perhaps include the PA as an appendix?

define

define

define

define

This should not be allowed.  If 3D editing is needed, that may be a measure of robustness of the semi-automated or automated segmentation algorithm.

# Endnotes

Moertel CG, Hanley JA. The effect of measuring error on the results of therapeutic trials in advanced disease. Disease 1976; 38: 388-394.

Quivey JM, Castro JR, Chen GT, Moss A, Marks WM. Computerized tomography in the quantitative assessment of tumour response. Br J Disease Suppl 1980; 4:30-34.

Munzenrider JE, Pilepich M, Rene-Ferrero JB, Tchakarova I, Carter BL. Use of body scanner in radiotherapy treatment planning. Disease 1977; 40:170-179.

Petrou M, Quint LE, Nan B, Baker LH. Pulmonary nodule volumetric measurement variability as a function of CT slice thickness and nodule morphology. Am J Radiol 2007; 188:306-312.

Bogot NR, Kazerooni EA, Kelly AM, Quint LE, Desjardins B, Nan B. Interobserver and intraobserver variability in the assessment of pulmonary nodule size on CT using film and computer display methods. Acad Radiol 2005; 12:948–956.

Erasmus JJ, Gladish GW, Broemeling L, et al. Interobserver and intraobserver variability in measurement of non-small-cell carcinoma lung lesions: Implications for assessment of tumor response. J Clin Oncol 2003; 21:2574–2582.

Winer-Muram HT, Jennings SG, Meyer CA, et al. Effect of varying CT section width on volumetric measurement of lung tumors and application of compensatory equations. Radiology 2003; 229:184-194.

Buckler AJ, Mozley PD, Schwartz L, et al. Volumetric CT in lung disease: An example for the qualification of imaging as a biomarker. Acad Radiol 2010; 17:107-115.

Radiological Society of North America. http://qibawiki.rsna.org/index.php?title=Main_Page, accessed 07 Sep 2009.

http://www.grand-challenge.org/index.php/Main_Page, accessed 23 December 2010.

http://www.preventcancer.org/uploadedFiles/Education/Conferences,_Workshops,_and_Educational_Programs/Day-2-Friday-1100-AM-Tony-Reeves.pdf, accessed 23 December 2010.

http://www.nist.gov/itl/iad/dmg/biochangechallenge.cfm, accessed 23 December 2010.

Petrick NP, Kim HJ, Clunie D, Borradaile K, Ford R, Zeng R, Gavrieldes MA, McNitt-Gray MF, Fenimore C, Lu J, Zhao B, Buckler AJ. Evaluation of 1D, 2D and 3D nodule size estimation by radiologists for spherical and non-spherical nodules through CT thoracic phantom imaging, SPIE, February 2011.