Protocol: DCEMRI Quantification

Version .16

3 March 2011

0. Executive Summary

This technique offers a robust, reproducible measure of microvascular parameters associated with human cancers based on kinetic modeling of dynamic MRI data sets. The rigor and details surrounding these data are described throughout the text of this document in various sub-sections.

1. Context of the Imaging Protocol within the Clinical Trial

The following sections describe how this imaging protocol interfaces with the rest of the clinical trial.

1.1. Utilities and Endpoints of the Imaging Protocol

In oncology, Phase 1 trials are generally conducted at 1-3 centers with the ability to recruit patients and conduct the complicated clinical study protocols associated with early development studies. Since these centers often do not have expertise in DCE-MRI and more than one center is typically involved, considerable effort is required to ensure consistent, reliable and fit-for-purpose quantitative DCE-MRI results are obtained reliably at all clinical sites over the duration of the trial. When these trials are sponsored by the biopharmaceutical industry, imaging core labs (also known as imaging contract research organizations, iCROs) are contracted to provide that effort. However, their approaches are proprietary and, in the absence of established guidelines, they are likely to differ among imaging core labs. When the trials are not industry-sponsored, they are generally conducted at a single site with considerable expertise in DCE-MRI. However, the drive for innovation all but ensures that there will be significant differences between academic sites. Hence, the guidelines provided in this profile will ensure that not only are the relative changes induced by treatment are informative, but that absolute changes can be compared across these studies.

1.2. Timing of Imaging within the Clinical Trial Calendar

The DCE MRI committee believes that all baseline evaluations should be ideally be within 14 days, but no longer than 30 days prior to the initiation of therapy. Otherwise the resulting functional tumor characterization may not reflect the status of the tumor prior to initiation of therapy. The interval between follow up scans within patients may be determined by current standards for good clinical practice or the rationale driving a clinical trial of a new treatment

| Timing Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

1.3. Management of Pre-enrollment Imaging

The principal investigator or co-investigators at the particular sites will be responsible for reviewing pre-enrollment imaging (e.g. CT or MRI examinations) that have been a component of routine clinical care. These data will serve as the requisite information to choose the target lesion(s) that will be used for DCE analysis upon enrollment However, only image acquisition and processing protocols that conform to, or exceed, the minimum design specifications described in this protocol are sufficient for quantifying tumor vascular parameters with the precision of measurement specified in the profile claims document. In practice, this will often require “baseline” scans to be repeated according to these guidelines when the objective is to quantify longitudinal changes within subjects.

| Enrollment Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

1.4. Management of Protocol Imaging Performed Off-schedule

**Describe the evaluation, handling and usage of imaging performed according to the Procedure below but not within the “on-schedule” timing window described in Section 1.2. (e.g. For what purpose(s) may such imaging be used (for clinical decision-making; for data analysis; for primary endpoints; for secondary endpoints; for continued subject eligibility evaluation; to supplement but not replace on-schedule imaging, etc.)? What characteristics or timing will make the imaging acceptable for the purpose? Is there normalization that should be done to account for the schedule deviation? What is the expected statistical impact of such imaging on data analysis? How should such imaging be recorded, archived, etc.)

| Management Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

1.5. Management of Protocol Imaging Performed Off-specification

**Describe the evaluation, handling and usage of imaging described below but not performed completely according to the specified Procedure. This may include deviations or errors in subject preparation, the acquisition protocol, data reconstruction, analysis, interpretation, and/or adequate recording and archiving of necessary data. (e.g. For what purpose(s) may such imaging be used (for clinical decision-making; for data analysis; for primary endpoints; for secondary endpoints; for continued subject eligibility evaluation; to supplement but not replace on-schedule imaging, etc.)? What characteristics or timing will make the imaging acceptable for the purpose? Is there normalization that should be done to account for the schedule deviation? What is the expected statistical impact of such imaging on data analysis? How should such imaging be recorded, archived, etc.)

| Management Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

1.6. Management of Off-protocol Imaging

**Describe the evaluation, handling and usage of additional imaging not described below. This may include imaging obtained in the course of clinical care or potentially for research purposes unrelated to the clinical trial at the local site. (e.g. For what purpose(s) may such imaging be used (for clinical decision-making; for data analysis; for primary endpoints; for secondary endpoints; for continued subject eligibility evaluation; to supplement but not replace on-schedule imaging, etc.)? What characteristics or timing will make the imaging acceptable for the purpose? Is there normalization that should be done to account for the schedule deviation? What is the expected statistical impact of such imaging on data analysis? How should such imaging be recorded, archived, etc.)

| Management Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

1.7. Subject Selection Criteria Related to Imaging

a. Absolute contraindications to MRI are not within the scope of this document. Suffice it to say that local policies for contraindications for absolute MRI safety should be followed .

b. Patient selection criteria will include and be guided by the Eastern Cooperative Oncology Group (ECOG) status (See Appendix 2) for full description of ECOG performance status). In specific, patients meeting ECOG status >= 2 will not be eligible for participation in the study, because historically, this patient profile has shown poor ability to meet the demands of the examination.

c. The QIBA DCE-MRI committee acknowledges that there are potential and relative contraindications to MRI in patients suffering from claustrophobia. Methods for minimizing this risk are at the discretion of the physician caring for the patient.

d. The QIBA DCE-MRI committee acknowledges that there are potential risks associated with the use of gadolinium-based contrast media. The default recommendations for intravenous contrast that follow assume there are no known contraindications in a particular patient other than the possibility of an allergic reaction to the gadolinium contrast agent. The committee assumes that local standards for good clinical practices (GCP) will be substituted for the default in cases where there are known risks.

e. Recent FDA guidelines (http://www.fda.gov/Drugs/DrugSafety/ucm223966.htm#aprooved), outline the safety concerns associated with using gadolinium based contrast agents in patients with impaired renal function. The DCE-MRI committee echoes these recommendations and advises reference to these standards when choosing patients in order to determine eligibility for entry into a DCE-MRI clinical trial.

f. Patients who have received an MRI with an extracellular Gadolinium based contrast agent should be ineligible for DCE-MRI trial until 24 hours have expired.

1.7.1. Relative Contraindications and Remediations

**Describe criteria that may require modification of the imaging protocol. (e.g. subjects with kidney insufficiency are contraindicated for Contrast CT in this protocol, at the physicians discretion, kidney function may be re-evaluated prior to imaging to see if the insufficiency has resolved, or the subject may be evaluated for dialysis, etc.)

| Contraindication | Remediations |
| --- | --- |
|  | Acceptable Target Ideal |

1.7.2. Absolute Contraindications and Alternatives

**Describe criteria that may fully disqualify the subject for the imaging protocol. If possible, identify possible alternative imaging protocols. (e.g. subjects with pacemakers are disqualified for this MRI protocol. Consider using CT protocol UPICT-31254 instead) These alternatives may also be useful for relative contraindications if remediations described in 1.7.1 are not possible or successful.

| Contraindication | Alternatives |
| --- | --- |
|  | Acceptable Target Ideal |

1.7.3. Imaging-specific Inclusion Criteria

**Describe imaging-specific inclusion criteria.

| Criteria | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

2. Site Selection, Qualification and Training

Typically clinical sites are selected due to their competence in oncology and access to a sufficiently large patient population under consideration. For DCE-MRI use as quantitative imaging biomarker it is essential to put some effort into an imaging capability assessment prior to final site selection for a specific trial. For imaging it is important to consider the availability of:

• appropriate imaging equipment and quality control processes,

• appropriate injector equipment and contrast media,

• experienced MR technologists for the imaging procedure, and

• processes that assure imaging protocol compliant image generation at the correct point in time.

2.1. Personnel Qualifications

The following qualifications are given for participants in the trial.

2.1.1. Technical

Technicians participating in the trial need the following training.

| Technical Expertise | Qualifications |
| --- | --- |
|  | Acceptable Target Ideal |

2.1.2. Physics

Medical physicists participating in the trial need the following training.

| Physics Expertise | Qualifications |
| --- | --- |
|  | Acceptable Target Ideal |

2.1.3. Physician

Physicians participating in the trial need the following training.

| Medical Expertise | Qualifications |
| --- | --- |
| Attending | Acceptable Target Ideal |
| Reading | Acceptable Target Ideal |

2.2. Imaging Equipment

1.5 T MR machines with 55-70 cm bores need to be available. The scanner needs to be under quality assurance and quality control processes (including preventive maintenance schedules) appropriate for quantitative MR imaging applications, which may exceed the standard requirements for routine clinical imaging or for MR facility accreditation purposes. The scanner software version should be identified and tracked across time. It might be beneficial to identify and qualify a second scanner at the site, if available. If this is done prior to the study start there will be no difficulties later on in case the first scanner is temporarily unavailable.

Injector qualification: A power injector is required for DCE-MRI studies. It needs to be properly serviced and calibrated.

| Device | Compliance Levels |
| --- | --- |
| Agent administration apparatus | Acceptable QIBA compliant device at acceptable level Target QIBA compliant device at target level Ideal QIBA compliant device at ideal level |
| Scanner | Acceptable QIBA compliant device at acceptable level Target QIBA compliant device at target level Ideal QIBA compliant device at ideal level |
| Reconstructor | Acceptable QIBA compliant device at acceptable level Target QIBA compliant device at target level Ideal QIBA compliant device at ideal level |
| Post-processing software | Acceptable QIBA compliant device at acceptable level Target QIBA compliant device at target level Ideal QIBA compliant device at ideal level |
| Analysis software | Acceptable QIBA compliant device at acceptable level Target QIBA compliant device at target level Ideal QIBA compliant device at ideal level |

2.3. Infrastructure

**List required infrastructure, such as subject management capabilities, internet capability, image de-identification and transmission capability.

| Infrastructure Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

2.4. Quality Control

Quality control is described in Section 12.

2.4.1. Procedures

See 12.1.1.

2.4.2. Baseline Metrics Submitted Prior to Subject Accrual

See 12.1.2.

2.4.3. Metrics Submitted Periodically During the Trial

See 12.1.3.

Additional task-specific Quality Control is described in sections below.

2.5. Protocol-specific Training

**Describe any necessary training specific to this protocol in the sections below.

2.5.1. Physician

**See …

2.5.2. Physics

**See …

2.5.3. Technician

**See …

3. Subject Scheduling

**Describe requirements and considerations for the physician when scheduling imaging and other activities, which may include things both related and unrelated to the trial.

3.1. Timing Relative to Index Intervention Activity

The following requirements are placed regarding timing relative to index intervention activity.

| Index Intervention Activity | Timing |
| --- | --- |
|  | Acceptable Target Ideal |

3.2. Timing Relative to Confounding Activities (to minimize “impact”)

**(e.g. Avoid scheduling a biopsy on a tumor within X days prior to the FGD-PET scan to evaluate tumor viability; Avoid scheduling the MRI scan within X hours following administration of TPA (for stroke) to the subject.)

| Confounding Activity | Timing |
| --- | --- |
|  | Acceptable Target Ideal |

3.3. Scheduling Ancillary Testing

**(e.g. order a blood draw to occur within X hours preceding the imaging procedure.)

| Ancillary Testing | Scheduling |
| --- | --- |
|  | Acceptable Target Ideal |

4. Subject Preparation

There are no specific patient preparation procedures for the MRI scans described in this protocol. The DCE-MRI committee acknowledges that there are specifications for other procedures that might be acquired contemporaneously, such as requirements for fasting prior to FDG PET scans or the administration of oral contrast for abdominal CT. Those timing procedures may be followed as indicated without adverse impact on these guidelines

4.1. Prior to Arrival

The local standard of care for acquiring MRI scans may be followed. For example, patients may be advised to wear comfortable clothing, leave jewelry at home, etc.

| Preparation | Compliance Levels |
| --- | --- |
|  | Acceptable Target deal |

4.2. Upon Arrival

Staff shall prepare the patient according to the local standard of care.

1. Patients should be assessed for any removable metal objects on their bodily surfaces that will be in the field of view.

2. Patient should be "comfortably positioned", in "comfortable clothes to minimize patient motion and stress (which might affect the imaging results) and any unnecessary patient discomfort.

4.2.1. Confirmation of Subject Compliance with Instructions

**(e.g. instructions to the admitting nurse/tech to confirm with the subject upon arrival that they have complied with each of the instructions in 4.1.)

| Instruction | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

4.2.2. Ancillary Testing

**(e.g. blood draws, weight/blood pressure measurement, etc. associated with the imaging and downstream actions relative to such testing)

| Testing | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

4.2.3. Preparation for Exam

**(e.g. empty bladder, removal of metal objects, etc.)

| Preparation | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

5. Imaging-related Substance Preparation and Administration

The DCE-MRI committee acknowledges that the use of intravenous contrast material is often medically indicated for the diagnosis and staging of cancer in many clinical settings.

5.1. Substance Description and Purpose

**Describe the substance and its purpose, stating any requirements needed.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

5.2. Dose Calculation and/or Schedule

Contrast Agent Dose Reduction Based On Creatinine Clearance: (renal function)

a. An extracellular gadolinium based contrast agent (e.g. Gd-DTPA) will be utilized.

b. Patient’s renal creatine clearance should be obtained, and estimated glomerular filtration rate (eGFR) determined in adults through well known and adopted formulas.

i. If eGFR < 60 ml/min/1.73 m2, then the subject should withdraw from the study.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

5.3. Timing, Subject Activity Level, and Factors Relevant to Initiation of Image Data Acquisition

a. Contrast injection should occur after the following imaging sequences have been acquired (See Section 6):

i. Anatomic imaging for localizing tumors

ii. Variable flip angle imaging for R1 map calculation

iii. Ratio map images for signal intensity normalization

b. Contrast injection should occur after at least 5 baseline volume imaging stacks have been acquired.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

5.4. Administration Route

a. Each subject should have an intravenous catheter with a gauge no smaller than 20 gauge which should be placed in the right antecubital fossa. Injection through a port-a-catheter, or permanent indwelling catheter is not allowed.

b. Contrast agent should be administered in a dynamic fashion, preferably with a power injector. At baseline and at each subsequent time-point, the same dose of contrast and rate of contrast administration should be performed as clinically safe. The rate of administration should be rapid enough to ensure adequate first-pass bolus arterial concentration of the contrast agent (generally 2-4 mL/sec).

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

5.5. Rate, Delay and Related Parameters / Apparatus

**Describe the rate, delay, and related parameters or apparatus. Place needed requirements.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

5.6. Required Visualization / Monitoring, if any

**Describe the required visualization or monitoring, placing any needed requirements.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

5.7. Quality Control

See 12.2.

6. Individual Subject Imaging-related Quality Control

See 12.3.

7. Imaging Procedure

This section describes the imaging protocols and procedure for conducting a DCE-MRI exam. Suitable localizer (scout) images must be collected at the start of exam, and used to confirm correct coil placement as well as selection of appropriate region to image. This will be followed by routine non contrast sequences, usually in the axial plane, to delineate number, location, and limits of tumor extension. Exact protocols for these imaging sequences may be determined by the local imaging norms. However, for imaging of areas subject to respiratory motion, care should be made to note the presence or respiratory suspension (end-inspiration vs. end- exhalation) or respiratory gated techniques as these maneuvers may displace the apparent location of tumor(s) relative to fixed non-moving anatomic landmarks.

7.1. Required Characteristics of Resulting Data

The DCE-MRI portion of the exam will consist of three components: (a) a ratio map series , for the acquisition of data useful for signal intensity normalization; (b) a variable flip angle series, for pre-contrast T1 mapping; and (c) a DCE-MRI protocol, which collects dynamic data during the passage of the contrast agent.  Detailed specifications for these protocols are as below:

a. Ratio Map Protocol (if ratio map intensity correction will be applied ): All parameters the same as for dynamic protocol, below, except: Single acquisition phase 15 degree flip angle Number of Signal Averages (NSA or NEX): Ideal: 8; Target: 4, Acceptable: 2  First acquire these images with phase array receive coil, then repeat second time with body receive coil. Machine gain settings should not be altered between these two sequences.

b. Variable Flip Angle T1 Mapping Protocol: All parameters the same as dynamic protocol, below, except: Single acquisition phase Ensure TR and TE values stay constant for all flip angles. Flip angles: Ideal: 7 or more flip angles (2, 5, 10, 15, 20, 25, 30 degree) Target: 6 flip angles (2, 10, 15, 20, 25, 30 degree) Acceptable: 5 flip angles (2, 10, 15, 20, 30 degree) Number of signal averages (NSA or NEX): Ideal: 4; Target: 2, Acceptable: 1  Ensure that machine gain settings are not reset automatically (using automated pre-scan features) between each flip angle acquisition so that system gain settings are identical for each flip angle acquisition.

c. DCE-MRI Protocol: Pulse Sequence: 3D fast spoiled gradient recalled echo or equivalent Coils: Body transmit coil, phased array receive coil No parallel imaging options  No magnetization preparation schemes Imaging Plane (thoracic/abdominal/pelvic): Coronal oblique acquisition slab (including appropriate vascular reference vessel(s)) Imaging Plane (brain/H&N/extremity): Axial acquisition slab (including appropriate vascular reference vessel(s)) Frequency encoding direction: S-I (Head-feet) frequency for oblique coronal and L-R frequency for axial. TE as short as possible: Ideal: < 1.5 ms, Target: 1.5-1.99 ms, Acceptable: 2.0-2.5 ms TR as short as possible: Ideal: <3 ms, Target: 3-4.99 ms, Acceptable: 5-7 ms Temporal resolution: Ideal: <5 s, Target: 5-8 s, Acceptable: 8-10 s Flip angle (degree): Ideal: 30-35, Target: 25-30; Acceptable: 20-25 Smaller flip angles will lead to potential saturation of the signal intensity vs. gadolinium concentration, particularly in vessels). Note should be made that SAR limits may affect the maximum allowable flip angle. Operators should use the maximal allowed flip angle when SAR limitations occur.  Receiver Bandwidth: greater or equal to ±31.25 kHz (or ~250 Hz/pixel) Region appropriate frequency encoding (FE) FOV (recommend 42 x 34 cm for body applications; 22-25 cm for brain/H&N ) ~80% phase encoding FOV  Partial Fourier (“fractional echo”) as needed to meet temporal resolution requirements Number of Slices: Acceptable: >=10 prior to zero fill. Ideal: as many as possible while maintaining ideal temporal resolution. Slice thickness: Ideal: <5 mm, Target: 5.1-6 mm, Acceptable: 6.1-8 mm  Matrix: 256 x 160 (before applying rectangular FOV) Number of acquisitions (phases): Sufficient to allow acquisition of at least 5 minutes of post injection data; including at least 5 phases acquired before contrast agent injection (baseline images).

7.1.1. Data Content

**Describe what the acquired images should contain/cover. (e.g. Anatomic Coverage, Field of View)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.1.2. Data Structure

**Describe how the data should be organized/sampled. (e.g. Spatial Resolution, Collimation Width, Slice Interval, Temporal Resolution)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.1.3. Data Quality

**Describe any needed measurements of quality. (e.g. Image Quality, Noise Levels, Motion Artifacts, Radiation Dose)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.2. Imaging Data Acquisition

The following sections describe the acquistion of imaging data.

7.2.1. Subject Positioning

Patient and coil positioning: When the general location of the target tumor(s) is known prior to DCE-MRI, for example glioma or local breast cancer evaluation, the patient set up for the MRI should be based on standard operating procedures for patient positioning and coil placement for clinical MRI of that body part.

When the subject under investigation may have uncertain tumor location(s), as is common in the setting of patients undergoing therapy for metastatic disease, it will often be necessary for the DCE-MRI study to be planned with reference to the most recent pre-DCE-MRI imaging (usually CT study). From this study, tumor burden and location should be assessed. Optimally, review of actual imaging by a radiologist involved in the DCE-MRI study planning should be made. At times, if such images are not available for direct review, review of imaging reports (CT, PET) detailing extent of disease is mandatory, both to confirm eligibility (presence of at least one “imageable” lesion) and to identify the preferred anatomic regions for DCE-MRI (chest, abdomen, pelvis, extremity). Review of prior diagnostic imaging may also be helpful to confirm cyctic or necrotic nature of certain lesions, assessments which may be challenging at the time of DCE-MRI planning based solely on T1 and T2 image sets. When multiple potential target lesions are available, the location of the most suitable lesion(s) should be noted. The most suitable lesion will depend on size, location relative to areas of pulsatile or respiratory artfacts, and presence or absence of necrosis or cystic. The schema for choosing the target lesion is provided below

DCE-MRI subject should be placed supine in the scanner position. Rarely, prone positioning can be used for patients who profess a preference was such positioning to ease their concerns with claustrophobia. Preferably, oral anxiolytics can be used for such patients prior to each DCE-MRI study.

When patient condition allows, placement of arms over the head may avoid undesirable wrap artifact for temporally optimized 3D spoiled gradient echo sequences for chest and abdomen lesions. However, such positioning often cannot be sustained by patients without excessive discomfort. In such cases, arms placed anteriorly over the chest or at the sides may be preferable. For larger patients, side-down arm positioning may require adjustment of the DCE-MRI imaging FOV to avoid undesirable wrap artifact.  Appropriate coil placement per area of examination (head, neck, breast, extremity) is then done. For lesions in the chest, abdomen, or pelvis, a torso array coil is then placed in the area of target lesion(s). Ideally, both anterior and posterior coils are centered over the expected target lesion location.

| Parameter | Compliance Levels |
| --- | --- |
| Size | Acceptable 2-3cm Target >6cm or 3-4 cm Ideal 4-6cm |
| Composition | Acceptable Predominantly necrotic or cyctic with at least 1 cm LD of solid enhancing/viable tumor. (Unacceptable: completely cyctic or necrotic lesions. Lesions with extensive hemorrhage) Target Predominantly solid, with some cystic change and/or necrosis Ideal Predominantly solid, no cyctic characteristics, little if any necrosis |
| Location | Acceptable Lesion in area without undo respiratory or pulsatile motion (liver-excluding lat segment left lobe, spleen, lung bases, lateral lungs, upper mediastiunum) (Unacceptable: hilar, para-cardiac medistiumun, lateral segment left lobe of liver.) Target Areas with some respiratory motion (kidneys, adrenal glands/lateral retroperitoneum, lateral chest wall, pancreas, lung apices, neck) Ideal Non-mobile locations (brain, skull, extremity, spine, bony pelvis, deep pelvis soft tissues, central retroperitoneum, posterior chest wall) |

7.2.2. Instructions to Subject During Acquisition

**(e.g., breath hold, etc.)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.2.3. Timing/Triggers

**(e.g., relative to administration of imaging agents; inter-time point standardization)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.2.4. Model-Specific Parameters

**Appendix G.1 lists acquisition parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

7.2.5. Archival Requirements for Primary Source Imaging Data

**See 11.3.

7.3. Imaging Data Reconstruction

**Describe the data reconstruction process including any inherent data correction, smoothing, etc. This may conceivably be performed on a different system than acquisition, and occasionally at a later time, particular for supplementary reconstructions.

7.3.1. Input Data to Be Used

**Place requirements on the input data to be used.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.3.2. Methods to Be Used

**Place requirements on the methods to be used.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.3.3. Required Characteristics of Resulting Data

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

7.3.4. Platform-specific Instructions

Appendix G.2 lists reconstruction parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

7.3.5. Archival Requirements for Reconstructed Imaging Data

See 11.4.

7.3.6. Quality Control

See 12.4.

8. Image Post-processing

There are no specific image post-processing requirements in this profile. No user-selected post-processing filters or image normalization methods should be used prior to data analysis as described in the next steps. If phased-array receiver coils are used, image combining and reconstruction should be according to standard manufacturer algorithms. DCE-MRI data will be signal normalized and converted into estimates of Gd concentration as described later in Section 9.

8.1. Input Data to Be Used

**Describe required input data and any necessary validation or adjustments which should be performed on it. (e.g. particular image series or views; raw, processed, both)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

8.2. Methods to Be Used

**Describe how the analysis should be performed. (e.g. algorithms to be used; where measurements should be taken; definition of key anatomical points or pathology boundaries; related annotations)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

8.3. Required Characteristics of Resulting Data

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

8.4. Platform-specific Instructions

**Appendix G.3 lists parameter values and/or instructions for specific models/versions that can be expected to produce data meeting the requirements of Section 8.3.

8.5. Archival Requirements

See 11.5.

8.6. Quality Control

See 12.5.

9. Image Analysis

Analysis of DCE-MRI data is carried out in a series of distinct steps:

• Generate a ratio image using Body and Phased Array coil images.

• Generate a T1 map using Variable Flip Angle Images.

• When required, apply time-series motion correction to the dynamic data.

• Correct signal intensities in the dynamic data using the ratio image.

• Convert SI(t) in the dynamic data to gadolinium concentration ([Gd](t)).

• Calculate an arterial input function.

• Identify the region or regions of interest in the dynamic data.

• Calculate vascular parameters.

Each of these steps is addressed in detail below.

9.1. Input Data to Be Used

**Describe required input data and any necessary validation or adjustments which should be performed on it. (e.g. particular image series or views; raw, processed, both)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

9.2. Methods to Be Used

9.2.1 Generate a Ratio Image (if ratio map intensity correction is to be applied)

The intent of this step is to optimize the information available for calculation of vascular parameters by combining the signal uniformity available using the inherent body coil with the improved SNR available using a torso phased array coil or similar. The required inputs for this process are two images at each slice location - one acquired with the inherent body coil, and another acquired with the selected phased array (PA) coil. These images should be prescribed identically to the dynamic series, and should be acquired immediately prior to the dynamic series, using the protocol as detailed in Section 5. The output is an image in which each pixel contains a scale factor which, when applied to the dynamic data (which will be acquired using the PA coil) will ideally be sufficient to correct any non-uniformity in signal intensity due to the coil. Motion effects, of course, will complicate the use of such ratio map intensity correction techniques and must be appropriately addressed when necessary. Furthermore, other intensity correction algorithms may be useful. Before using any such intensity correction algorithm, however, the robustness of the technique should be carefully evaluated in phantoms and in test scans from patients or, if appropriately approved, volunteers.

If the ratio map intensity correction approach is utilized, the following technique should be used: The ratio image is calculated as follows:

1. Normalize the range of the two data sets IB and IPA. a. Identify the minimum value in each image. b. Subtract the minimum value from each image. c. Identify the maximum value in each image. d. Multiply IB by the ratio maxPA / maxB to yield a new IB image.

2. Calculate the ratio image IR=IB/IPA.

3. Reduce noise due to low signal regions (e.g. air space) and poorly registered regional transitions. a. Set the ratio = 1.0 at all points where the signal value in either IB or IPA is less than 10. b. Process the resulting IR using a median filter with support 21x21.

9.2.2 Generate a T1 Map

The intent of this step is to provide a complete map of pre-contrast T1 values for the imaged slab. These values will then be used in the signal formation model based conversion of changes in signal intensity to gadolinium concentration. The required inputs for this process are two or more images acquired with different flip angles, preferably covering a broad range above and below the Ernst angles for the acquired TR and range of T1 values expected in the human body before contrast administration. These images should be prescribed identically to the dynamic series, and should be acquired immediately prior to the dynamic series. The output of this step is an image of T1 values which can be co-registered to the dynamic series and used in subsequent calculations. The T1 values at each voxel location are calculated as follows [1]:

1. Create a vector x containing the signal intensity at each flip angle divided by the tangent of the flip angle.

2. Create a vector y containing the signal intensity at each flip angle divided by the sine of the flip angle.

3. For the n acquired flip angles create a set of points (x0,y0)… (xn,yn).

4. Fit a line with slope s to the set of points defined in Step 3. 5. T1 = -TR/ln(s).

9.2.3 Apply Motion Correction to the Dynamic Data

The intent of this step is to correct for patient motion that occurs between acquired phases of the dynamic data due to respiration, swallowing, and other involuntary movements. This step is not intended to correct ghosting artifacts that can occur along the phase encoding direction within a particular image due to patient motion during acquisition. These artifacts are more or less intractable unless the motion is regular and easily modeled, and are best addressed by adjusting the phase/frequency encoding scheme to minimize their impact on structures of interest.  In general, simple rigid shift or affine transform based registration methods will not be adequate for this step, due to the fact that the movement in question is typically limited to specific regions within the image – for example, the liver in a coronal scan of the abdomen may move substantially with respiration while the bulk of the body remains relatively motionless. Fully deformable registration methods based on optical flow may provide good results in some cases [2, 3]. However, these methods will frequently fail for the phases immediately surrounding the contrast injection. Semi-automated registration in which a user identifies the target tumor and only information drawn from that region is used to generate phase to phase shifts provides an alternative approach. This allows rigid shift methods using mutual information [4], which tend to be more robust than optical flow methods, to be employed. Finally, registration may be carried out manually.  Data corrupted with motion must be either corrected prior to analysis or discarded for subsequent pharmacokinetic analysis.

9.2.4 Correct Signal Intensities Using the Ratio Image

The intent of this step is to correct for signal intensity variations introduced by non-uniform spatial sensitivity in the phased array coil. This step is carried out by multiplying the signal value at each pixel in each image in the dynamic data by the value in the corresponding pixel in the ratio image calculated in 9.2.1.

9.2.5 Convert SI(t) in the Dynamic Data to [Gd](t)

The intent of this step is to convert the arbitrary signal intensity units in the dynamic data into units of gadolinium concentration. This step should be applied after the regions of interest for analysis have been defined, but prior to the calculation of vascular parameters. Two methods for accomplishing this are defined below.

9.2.5.1 Conversion Using a Signal Formation Model Gadolinium concentration at each image pixel is given by )(eq 1):

Here T10 is the pre-contrast T1 at that pixel, obtained from T1 map from Section 8.2, and Rgd is the relaxivity of Gd (obtained from manufacturer’s specifications).

T1(t) can be derived from the SPGR imaging equation and is given by the following expressions (eqs 2-4): Let

where α is the flip angle, TR is the repetition time, and S(t) and S(0) are the signal intensities at time t and pre-contrast baseline respectively in the DCEMRI sequence (eq 5).  Then,

9.2.5.2 Conversion Using a Look-Up Table

This method requires that a phantom containing a range of concentrations of gadolinium and a range of baseline T1 values (generally obtained via different concentrations of copper sulfate or a similar compound) is scanned using the dynamic protocol on each scanner that will be used for the study. Data from these phantoms can then be used to construct a look-up table relating baseline T1, signal delta, and gadolinium concentration. This look-up table can then be used during the analysis process to directly convert signal delta to gadolinium concentration prior to parameter calculation.

9.2.5.3 Conversion Using a Fixed T1 values

In this method, the equations in 8.5.1 are used, but instead of using the pre-contrast T10 from the T1 map, fixed known tissue T10 values may be used. This helps make the computation less susceptible to B1 inhomogeneity errors in T1 mapping.

9.2.6 Calculate an Arterial Input Function

The intent of this step is to generate an accurate, patient-specific arterial input function to serve as an input to the vascular model. One way to accomplish this is to have an analyst draw a manual ROI within an artery, and use the mean enhancement curve within that ROI as the subject-specific AIF, as described by Vonken et al [5]. It has been demonstrated previously that this method has significant variability associated with it [6], due primarily to the spatially- and temporally-varying flow artifacts found in major arteries. A better option is to make use of an automated search technique to generate a locally optimal AIF. Several methods of accomplishing this have been described previously [6,7].

9.2.7 Identify Regions of Interest in the Dynamic Data

The intent of this step is to define the boundaries of the tumors of interest in the dynamic data in order to determine which voxels will contribute to the parameter set for each tumor. This step will generally be carried out by a trained and qualified radiologist using either manual tracing or a semi-automated boundary identification tool [6].

9.2.8 Calculate the Vascular Parameters  The intent of this step is to generate the parameter set which will be used to characterize the tissues of interest. Parameters will be calculated based on the standard Tofts model [8], which is derived from the Kety equations [9]. The vascular bed is modeled as a linear system, such that (eq 6):

with impulse response h(t) given by (eq 7):

where Ktrans is the volume rate constant between blood plasma and extra-cellular extra-vascular space (EES) and kep is the rate constant between the EES and blood plasma. Given the tissue uptake curve Ct(t) and the AIF Cp(t) , Ktrans and kep are estimated using a gradient-descent energy minimization scheme. Delay correction should be performed to shift the AIF curve to match the arrival time of the tumor curve prior to curve fitting.

A full parameter set will be calculated for each voxel within the defined tumor boundaries. Parameters may be reported out either as mean and median values per tumor or as histograms.

The blood normalized IAUGCBN is measured from the area under the concentration curve up to 90 seconds post injection, normalized by dividing the area under the arterial input function curve also up to 90 seconds post injection.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

9.3. Required Characteristics of Resulting Data

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

9.4. Platform-specific Instructions

Appendix G.4 lists parameter values and/or instructions for specific models/versions that can be expected to produce data meeting the requirements of Section 9.3.

9.5. Archival Requirements

See 11.6

9.6. Quality Control

See 12.6.

10. Image Interpretation

Derivation of quantitative parameters characterizing the response associated with a lesion of interest from parameter maps is a multistep process, most if not all of which are being studied by on-going research. There are several choices that can be made at any of these steps, and the effect of these choices on the validity of results and variability of parametric maps has not yet been fully characterized.

10.1. Input Data to Be Used

**Describe required input data and any necessary validation or adjustments which should be performed on it. May also specify data which should not be used until after the clinical trial interpretation is recorded. (e.g. particular image series or views; before and after processing versions of images to evaluate/validate the effects of processing; analysis results)

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

10.2. Methods to Be Used

10.2.1 Elimination of end slices.

i. One shortcoming of the 3D fast spoiled gradient recalled echo technique used to acquire the dynamic images is that initial and end slice locations give inaccurate results due to wraparound artifact and variability in excitation profile. Image analysis can begin by discarding the data on these end slices . One method that can be used to determine which slices to discard is to closely examine the T1 maps obtained at the initial and end slice locations. Marked non-physiologic overestimations of T1 on initial and end slices are indicative of artifact.

10.2.2 Tumor ROI Definition.

The first step in the extraction of quantitative parameters (Ktrans or IAUGCBN) associated with a particular neoplastic or other lesion is to segment this lesion from adjacent tissues. Which techniques of segmentation are ideal or even acceptable for a given application is the subject of on-going research, but it is clear that the segmentation techniques used must be tailored to the particular organ system being studied with DCE-MRI. The following guidelines are proposed:

i. We do not recommend an analysis scheme where an operator defines a lesion by placing regions of interest directly on parameter maps as that will introduce bias into the results

ii. Less subjective results can be obtained by using correlative imaging to define the lesion. These correlative images may be obtained at the same imaging session but not directly related to the DCE-MRI images (for example, a T2-weighted image of an organ which clearly delineates lesions and their boundaries). Correlative images should be obtained in the same imaging plane as the DCE-MRI series, with simlar FOV and matrices if feasible. In this scenario, a registration step will likely be required (see step x below).

iii. An alternative approach which may be helpful if the lesion is well delineated on contrast-enhanced T1-weighted MRI is to create summation images (images obtained by adding together images obtained on the dynamic series for each slice location). The summation images can be used to segment the lesion on one or more slices, and because these segmentations are (in the absence of patient or organ motion) registered to the dynamic series, the segmentations can be used to directly extract lesion-based parameters from parameter maps .

iv. Because of the presence of image noise on source images of the dynamic series, along with time-dependent changes in signal intensity which may blur or even obliterate the border between lesion and background tissue, analysis schemes in which lesions are segmented independently on each image of the dynamic series should be avoided where possible. In the case of moving organs, it may be necessary to segment the lesion of interest on early (preferably, before arrival of the contrast bolus) or late dynamic images and estimate the position of the segmented lesion in intermediate time points.

v. Tumor size and location on longitudinal studies should be considered in the design of an analysis scheme. Recall that the claims of this profile are only applicable to lesions 2cm or more in size. If the lesion is large in proportion to the volume imaged by DCE-MRI precautions should be taken to maximize the possibility that the same portion of the lesion will imaged on longitudinal studies. In general, this requires careful scan location set up on follow-up studies in order to match the same anatomic positions imaged in target organs on earlier studies. However, because of differences in patient angulation on follow-up studies the same anatomic locations may not be imaged on each study. In this case, an analysis scheme that discards image data from locations that are not included in the imaged volume (after end slice elimination) of all relevant studies is favored. This can be accomplished by registration of images obtained from the dynamic sequences of all studies (for example summation images) to a high resolution anatomic images obtained (for example) at the initial time point.

Although lesions can be segmented using manual techniques, several techniques are available that allow a semi-automated approach to be used. These techniques include level-set, watershed, and fuzzy k-means clustering techniques, and are beyond the scope of this document. The training of operator or operators in performing segmentations should be documented, preferably with training sets.

10.2.3 Registration of segmentations and parameter maps.

Unless the segmentations are derived from relatively motion-free or motion-corrected dynamic images (for example, summary images) image registration techniques may need to be used to place the segmentations and parameter maps into a single anatomic framework. The choice of registration technique to be used depends upon the organ system being imaged; the details of this are beyond the scope of this document. In performing registration techniques, either images aligned with the parameter maps or correlative images upon which the segmentation was performed are used as the target image for registration. The registered images are then interpolated from the source images. In interpolating parameter maps to match correlative images, tri-linear techniques are favored to avoid artifacts that may be associated with more advanced interpolation techniques.

10.2.4 Extraction of figures of merit . To create figures of merit from Ktrans or IAUGCBN parameter maps, each pixel value within the tumor ROI as applied to the parameter map (after elimination of end slices) should be considered. Median, mean and standard deviation of the pixel values should be calculated, and the median is considered the primary figure of merit. Situations requiring special consideration are listed, as follows:

i. Whole lesion curve analysis. As an alternative to analysis of pixel-by-pixel parameter maps, an analysis technique in which a single concentration-time curve within a lesion is extracted and a single Ktrans or IAUGCBN is calculated and reported is acceptable if specifically justified by data quality or lesion motion .

ii. Multiple lesions. In a patient with multiple lesions due to metastatic disease, each lesion should be reported and tracked separately. In a patient with multiple lesions due to recurrent local tumor (for example, recurrent glioblastoma) per-patient figures of merit should be reported by aggregating the results of the multiple lesions.

Choice of time point for segmentation. As a rule, the Ktrans or IAUGCBN at a given time point should be extracted using tumor ROIs segmented from the same imaging examination. However, in the situation where anti-angiogenic therapies are evaluated and post-therapy imaging is performed within 72 hours of initial treatment with the anti-angiogenic agent, it is acceptable to use a recent (within 1 week) pre-therapy time point to provide the segmentation used to define the lesion on the immediate post-therapy imaging session. In this case, it is presumed that changes in the appearance of lesions on immediate post-therapy study are due to immediate decreases in permeability or blood flow rather than decrease in lesion volume .

I would consider adding an additional statement. Not all may agree. /mr:  In settings were analysis is performed retrospectively, all time points should be made available to the reader simultaneously to allow for consistency in choice of tumor(s) for segmentation, and to ensure that similar regions of large tumors have been sampled and segmented. In the case of manual AIF segmentation, such work flow analyses also allow for greater standardization of the region of the aorta or other artery used in the analysis. In such settings, the reader should be blinded to the nature of each time point, so that inherent bias in tumor and/or AIF segmentation does not influence the results.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

10.3. Required Characteristics of Resulting Data

**Place requirements on characteristics of resulting data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

10.4. Platform-specific Instructions

Appendix G.5 provides instructions for specific models/versions that can be expected to produce data meeting the requirements of Section 10.3.

10.5. Archival Requirements

See 11.7.

10.6. Quality Control

See 12.7.

11. Archival and Distribution of Data

Archival and data distribution procedures are recommended so that all analysis results can be recomputed for verification and validation purposes. In addition to saving of all original images in DICOM formats, the following information must be archived along with the image data:

i. Detailed specification of the input function selection. This may include a binary mask of pixels selected for arterial input function, or may consist of a tabulated text file containing RAS coordinates (or x,y,z) co-ordinates of the AIF pixel locations.

ii. Saving all parameters and users inputs required for image registration if used. Time-series image registration may be used to align data spatially over time. Any parameters which control the performance of the registration algorithm (metric used, optimization parameters, user click points/sub regions used for alignment, etc) must be stored in suitable format. It is preferred to save the registration transform parameters so that identical registration can be reproduced in a multi-center environment.

iii. All regions of interest where analysis is performed and statistics are performed should be saved .

When data are distributed, such data must be properly anonymized according to HIPAA or other patient information protection regulations and can only be distributed after meeting all requirements specified by the local IRB.

11.1. Central Management of Imaging Data

The following requirements are placed on the central management of imaging data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

11.2. De-identification / Anonymization Schema(s) to Be Used

The following requirements are placed on de-identification schema(s).

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

11.3. Primary Source Imaging Data

The following requirements are placed on primary source imaging data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

11.4. Reconstructed Imaging Data

The following requirements are placed on reconstructed data.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

11.5. Post-Processed Data

The following requirements are placed on post-processed data.

| Aarameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

11.6. Analysis Results

The following requirements are placed on analysis results.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

11.7. Interpretation Results

The following requirements are placed on interpretation results.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12. Quality Control

**This is the section where all aspects of Quality Control are identified and described.

12.1. QC Associated with the Site

The following sectins consider various aspects of quality control.

12.1.1. Quality Control Procedures

**Describe required procedures and documentation for routine and periodic QC for the site and various pieces of equipment.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.1.2. Baseline Metrics Submitted Prior to Subject Accrual

**List required baseline metrics and submission details.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.1.3. Metrics Submitted Periodically During the Trial

**List required periodic metrics and submission details.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.2. QC Associated with Imaging-related Substance Preparation and Administration

The following requirements are placed on QC associated with imaging-related substance preparation and administration.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.3. QC Associated with Individual Subject Imaging

The following sections describe aspects of QC associated with individual subject imaging.

12.3.1. Phantom Imaging and/or Calibration

The following requirements are placed on phantom imaging and/or calibration.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.3.2. Quality Control of the Subject Image and Image Data

Tumor selection and Image QC prior to the DCE-MRI scan

Axial imaging pre-gadolinium should be performed using available sequences per institutional routine. Imaging in an end-expiratory breath-held position is encouraged, especially when imaging the chest or upper abdomen. Both T1 and T2 imaging should be performed and inspected to identify putative target lesions that may be unacceptable for DCE-MRI due to the presence of previously unidentified cystic or hemorrhagic areas.

Once the target lesion is identified, the DCE-MRI volume (oblique coronal or oblique sagittal slab) should be prescribed, using the axial T1W or T2W slice through the center of the target lesion, with the center of the slab bisecting both the target tumor and the representative artery (usually the aorta). ). If local reference tissue is included in the FOV, it allows the use of reference tissue methods to measure the AIF. It is recommended for central laboratory quality assessment that the grid lines representing each slice of the DCE-MRI 3D slab be overlaid on this representative pre-Gad T1 or T2 image, and saved as a separate series for exportation with the remainder of the DCE-MRI study. This overlay image should serve as a guide for reproducible placement of the DCE-MRI 3D slab at subsequent patient visits.

A representative 3D SPGR sequence is then obtained using the local torso array coils. Breath-hold (end-expiration) imaging is recommended to minimize the artifact from this image set. This sequence should be inspected for additional artifact (especially phase wrap) to identify whether such artifact occur in areas that overlap either the target tumor lesion or the arterial vessel. As the image contrast in this sequence is often low, an anatomic expert should be available to aid the technologist in this assessment. If unacceptable wrap is seen, adjustments to the imaging protocol (increase in the FOV, decrease in the degree of rectangular phase reduction—if allowed without unacceptable increase in temporal resolution) should be undertaken with repeat imaging. Patient repositioning may also be attempted at this point to reduce artifacts. If, despite all effort, an acceptable DCE-MRI 3D slab image set cannot be obtained, then an alternate target tumor should be selected, if possible .

Centralized post acquisition DCE-MRI quality control and assessment

In practice, DCE-MRI image sets must be assessed for quality before quantitative analyses are performed. While it is beyond the scope of this profile to detail estimates of DCE-MRI technical “success” in clinical practice or in clinical trial settings, it is recognized that within a clinical trial cohort, a fraction of DCE-MRI data sets may be considered suboptimal due to errors in patient selection, DCE-MRI exam prescription, magnet performance during the DCE-MRI study, or other factors, any or all of which may lead to spurious errors in quantification of tumor Ktrans or IAUCG. Recognition of these factors will mitigate reporting of spurious DCE-MRI changes in patients, and will facilitate DCE-MRI technical assessment.

In both technical assessment and early phase clinical trials, censuring of data for reporting should be discouraged, as this may lead to overt or inadvertent biases in the ensuing results. However, it may be appropriate to assess individual data sets for overall technical success. In such cases, in vivo results—be they for reproducibility assessment, prognostic factors, or response to therapeutic interventions—may be reported for both the entire cohort (analogous to an intention-to-treat reporting) and for the subset of subject with technically “adequate” studies. The remainder of this section refers to methods for evaluating the technical adequacy of data sets.

A) Confirming patient/lesion suitability for DCE-MRI

As discussed in section XX above, patients suitable for DCE-MRI analysis must possess at least one tumor of sufficient size. This requires that there be at least one tumor (designated the DCE-MRI target lesion) with a longest diameter >=2cm in anatomically fixed locations (defined in section XX) or >=3cm or greater in anatomically non-fixed areas. Furthermore, the target lesion for DCE-MRI analysis should be well removed from areas subject to large degrees of cardiac pulsatility artifact (e.g., hilar/subhilar lung regions, lateral left lobe of the liver), as these artifacts may yield unacceptable signal intensity fluctuations, rendering DCE-MRI quantitative imaging biomarkers of questionable validity.

Depending on the specifications of individual protocols, use of target lesions that violate specified acceptable target lesion attributes should be rejected. In rare instances, a second non-target lesion meeting requirements for target lesion status may be visible within the analyzable DCE-MRI volume. In such cases, these lesions may be analyzed as a substitute for the intended DCE-MRI target lesion.

B) Confirming fidelity of DCE-MRI exam prescription

The DCE-MRI exam includes a set of images used to quantify baseline (pre-gadolinium) tumor R1 (1/T1) values (termed T1 mapping series), a second set used for removal of local coil influences and a dynamic gadolinium enhanced series. These series must adhere to required prescription constraints in order for DCE-MRI quantification to be valid.

Unless specifically indicated (e.g., flip angle variation in a T1-mapping series), all imaging parameters (e.g., TR, TE) and the slab imaging geometry must be held constant thoughout the DCE-MRI series. Gross deviations from accepted parameter values or geometric prescription will cause the quantification to be invalid. Furthermore, automatic gain recalibration must be disabled to ensure that signal intensities are scaled similarly throughout the exam prescription.

The DCE-MRI series must also be assessed for adequate quality. The dynamic enhanced series must be run with adequate temporal resolution (discussed in section XX) to ensure accurate quantification of the arterial input function and tumor enhancement curve. Baseline (pre-gadolinium) imaging of sufficient duration must be obtained to ensure that the correct baseline tumor signal intensity is obtained. The total scan duration must also be adequate to ensure sampling of the tumor enhancement curve.

C) Fidelity of contrast administration

During the dynamic enhanced imaging, the gadolinium bolus must be adequately administered. Bolus administration is required to for adequate dynamic range of contrast enhancement. For longitudinal DCE-MRI studies, the type and dose of gadolinium contrast must remain the same. Deviations form this may require rejection of a subset of patient exams for analysis

D) Fidelity of DCE-MRI slab placement

For axial slab placement (brain, neck, extremities) the DCE-MRI slab should be centered with the center of the target tumor. For torso imaging, the oblique coronal slab should bisect the target tumor and the representative arterial vessel used to evaluate the arterial input function (usually the descending thoracic or abdominal aorta). In the event that the center slice of the slab does not bisect these two structures, the slab must be wide enough such that the center of both the target tumor and the aorta do not lie in the last 2 slices at the edge of the slab . The number of acceptable slices will therefore depend on the number of slices in the slab prescription. If 12 slices are prescribed, then the tumor and the aorta must be identified between slices 3 and 10.

E) Reproducibility of DCE-MRI study prescription

Repeat DCE-MRI examination, whether for determination of test-retest reproducibility of DCE-MRI quantification or for evaluation of change in tumor DCE-MRI characteristics in the setting of an anti-tumor therapy, must be performed with strict adherence to protocol characteristics as defined by the initial imaging session. The following parameters should be assessed in evaluating the adequacy of the DCE-MRI study

1. Patient positioning should be similar between the two sessions. Prone positioning for alleviation of claustrophobia should not be undertaken unless such positioning was performed at the initial session. Arm positioning (by sides or above the head) should adhere to original positioning. Number of and placement of local receive coils should be stable between DCE-MRI studies.

2. Patients should be land-marked in a near-identical manner as performed for the initial scan. Stability of patient landmarking should be assessed by evaluation of the location of a specific fixed area of anatomy (such as a certain spinal level or other rigid osseous landmark). The z-position of this landmark should not deviate by more than XX mm between the two (or more) examinations.

3. Study parameters should not have varied between the two sessions. Specifically, the TR, flip angle(s), and DCE-MRI slab acquisition parameters should be identical. When oblique slab prescription is utilized, the obliquity of the slab prescription should be as close as possible to that used in the original study. At best, the obliquity of the slab plane should not deviate by more than XX degrees from the prior placement.

4. Reproducibility of slab geometry placement by the technologist should be assessed via evaluation of the DCE-MRI image sets. Slice location overlays on representative axial images should be used to determine if the 3D oblique slab was positioned reproducibly. While slight variations in slab obliquity are expected and acceptable, the relationship between the slab and the center of the target lesion and artery should be relatively constant for each DCE-MRI session. In addition, z-centering of the slab should be constant between studies, as gross variation in z-positing of the center of the slab relative to the tumor and artery may lead to unacceptable deviations in signal intensities as portions of the image extend toward the ends of the magnet bore. For larger bore magnet systems, the tumor center should be within XX mm of isocenter in the z position. For short bore magnets, the position of the tumor in relation to isocenter should be XX mm.

F) Assessing image artifacts during DCE-MRI study

As the DCE-MRI study optimizes contrast and temporal resolution, volumetric coverage during the DCE-MRI study is minimized. Such MRI prescriptions may lead to image artifacts, including wrap (aliasing). These artifacts should be relatively constant during the MRI study. Such artifacts are acceptable as long as they do not interfere with the evaluation of signal intensity for the target DCE-MRI lesion and/or the arterial vessel or reference tissue. Fixed aliasing artifacts across any of these structures will render the DCE-MRI image set non-analyzable

Motion artifacts are inevitable in the DCE-MRI studies. As the dynamic enhanced study is performed during free respiration, tumors located in the chest and upper abdomen will undergo motion during the study, and as such, respiratory artifacts will be present. However, the magnitude of these artifacts should be such that they do not obscure the tumor or representative arterial vessel during the course of the dynamic enhanced imaging. If individual time points demonstrate such artifacts, they should be discarded from the dynamic enhanced series.

Additional artifacts, due to magnet instability or random electronic noise may also be seen during the course of DCE-MRI imaging studies. The DCE-MRI series should be assessed for such artifacts. Gross patient motion artifacts may also be encountered in the course of a DCE-MRI study. Such artifacts arise when the patient undergoes voluntary or involuntary shifting of position during the DCE-MRI study. All DCE-MRI image sets should be assessed for evidence of such motion. If gross motion within or between DCE-MRI series is noted, the DCE-MRI series should be rejected as non-analyzable.

One method for assessing for random magnet instability or gross patient artifacts is to evaluate the dynamic signal intensity of a region of interest in non-enhancing adipose tissue in the posterior aspect of the patient, as this area of anatomy should be relative motion-free during the course of the dynamic enhanced imaging. If the signal intensity is seen to deviate by more than 10% of the mean signal intensity, then either random machine noise or gross patient motion should be suspected. If the deviation is transient (e.g., noise “spike”) the time point or points affected may be discarded and the remainder of the DCE-MRI series evaluated. If the deviation is fixed over a portion of the dynamic enhanced image set, this is usually a sign of gross patient motion. Such data sets should be rejected as non-analyzable.

G) Monitoring and reporting quality assessment in DCE-MRI studies

As discussed earlier in this section, it is expected that within a given cohort, a subset of DCE-MRI studies may provide poor quantified data due to errors discussed above. In general, protocols including DCE-MRI assessment should discuss specific methods for handling and reporting such data sets. It is the opinion of the committee that no data sets should be deemed quantitative DCE-MRI “failures” without explicit documentation of the reasons for such failure. Furthermore, depending on the nature of the study endpoints, DCE-MRI analysis should proceed when feasible for all such data sets, regardless of image QC measures. In such cases, result reporting my include cohort-wide and subset analysis, the latter including only those data sets that meet pre-specified quality control standards. At a minimum, the number of DCE-MRI cases and patients excluded from analysis, and the reasons for such exclusions, should be specified in all result reporting.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.4. QC Associated with Image Reconstruction

The following requirements are placed on QC associated with image reconstruction.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.5. QC Associated with Image Processing

The following requirements are placed on QC associated with image processing.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.6. QC Associated with Image Analysis

The following requirements are placed on QC associated with image analysis.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

12.7. QC Associated with Interpretation

The following requirements are placed on QC associated with interpretation.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13. Imaging-associated Risks and Risk Management

MR safety considerations are to be established individually at each institution according to each institutions' radiology departmental guidelines and institutional review board (IRB) considerations to include policy guidelines on (1) laboratory screening for renal dysfunction prior to gadolinium based contrast administration (2) contrast administration in pregnant patients and in patients who are lactating (3) policy on patients receiving gadolinium based agents who have a positive history of a previous adverse event or events to iodinated or gadolinium based contrast agents to include serious and non-serious adverse events. The American College of Radiology Manual on Contrast Media Version 7 2010 can serve as a referenced guideline for each institutional policy development. This manual reflects policy statements previously released by the Food and Drug Administration (FDA) in the United States and its counterpart in the European Union, The Committee for Medicinal Products for Human Use (CHMP).

13.1. Radiation Dose and Safety Considerations

The following requirements are placed regarding radiation dose and safety considerations.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13.2. Imaging Agent Dose and Safety Considerations

The following requirements are placed regarding imaging agent dose and safety considerations.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13.3. Imaging Hardware-specific Safety Considerations

The following requirements are placed regarding hardware-specific safety considerations.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13.4. Management and Reporting of Adverse Events Associated with Imaging Agent and Enhancer Administration

The following requirements are placed regarding management and reporting of adverse events associated with imaging agent and enhancer administration.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

13.5. Management and Reporting of Adverse Events Associated with Image Data Acquisition

The following requirements are placed regarding management and reporting of adverse events associated with image data acquisition.

| Parameter | Compliance Levels |
| --- | --- |
|  | Acceptable Target Ideal |

Appendix A: Acknowledgements and Attributions

**

Appendix B: Background Information

A growing understanding of the underlying molecular pathways active in cancer has led to the development of novel therapies targeting VEGF, EGFR-tk, PI3-k, mTOR , Akt and other pathways. Unlike the conventional cytotoxic chemotherapeutic agents, many of the molecularly-targeted agents are cytostatic, causing inhibition of tumor growth rather than tumor regression. One example is anti-angiogenesis agents, which are presumed to act through altering tumor vasculature and reducing tumor blood flow. In this context, conventional endpoints such as tumor shrinkage may not be the most effective means to measure therapeutic responses. Functional imaging is an important candidate biomarker to predict and monitor targeted treatment response and to document pharmacodynamic response.

Dynamic contrast enhanced magnetic resonance imaging (DCE-MRI) represents an MRI-based method to assess tumor vascularity by tracking the kinetics of a low-molecular weight contrast agent intravenously administered to patients that highlights the tumor vasculature. The emerging importance of angiogenesis as a cancer therapy target makes assays of vascularity important to clinical research and future clinical practice related to targeted cancer therapy. There are multiple literature reports of the application of DCE-MRI to predict and detect changes associated with angiogenesis targeted therapy (Wedam, 2006, Rosen, 2004; Dowlati, 2002; Stevenson, 2003, Morgan et al, JCO 2003, Flaherty et al Cancer Biol Ther 2008, Liu et al JCO 2005, Drevs et al JCO 2007). Further, there is interest in the application of quantitative DCE-MRI to characterize contrast enhancing lesions as malignant in several organ systems including breast and prostate.

QIBA recognizes the potential importance of DCE-MRI as a functional imaging biomarker of angiogenesis. As a result, the DCE-MRI QIBA committee was formed to define the basic standards for DCE-MRI measurements and quality control.

The DCE-MRI technical committee is composed of scientists representing the imaging device manufacturers, image analysis laboratories, biopharmaceutical industry, academia, government research organizations, and professional societies, among others. All work is classified as pre-competitive. The goal of the DCE-MRI committee is to define basic standards for DCE-MRI measurements and quality control that enable consistent, reliable and fit-for-purpose quantitative Ktrans (1) and IAUGCBN (2) results across imaging platforms, clinical sites, and time.

This committee has developed the following profile claim for in vivo DCE-MRI applications based on the recommended standards, protocols, and quality control programs as well as previously published findings [Ng, Ferl, Roberts, Galbraith, Ashton, Padhani] and preliminary measures obtained by the committee.

Appendix C: Conventions and Definitions

Definitions: Review this document and define relevant terms and notations here.

Acquisition vs. Analysis vs. Interpretation: This document organizes acquisition, reconstruction, post-processing, analysis and interpretation as steps in a pipeline that transforms data to information to knowledge. Acquisition, reconstruction and post-processing are considered to address the collection and structuring of new data from the subject. Analysis is primarily considered to be computational steps that transform the data into information, extracting important values. Interpretation is primarily considered to be judgment that transforms the information into knowledge. (The transformation of knowledge into wisdom is beyond the scope of this document.)

Bulls-eye Compliance Levels Acquisition parameter values and some other requirements in this protocol are specified using a “bullseye” approach. Three rings are considered from widest to narrowest with the following semantics:

ACCEPTABLE: failing to meet this specification will result in data that is likely unacceptable for the intended use of this protocol.

TARGET: meeting this specification is considered to be achievable with reasonable effort and equipment and is expected to provide better results than meeting the ACCEPTABLE specification.

IDEAL: meeting this specification may require unusual effort or equipment, but is expected to provide better results than meeting the TARGET.

An ACCEPTABLE value will always be provided for each parameter. When there is no reason to expect better results (e.g. in terms of higher image quality, greater consistency, lower dose, etc.), TARGET and IDEAL values are not provided.

Some protocols may need sites that perform at higher compliance levels do so consistently, so sites may be requested to declare their “level of compliance”. If a site declares they will operate at the TARGET level, they must achieve the TARGET specification whenever it is provided and the ACCEPTABLE specification when a TARGET specification is not provided. Similarly, if they declare IDEAL, they must achieve the IDEAL specification whenever it is provided, the TARGET specification where no IDEAL level is specified, and the ACCEPTABLE level for the rest.

Appendix D: Documents included in the imaging protocol (e.g., CRFs)

**(Material the site needs to submit)

**Subject preparation

**Imaging agent dose calculation

**Imaging agent

**Image data acquisition

**Inherent image data reconstruction / processing

**Image analysis

**Interpretation

**Site selection and Quality Control

**Phantom Imaging and Calibration

Appendix E: Associated Documents (derived from the imaging protocol or supportive of the imaging protocol)

**e.g. the Imaging Charter, Site Manual, Standard Operating Procedures, etc.

Appendix F: TBD

Appendix G: Model-specific Instructions and Parameters

The presence of specific product models/versions in the following tables should not be taken to imply that those products are fully compliant with the QIBA Profile. Compliance with a profile involves meeting a variety of requirements of which operating by these parameters is just one. To determine if a product (and a specific model/version of that product) is compliant, please refer to the QIBA Conformance Document for that product. G.1. Image Acquisition Parameters The following technique tables list acquisition parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

These technique tables may have been prepared by the submitter of this imaging protocol document, the clinical trial organizer, the vendor of the equipment, and/or some other source. (Consequently, a given model/version may appear in more than one table.) The source is listed at the top of each table.

Sites using models listed here are encouraged to consider using these parameters for both simplicity and consistency. Sites using models not listed here may be able to devise their own acquisition parameters that result in data meeting the requirements of Section 7.1 and conform to the considerations in Section 13.

In some cases, parameter sets may be available as an electronic file for direct implementation on the imaging platform.