X.  Imaging Protocol (Italicized text within draft consensus statements indicates editorial suggestions compatible with the intent of the reference source materials but not extracted directly from those source materials)

- Executive Summary

Provide a brief (less than 250 words) synopsis to let readers quickly determine if this imaging protocol is relevant to them.  Sketch key details such as the primary utility, imaging study design, specific aims, context, methods, expected results, risks, and deliverables.

Consolidated Statement - (draft pending completion of the extraction process) – This UPICT Protocol is intended to guide the performance of whole-body FDG-PET/CT within the context of single- and multi-center clinical trials of oncologic therapies by providing acceptable (minimum), target, and ideal standards for all phases of the imaging examination as defined by the UPICT Template (ref) with the aim of minimizing intra- and inter-subject, intra- and inter-platform, inter-examination, and inter-institutional variability of primary and/or derived data that might be attributable to factors other than the index intervention under investigation.  The specific potential utilities for the FDG-PET/CT study(ies) as performed in accordance with this Protocol within any particular clinical trial could be to utilize qualitative, semi-quantitative, and/or quantitative data for single time point assessments (e.g., diagnosis, staging, eligibility assessment, investigation of predictive and/or prognostic biomarker(s)) and/or for multi-time point comparative assessments (e.g., response assessment, investigation of predictive and/or prognostic biomarker(s)).  More generally, such standardization of FDG-PET/CT within the conduct of clinical trials should 1) support internal decision-making in drug, biologic, and device development, 2) provide data to support registration and market-label indications, and 3) allow the eventual qualification of one or more imaging biomarkers (perhaps as surrogates for clinical endpoints) by supporting meta-analyses of multiple clinical trials (possibly over different compounds or devices and as contributed by different companies).

This document does include specifications for the performance of CT for the purposes of attenuation correction and/or localization, but does not address the performance of diagnostic CT within the context of FDG-PET/CT; although the integration of diagnostic CT in conjunction with FDG-PET/CT for oncology is acknowledged as potentially useful and appropriate.  When the integration of diagnostic CT is desired as part of the imaging protocol within the clinical trial, specifications for the CT portion of the imaging protocol may be derived from other UPICT protocol(s).

While focused primarily on the use of FDG-PET/CT in the conduct of oncologic clinical trials, this protocol also may have utility for guiding the performance of high quality imaging studies in clinical practice.

- Context of the Imaging Protocol within the Clinical Trial

Describe how this imaging protocol interfaces with the rest of the clinical trial.

General Note Regarding Study Design:  If quantitative FDG-PET/CT is to be used towards either primary, secondary, or exploratory aims, the study should include specific directions as to the management of subjects with abnormal fasting blood glucose measurements whether known to be diabetic or not.  While there is a paucity of scientific data to suggest that subjects with abnormal blood glucose measurements should be excluded from clinical trials that use FDG-PET/CT scan data, it is important to define how such subjects and the data from their imaging studies are managed to ensure comparability of imaging data within and among clinical trials.  Specifically, consideration should be given to the exclusion of subjects with abnormal fasting blood glucose when quantitative FDG-PET/CT is being used as the study’s primary endpoint.

  - Utilities and Endpoints of the Imaging Protocol

Describe one or more utilities or endpoints this Imaging Protocol could serve in a Clinical Trial.

(e.g. to determine eligibility of potential subjects in the clinical trial; to triage eligible subjects into cohorts based on stage or severity of disease; to assess response to treatment; to establish the presence of progression for determining TTP, PFS, etc.; to monitor for adverse events; to establish a database for the development, optimization, and validation of imaging biomarkers, etc.)

Consolidated Statement – The specific utility(ies) for which the FDG-PET/CT imaging study(ies) in the clinical trial authored using this UPICT Protocol should be clearly stated within the documentation of the clinical trial.  This Protocol has been derived from various referenced standards documents and publically listed clinical trials.  The specific utilities for the FDG-PET/CT imaging stated in those source materials include:

- diagnosis and staging of tumors (EU, Neth, ACRIN protocol #s 6671, 6685)

- prognostic stratification / biomarker (Neth, Hallett, ACRIN protocol # 6685)

- treatment planning or triage (ACRIN protocol # 6685)

- edge detection of tumors in radiotherapy planning (EU)

- lesion localization and characterization (EU, ACRIN protocol #s 6671, 6685)

- evaluate and quantitate tumor response / predictive stratification / biomarker (EU, Neth, Hallett, NCI, ACRIN protocol #s 6665, 6678, SNM GHS, PERCIST)

- correlation between imaging and tissue biomarkers and/or pathway activity (ACRIN protocol # 6665)

  - Timing of Imaging within the Clinical Trial Calendar

Describe for each discrete imaging acquisition the timing that will be considered “on-schedule” preferably as a “window” of acceptable timing relative to other events in the clinical trial calendar.  Consider presenting the information as a grid which could be incorporated into the clinical trial calendar.

Consolidated Statement - The study protocol should specifically define an acceptable time interval that should separate the performance of FDG-PET/CT image acquisition from both (1) the index intervention and (2) other interventions (e.g. chemotherapy, radiotherapy or prior treatment).  If response assessment will be based on serial FDG-PET/CT imaging studies, the time interval between the baseline study and the initiation of treatment should also be specified as well as the time intervals between subsequent FDG-PET studies and cycles of treatment.  Additionally, the study protocol should specifically define an acceptable timing variance for performance of FDG-PET/CT around each time point at which imaging is specified (i.e., the acceptable window of time during which the imaging may be obtained “on schedule.”  The timing interval and window are entirely dependent upon 1) the utility for the FDG-PET/CT imaging within the clinical trial, 2) the clinical question that is being investigated, and 3) the specific intervention under investigation.  There is some difference of opinion based on the reference source and the specific index intervention.  Suggested parameters for timing of FDG-PET/CT within oncologic trials include:

- When results of FDG-PET/CT are a study entry criterion the scans ideally should be performed within 21 days before initiation of the therapeutic intervention.  It should be noted that tumors with low FDG uptake (also see Sections 9 and 10) may not be suitable for follow-up studies of treatment response with PET (PERCIST cut off as an example).

- For FDG-avid and evaluable tumors, the minimum interval between the last dose of chemotherapy or biologic therapy and FDG-PET ideally should be 10 days (EU) with an acceptable interval of 14 days (Neth, NCI), if possible;

- As an alternative if FDG-PET/CT is being used during an ongoing treatment schedule (perhaps as an early predictor of response), the test should be performed at an interval within the treatment schedule that is determined by factors including, but not limited to, the type of treatment, specific cancer diagnosis, specific treatment target, and details of the treatment schedule itself.  For example, if the FDG-PET/CT will be performed between cycles that have no “break,” the scan might be performed as close to the start of the next cycle as possible (EU).  However, if the FDG-PET/CT will be performed within a treatment plan that incorporates periodic “breaks” between sets of treatment cycles, the scan might be performed shortly after the completion of the preceding cycle rather than after the “break” and therefore prior to the next cycle.

- In trials of or including radiation treatment, an interval up to 4 months may be required (Neth), although many investigators recommend a minimum delay after radiation therapy of 6-8 wks or longer before performing the post-treatment FDG-PET study (NCI).  Studies evaluating completeness of response should be performed later, however investigational studies used to modify therapy or predict outcome may be performed during therapy.

- When FDG-PET/CT is used for post-treatment response assessment in lymphoma, imaging should not be performed before at least 3 weeks after chemotherapy and preferably 8 – 12 weeks after completion of radiotherapy per the consensus statement of the Imaging Subcommittee of the IHP in Lymphoma as published in JCO 25(5) p572 (EU).  For intra-therapy evaluation please see bullet #3 above.

- An issue that must be addressed in the study-specific clinical trial protocol is the specific windows about each time point that would constitute an appropriate variance for that specific clinical trial

  - Management of Pre-enrollment Imaging –

Describe the evaluation, handling and usage of imaging performed prior to enrollment.

Clearly identify purposes for which such imaging may be used: eligibility determination, sample enrichment, stratification, setting the measurement base-line, etc.

(e.g. What characteristics or timing will make the imaging acceptable for the purpose?

Will digitized films be accepted?
        Will low-grade images be annotated and/or excluded from parts of the trial?
        Is there normalization that should be done to improve low-grade priors?

How should such imaging be obtained, archived, transferred, etc.)

Consolidated Statement:  The imaging protocol must contain documentation as to how pre-enrollment imaging should be managed; specifically 1) will imaging obtained prior to enrollment be used as baseline imaging and 2) if so under what specific conditions.  It is suggested that the specific conditions should take into account technical factors related to the imaging platforms (PET and CT) as well as the biology of the disease and the specific interventions used in the trial.  In general, scans performed as standard clinical care on PET/CT scanners that have not been previously qualified for the clinical trial and/or not in conformance with the imaging protocol would not be acceptable for the clinical trial.  One reference suggests that PET/CT scanning performed within eight weeks prior to initiation of drug therapy could be used as the baseline study (ACRIN 6665).  While another source states that if the pre-enrollment PET/CT was performed on an imaging platform not approved for use in the trial or otherwise does not meet trial requirements, the scan should be repeated within the trial budget; but that studies that are performed on approved scanners and otherwise conforming to all trial specification will be accepted as baseline studies and will be subjected to the same QA as studies performed after registration (and even if originally obtained as standard of care imaging therefore might need to be repeated (ACRIN 6671, 6678).

  - Management of Protocol Imaging Performed Off-schedule

Describe the evaluation, handling and usage of imaging performed according to the Procedure below but not within the “on-schedule” timing window described in Section 1.2.

(e.g. For what purpose(s) may such imaging be used (for clinical decision-making; for data

analysis; for primary endpoints; for secondary endpoints; for continued subject eligibility

evaluation; to supplement but not replace on-schedule imaging, etc.)?

What characteristics or timing will make the imaging acceptable for the purpose?

Is there normalization that should be done to account for the schedule deviation?

What is the expected statistical impact of such imaging on data analysis?

How should such imaging be recorded, archived, etc.)

To Be Authored.

  - Management of Protocol Imaging Performed Off-specification

Describe the evaluation, handling and usage of imaging described below but not performed completely according to the specified Procedure.  This may include deviations or errors in subject preparation, the acquisition protocol, data reconstruction, analysis, interpretation, and/or adequate recording and archiving of necessary data.

(e.g. For what purpose(s) may such imaging be used (for clinical decision-making; for data

analysis; for primary endpoints; for secondary endpoints; for continued subject eligibility

evaluation; to supplement but not replace on-schedule imaging, etc.)?

What characteristics or timing will make the imaging acceptable for the purpose?

Is there normalization that should be done to account for the schedule deviation?

What is the expected statistical impact of such imaging on data analysis?

How should such imaging be recorded, archived, etc.)

Consolidated Statement:  Criteria should be included in the protocol that define acceptable, target, and ideal FDG-PET/CT imaging specifications and parameters.  Imaging studies judged to be sub-optimal, if performed for “standard of care” could be repeated at the discretion of the site if the site deems the scan clinically unacceptable (ACRIN 6671).  If the scan is judged unacceptable for research purposes, the study may be repeated as dictated by the protocol and informed consent.  The protocol should then state how the cost of such repeated studies should be managed within the trial budget (ACRIN 6678).

  - Management of Off-protocol Imaging

Describe the evaluation, handling and usage of additional imaging not described below.  This may include imaging obtained in the course of clinical care or potentially for research purposes unrelated to the clinical trial at the local site.

(e.g. For what purpose(s) may such imaging be used (for clinical decision-making; for data

analysis; for primary endpoints; for secondary endpoints; for continued subject eligibility

evaluation; to supplement but not replace on-schedule imaging, etc.)?

What characteristics or timing will make the imaging acceptable for the purpose?

Is there normalization that should be done to account for the schedule deviation?

What is the expected statistical impact of such imaging on data analysis?

How should such imaging be recorded, archived, etc.)

To Be Authored

  - Subject Selection Criteria Related to Imaging

    - Relative Contraindications and Remediations

Describe criteria that may require modification of the imaging protocol.

(e.g. subjects with kidney insufficiency are contraindicated for Contrast CT in this protocol, at the physicians discretion, kidney function may be re-evaluated prior to imaging to see if the insufficiency has resolved, or the subject may be evaluated for dialysis, etc.)

Consolidated Statement – Inability to comply with or tolerate the performance of FDG-PET/CT imaging may be a relative exclusion criterion for subjects in a clinical trial that depends upon FDG-PET/CT for a primary or secondary endpoint.  Examples of such relative contraindications include inability to remain motionless for the duration of the scan time or to lie flat for any number of reasons (e.g., severe congestive heart failure).  However, such relative exclusion criteria are not unique to FDG-PET/CT.  A plasma glucose level above the threshold as defined in Section 4.2.2 may necessitate the rescheduling of the FDG-PET/CT test to another day when the plasma glucose level is less than the defined threshold.  For this reason, subjects at risk for elevated plasma glucose levels should be scheduled early during the timing interval as specified in Section 1.2 so that if the test must be rescheduled the test date will still fall within the acceptable timing interval (See Section 1.2) so as to avoid a protocol deviation.  In addition, it is suggested that for subjects who are known diabetics that three serial morning fasting blood glucose determinations (using home test kits) with values of less than 200 mg/dl be obtained prior to scheduling the FDG-PET/CT test in order to assure that the test results may be valid within the context of the trial (see Sections 1.7.2, 3 and 4.2.2).  Relative contraindications become absolute (i.e., Imaging Exclusion Criteria) when they cannot be remediated.  When the FDG-PET/CT imaging endpoint is a trial endpoint, the subject would then be excluded from the trial.

    - Absolute Contraindications and Alternatives

Describe criteria that may fully disqualify the subject for the imaging protocol.  
 If possible, identify possible alternative imaging protocols.

(e.g. subjects with pacemakers are disqualified for this MRI protocol.  Consider using CT protocol UPICT-31254 instead)

These alternatives may also be useful for relative contraindications if remediations described in 1.7.1 are not possible or successful.

Consolidated Statement – The protocol should specifically define a threshold plasma glucose level that should represent an absolute exclusion criterion for participation in any clinical trial that depends on FDG-PET/CT imaging for any primary or a quantitative secondary endpoint if the plasma glucose level cannot be maintained below that threshold level using the diabetic management procedures as described in Section 4.2.2.  Threshold plasma glucose levels for inclusion as suggested by referenced standards documents and publicly listed clinical trials include:

- A plasma glucose level: ≤126 mg/dl (≈7.0 mmol/L) (EU)

- Blood glucose levels: ≤150 mg/dl (≈8.3 mmol/L) (ACRIN 6678)

- Blood glucose levels: ≤200 mg/dl (≈11.1 mmol/L) (ACRIN 6672, Neth)

- Per the SNM global harmonization summit, subjects known to be diabetic who have three serial fasting morning blood glucose levels of >200 mg/dl (despite adequate medical management) prior to the baseline or initial FDG-PET/CT study should be excluded from a clinical trial in which quantitative FDG-PET/CT is used for a primary endpoint.  When FDG-PET/CT is used towards secondary and/or exploratory endpoints the trial should specifically state whether subjects with fasting blood glucose levels >200 mg/dl will be included or excluded; and if included how the data from such subjects will be managed.  Furthermore, there are specific clinical trial purposes (e.g., pD determination) for which fasting blood glucose levels >200 mg/dl are acceptable.  Finally, there is a scientific gap in knowledge regarding the relationship between fasting blood glucose level and the effect on quantitative and qualitative FDG-PET/CT.  It is recommended that investigators utilize pooled data from studies performed under rigorous protocols (such as the UPICT Oncologic FDG-PET/CT protocol) to investigate this relationship – including data from subjects with fasting blood glucose levels >200 mg/dl. (SNM GHS)

Many clinical trials exclude subjects who are pregnant (or suspect they are pregnant) or breastfeeding when FDG-PET/CT is being used as a primary or secondary endpoint.  However, such potential subjects may already be excluded on the basis of the index intervention under investigation without regard to the use of FDG-PET/CT.

Additional suggested exclusion criteria include weight exceeding table limits (300 - 450 lbs for most current PET/CT scanners) and subjects with a history of life-threatening allergic / anaphylactoid reactions to any contrast media if contrast is being used in the study. (ACRIN 6671)

Relative contraindications become absolute (i.e. Imaging Exclusion Criteria) when they can no longer be remediated.  When the FDG-PET/CT imaging endpoint is a trial endpoint, the subject would then be excluded from the trial.

CONSENSUS:

Baseline (or for study recruitment):

1.  When the quantitative FDG-PET/CT results will be used as the primary endpoint for determining the primary aim in the clinical trial, it is suggested that diabetic subjects unable to achieve a fasting blood glucose level (without the use of short-acting / regular insulin within four hours preceding the FDG-PET/CT study) of <200 mg/dl prior to the baseline study be excluded from the trial.  If such subjects are included in the trial, it is suggested that subgroup analysis be performed to assess the effect, if any, that the inclusion of these subjects has on the study outcome.  The clinical trial protocol should state how the data generated by such subjects will be managed.  See also note on General Study Design Considerations Section 1.0.

2.  When the quantitative FDG-PET/CT results will be used as secondary endpoint(s) or exploratory endpoint(s) or if only qualitative FDG-PET/CT results are used for either primary or secondary or exploratory endpoint(s) in the clinical trial, it is suggested that the trial explicitly state whether subjects unable to achieve a fasting blood glucose level of <200 mg/dl prior to the baseline study will be included or excluded from the clinical trial.  If these subjects are not excluded, it is suggested that analysis of results be performed with and without the inclusion of these data to determine if the results are affected by their inclusion.  The clinical trial protocol should state how the data generated by such subjects will be managed.

On study management (potential for censor or special data management):

1.  When subsequent (follow up) FDG-PET/CT scans are obtained despite fasting blood glucose measurements outside of the inclusion threshold (typically <200 mg/dl) on subjects originally within the inclusion threshold, it is suggested that the magnitude of the change (in absolute value and in percentage change) in fasting blood glucose measurement be taken into account when making decisions related to censoring the subject(s) or managing their data.

2.  Consideration might be given to censoring subjects with significant (the definition of a significant change must be defined on a trial by trial basis) changes in fasting blood glucose between baseline and follow up studies (even if the fasting blood glucose levels are below the threshold for exclusion) when quantitative FDG-PET/CT is to be used as the primary study endpoint.  If these subjects are not censored, consideration might be given to data analysis to determine the effect, if any, that the inclusion of such subjects has on the trial outcome and/or using correction factors in calculating outcome values.  Of note, significant changes in blood glucose from baseline to follow up scans could be in either direction.  The clinical trial protocol should explicitly state how the data generated by such subjects will be managed.

    - Imaging-specific Inclusion Criteria

Describe inclusion criteria that are specifically related to the imaging portion of the study.

Consolidated Statement - One source states that for clinical trials with longitudinal FDG-PET measurements as a primary endpoint might require a minimum tumor FDG-avidity based on the SUV (e.g., >4, tumor to background ratio of >1.5 hepatic mean) at baseline in order to remain on or to be eligible for participation on the study and have subsequent follow-up FDG-PET/CT scans (ACRIN 6678).  There may also be lesion “size” threshold (RECIST, WHO, volume) and/or lesion multiplicity (stage) threshold for eligibility (See also sections 9 and 10).

- Site Selection, Qualification and Training (See also Section 12 relative to QC)

  - Personnel Qualifications –

    - Technical

Consolidated Statement:  Technologist:  Appropriate education, training, and certification of technologists is required to perform PET/CT. Representatives from the Society of Nuclear Medicine Technologist Section (SNMTS) and the American Society of Radiologic Technologists (ASRT) met in 2002 and published specific recommendations (J Nucl Med Technol. 2002;30:201204).  (Delbeke)

    - Physics

Consolidated Statement:  Physicist: The SNM considers certification and continuing education in the appropriate sub- field(s) to demonstrate that an individual is competent to practice one or more of the subfield(s) of medical physics and to be a qualified medical physicist. The SNM recommends that the individual be certified in the appropriate subfield(s) by the American Board of Radiology (ABR) or the American Board of Science in Nuclear Medicine (ABSNM).  (Delbeke)

    - Physician

Consolidated Statement:  Physician: Imaging experts interpreting PET/CT scans should have appropriate training in both PET and CT. A working group of representatives from the American College of Radiology, the Society of Nuclear Medicine (SNM), and the Society of Computed Body Tomography and Magnetic Resonance agree only appropriately trained, qualified physicians should interpret PET/CT images (J Nucl Med. 2005;46:12251239). This working group has also recommended the number of continuing medical education credits earned and the number of cases interpreted that would demonstrate adequate training. (Delbeke)

    - Other (e.g., radiochemistry, radiobiologist, pharmacist, etc.)

For oncologic FDG-PET/CT the qualifications of the personnel involved in the preparation of the FDG should be appropriate to comply with the FDA part 212 specifications or the international equivalent as appropriate.

  - Imaging Equipment

List required equipment and software packages, such as CT scanners, image processing workstations, and analysis packages.  Specific capabilities of the equipment are described in later sections of this document.  Include ancillary equipment in this section as well.

Consider discussing the trade-off between accrual rates and the “bullseye rings of compliance” described in Section 7.2.

Each site needs to have PET/CT system(s) (SNM GHS); hence,

- Multiple references suggest that integrated PET/CT scanners are preferable to be used for imaging based on increased accuracy for lesion localization and characterization than that obtained from the results obtained from PET and CT separately and interpreted sided by side or following software based fusion of the PET and CT datasets (EU).

- PET scanners that utilize NaI detectors are excluded. (NCI and ACRIN)

ACRIN trials require that PET and PET/CT scanners are qualified in order to participate in the trial and all data must be acquired on an ACRIN qualified scanner.

An important aspect of quantitative multi-center PET imaging studies and therefore integral to the qualification of imaging platforms is the cross-calibration of scanner performance across various imaging sites.  Several societies, organizations and clinical trials networks, such as the NCI, ACRIN, EORTC, EANM and SNM, etc. have developed multi-center clinical trials imaging guidelines and have set up or are setting up PET/CT system validation and site accreditation programs to ensure that data collected using these systems are comparable, i.e. can be exchanged.  These site accreditation programs use different phantoms for this purpose, among the performance characteristics that are tested are:  (1) the verification of a correct (cross-) calibration of the PET/CT system (against a dose calibrator) (ACRIN, Neth, EU, SNM), (2) scanner normalization and uniformity (ACRIN), and (3) the assessment of 2D or 3D SUV recovery coefficients (thereby essentially assessing contrast recovery and/or partial volume effects as a function of sphere size or rod diameter) (Neth, EU, SNM).  Despite the differences in the implementation of scanner validations, all site accreditation programs aim to assess image quality on some or all of these main image characteristics.  Future work should focus on further aligning the activities of these societies, either by harmonizing the scanner validation platforms/phantoms and development of a equivalent scanner multi-center QC program.  The latter should be feasible considering the good agreement between the societies regarding the image characteristics to be verified.  At present there is a strong interest from all groups in establishing a common FDG PET standard.

The SNM GHS states that site qualification by a standardized method (including, but not limited to, documentation of a rigorous quality control program incorporating the use of a uniform phantom to verify scanner normalization and calibration) is the minimum acceptable for clinical trials and use of a standardized multi-compartmental phantom (to additionally evaluate detectability, resolution and contrast recovery) at all sites for this purpose is the target.  (SNM GHS)  For a detailed discussion with materials and methods see Section 12.1.1 (and Appendix D???).

Initial and ongoing periodic QC for CT as used for attenuation correction and localization is included within the scope of this document (see 12.1.1 and Appendix D for detail).  However, QC for diagnostic CT performed in conjunction with oncologic FDG-PET/CT is not included within the scope of this document.  Documentation for diagnostic CT may be obtained from other UPICT documents (see url XXXXXX).

The sites also need to have all the ancillary equipment for conduct of the trial including, but not limited to, appropriately calibrated glucose measuring device, dose calibrators, stadiometer to measure height, and scales to weigh subjects.  See Section 12.1.1 for quality control. (SNM GHS)

Details of several such accreditation programs are contained in Section 12 and Appendix D.

  - Infrastructure – Not specified

List required infrastructure, such as subject management capabilities, internet capability, image de-identification and transmission capability.

  - Quality Control

    - Procedures

See 12.1.1.

    - Baseline Metrics Submitted Prior to Subject Accrual

See 12.1.2.

    - Metrics Submitted Periodically During the Trial

See 12.1.3.

Additional task-specific Quality Control is described in sections below.

  - Protocol-specific Training

Consolidated Statement – By definition Section 2.5 inclusive of all subsections will be directed by the specific needs and details of the protocol within which the FDG-PET/CT will be performed.  For non-protocol specific training and qualifications please see Sections 2.1 – 2.4 inclusive of all subsections.

    - Physician

See 10.5

    - Physics

See …

    - Technician

See …

- Subject Scheduling

Describe requirements and considerations for the physician when scheduling imaging and other activities, which may include things both related and unrelated to the trial.

Consolidated Statement:  Prior to scheduling potential and/or already accrued subjects for FDG-PET/CT with its inherent (albeit minimal) risks confirmation of appropriateness for imaging (e.g., history, physical examination, staging, biopsy for diagnosis, etc.) should be performed and documented.   Scheduling diabetic subjects may require special attention (please see Section 4.2.2 for additional details) and therefore this should be specifically queried at the time of scheduling.  At the time of scheduling, the study team should determine that inclusion of the subject does not violate any of the study-specific inclusion and exclusion criteria pertinent to the FDG-PET/CT study. (SNM GHS)  For considerations related to the scheduling of subjects who are known to be diabetic please also see Sections 1.7.2 and 4.2.2.

- Additional scheduling recommendations for diabetic subjects are suggested by two references (EU and Neth).  These include the following:

  - For type I diabetes:

    - Ideal to achieve euglycemia prior to PET study

    - Schedule study for late morning by eating normal breakfast at 7 am and taking normal amount of  insulin; then fast for at least 4 hours till exam

  - For type II diabetes:

    - Schedule study for late morning

    - Comply with at least 4 hour fast till exam

    - Continue oral medication (hypoglycemic) as usual

- One reference (ACRIN) suggests the following for diabetic management:

  - Diabetic subjects should be scanned early in the morning before the first meal, and doses of insulin and/or hypoglycemic medication should be titrated appropriately in consultation with the subject's referring physician.

Before scheduling an FDG-PET study, diabetic subjects should test their ability to maintain reasonable plasma glucose levels after fasting, while avoiding insulin close to the time that FDG would be administered.

- For known diabetic subjects with anticipated fasting blood glucose measurements for the day of the examination between 126 mg/dl and 200mg/dl, the following scheduling recommendations apply:

  - Ideal / Target:  Type I and Type II diabetic subjects should be scanned early in the morning before the first meal, and doses of insulin and/or hypoglycemic medication should be withheld if glucose levels remain in the acceptable range.  This should be established from morning blood glucose levels prior to the study.

  - Acceptable:  Type I and Type II diabetic subjects, who cannot reliably attain acceptable glucose levels early in the morning, should be scheduled for late morning, and should eat a normal breakfast at 7 am and take their normal morning diabetic drugs; then fast for at least 4 hours till exam.  This strategy is acceptable only for:

    - Non-quantitative PET/CT, or

    - Endpoints that are not for the primary aim, or

    - Subjects whose baseline study was performed with a FBG <200 mg/dl, but who have become uncontrolled hyperglycemics secondary to treatment effect, disease progression, or are being evaluated for exploratory endpoints

  - In each case, the goal is to achieve a fasting blood glucose with the prescribed range (e.g., ≤126,, ≤150,, or ≤200 mg/dl dependent on the clinical status of the subject, mechanism of therapy, and the utility of the FDG-PET/CT test in the clinical trial) (SNM GHS)

CONSENSUS:

Baseline exam scheduling:

1..For known diabetic subjects with a history of adequate control achieved without the use of short-acting insulin (continuous infusion or periodic frequent administration) that might be needed within the four hours preceding the proposed scheduled FDG-PET/CT study (routinely <200mg/dl daily fasting blood glucose), the exam should be scheduled as follows:

For known diabetic subjects with anticipated fasting blood glucose measurements for the day of the examination between 126 mg/dl and 200mg/dl, the following scheduling recommendations apply:

  - Ideal / Target:  Type I and Type II diabetic subjects should be scanned early in the morning before the first meal, and doses of insulin and/or hypoglycemic medication should be withheld if glucose levels remain in the acceptable range.  This should be established from morning blood glucose levels prior to the study.

  - Acceptable:  Type I and Type II diabetic subjects, who cannot reliably attain acceptable glucose levels early in the morning, should be scheduled for late morning, and should eat a normal breakfast at 7 am and take their normal morning diabetic drugs; then fast for at least 4 hours till exam.  This strategy is acceptable only for

    - Non-quantitative PET/CT, or

    - Endpoints that are not for the primary aim, or

    - Subjects whose baseline study was performed with a FBG <200 mg/dl, but who have become uncontrolled hyperglycemics secondary to treatment effect, disease progression, or are being evaluated for exploratory endpoints

2.  For known diabetic subjects with a longstanding or recent history of inadequate control or without routine daily fasting blood glucose testing results, it is suggested that three serial daily fasting blood glucose measurements be obtained with a value of <200mg/dl prior to scheduling the FDG-PET/CT examination.  If the fasting blood glucose measurements obtained during these serial tests cannot be maintained under the prescribed limit, please see section “inclusion/exclusion criteria, 1.7.1 and 1.7.2.”

3.  For subjects who achieve adequate control of FBG through the use of continuous or frequent use of regular insulin (i.e., within four hours preceding the proposed scheduled FDG-PET/CT study), please see section “inclusion/exclusion criteria, 1.7.1 and 1.7.2.”  Specifically, subjects who achieve adequate blood glucose control through the use of continuous infusion insulin pumps or frequent use of regular insulin should not be excluded from participation in a clinical trial if blood glucose levels are maintained at an adequate level (as defined to follow) after at least four hours without insulin infusion or administration.

4.  For subjects not known to be diabetic, please see sections 4, 5, and 7 (i.e., perform the procedure as scheduled).

Follow-up or repeat exam scheduling:

1..For known diabetic subjects with a history of adequate control (routinely <200mg/dl daily fasting blood glucose), the exam should be scheduled as per number one above (baseline exam scheduling).

2.  For known diabetic subjects with a longstanding or recent history of inadequate control or without routine daily fasting blood glucose testing results, it is suggested that three serial daily fasting blood glucose measurements be obtained with a value of <200mg/dl prior to scheduling the FDG-PET/CT examination.  If the fasting blood glucose measurements obtained during these serial tests cannot be maintained under the prescribed limit, the subject should be referred for best efforts at medical management to control the fasting blood glucose level and scanned at a time appropriate to that medical management plan (keeping in mind the previously stated precautions related to the use of insulin immediately prior to the scan acquisition).  The data interpretation and analysis plan should specify how the data from such subjects is managed.  Please see section “inclusion / exclusion criteria, section 1.7.1 and 1.7.2 for further information.”

3.  For subjects not known to be diabetic, please see Section 4, 5, and 7 (i.e., perform the procedure as scheduled).

4.  For subjects who are either diabetic or non-diabetic prior to the investigational protocol, but in whom the therapeutic intervention or deterioration of diabetic control for general health reasons is known or suspected to cause hyperglycemia, it is suggested that one or more fasting blood glucose measurements should be obtained prior to scheduling the FDG-PET/CT exam so as to ascertain if pre-imaging hyperglycemic management should be optimized prior to the imaging test.  For subjects with persistent fasting blood glucose measurements (>200 mg/dl), please see the “inclusion/exclusion criteria, 1.7.1 and 1.7.2” and “upon arrival testing, 4.2.2” sections.

  - Timing Relative to Index Intervention Activity

Consensus – Please see Section 1.2

  - Timing Relative to confounding Activities (to minimize “impact”)

(e.g. Avoid scheduling a biopsy on a tumor within X days prior to the FDG-PET scan to evaluate tumor viability; Avoid scheduling the MRI scan within X hours following administration of TPA (for stroke) to the subject.)

Consolidated Statement – Activities, tests, and interventions that might increase the chances for false positive and/or false negative FDG-PET/CT studies should be avoided prior to scans.  The allowable interval between the potentially confounding event and the imaging test will be dependent on the nature of the confounder.  For example, a percutaneous or excisional biopsy of a suspicious mass may cause focally increased FDG-PET activity or might lead to the appearance of a non-malignant mass (e.g., hematoma) on the CT portion of the study.  A percutaneous ablation procedure of a known malignant focus may cause focally increased FDG-PET activity and/or an immediate post-ablation increase in the apparent volume of the ablation target lesion.  The time of onset and the duration of the increased FDG-PET activity and/or the change in lesion volume might be different for these two different confounding factors.

If iodinated contrast is to be used for the CT portion of the PET/CT study, conflict with other tests and treatments should be avoided congruous with community standards of care (e.g., thyroid scan).

  - Scheduling Ancillary Testing

(e.g. order a blood draw to occur within X hours preceding the imaging procedure.)

Consolidated Statement – Avoid scheduling tests that might confound the qualitative or quantitative results of the FDG-PET/CT study within the time period prior to the scan.  For example, a glucose tolerance test should not be scheduled during the 24 hours prior to the performance of FDG-PET/CT.  Similarly, other tests that might involve increasing plasma glucose, insulin, or corticosteroid levels should also be avoided.  Exercise cardiac stress testing should be avoided during the twenty-four (24) hours prior to the performance of FDG-PET/CT.  Similarly, other tests that might involve vigorous exercise and thereby increase muscle metabolic function should also be avoided.

- Subject Preparation

  - Prior to Arrival

Describe the presence/absence/timing of subject activities that may impact the procedure or results. These items should typically result in instructions for the subject at time of scheduling, or reminders sent to the subject shortly prior to imaging.

(e.g. oral and/or IV intake, vigorous physical activity, non-protocol-related medical interventions, etc.)

Consolidated Statement - The main purpose of subject preparation is to reduce tracer uptake in normal tissue (kidneys, bladder, skeletal muscle, myocardium, brown fat) while maintaining and optimizing tracer uptake in the target structures (tumor tissue).  Below is a generally applicable protocol to address (1) Dietary, (2) Fluid Intake, and (3) Other activities that may impact the FDG-PET/CT procedure or results.

(1) Dietary (for the management of previously known or unknown diabetic subjects please see section 4.2.2):

  - According to two sources, subjects should not eat any food for an absolute minimum (acceptable level) of 4 hours prior to start of FDG-PET study (ACRIN), although the target pre-test fasting period is recommend as a 6 hour minimum (Neth, EU).  This can be achieved as follows:

- Subjects scheduled to undergo the PET study in the morning should not eat after midnight and preferably have a light meal during the evening prior to the PET study.

- Subjects scheduled for an afternoon PET study may have a light breakfast before 8 am.

- Medication can be taken as prescribed (see Section 4.2.2 for diabetic management)

  - Two sources have stated that a low carbohydrate diet should be followed for 24 hours before the study, culminating with fasting for the final six hours. (NCI, ACRIN)

  - One large study (ref) has suggested that a high-fat, low-carbohydrate meal is preferred for the last meal prior to commencing the period of fasting; as well one study has suggested that beta blockade may be useful to decrease brown fat uptake.  Although there are insufficient data to recommend these strategies as routine at this time (SNM GHS)

However, on the basis of the SNM harmonization summit the acceptable and target timing for discontinuing enteral nutrition is at least six (6) hours prior to the anticipated time of FDG administration. (SNM GHS)

(2) Fluid Intake:

All references note that adequate hydration (before and after FDG administration) is important (both to ensure a sufficiently low FDG concentration in urine (less artifacts) and for radiation safety reasons). Whichever hydration strategy is used (how much and when to administer), the protocol should be uniform among sites during a trial.   Specific hydration recommendations include: one reference (NCI) suggests oral intake of at least 710-1665 mls of water while fasting, an additional reference (ACRIN) recommends consumption of two to three 8-12 oz water (??? = 473-1065 mls) while fasting, and two references (Neth and EU) suggest 1 liter during 2 hours prior to FDG administration.

Subjects should be well hydrated.  They should be asked to drink at least 500 ml of non-sugar, non-caffeine containing fluid during two hours prior to the study.  If IV contrast is to be injected as part of the study, subjects should be asked to drink more fluid (total of 1 liter) during the two hours prior to the study. The fluid administered should not contain glucose or caffeine.  It is acceptable for subjects to receive non-glucose containing IV solutions such as normal or dilute saline. Lactated Ringer's solution is not acceptable and should be discontinued.  This hydration strategy should be modulated as clinically appropriate in subjects with certain medical conditions including, but not limited to CHF, RF, and fluid retention for example. (SNM GHS)

While two sources (EU, ACRIN) indicate that parenteral nutrition and intravenous fluids containing glucose should be discontinued at least 4 (acceptable) - 6 (target) hours before the PET examination.  The infusion used to administer intravenous prehydration must not contain any glucose.  However, on the basis of the SNM harmonization summit, the acceptable and target are both set to six (6) hours in the consensus protocol.  (SNM GHS)

(3) Other Activities:

The subject should avoid strenuous or extreme exercise before the PET exam to minimize uptake of the radiotracer in muscle (ACRIN), for a minimum acceptable time period of at least 6 hours (EU) but preferably for the target time period of 24 hours (NCI, EU, ACRIN) prior to the PET exam.  Based on the SNM global harmonization summit the acceptable and target time for avoidance of strenuous physical activity and cold exposure is 24 hours prior to the study.

Other activities that might be avoided are contained in sections 3 and 3.2.

Insert screening for claustrophobia here with ACRIN attribution.

Insert screening and pretreatment for iodinated contrast issues here with ACRIN attribution – separate paragraph.

  - Upon Arrival

    - Confirmation of subject compliance with instructions

(e.g. instructions to the admitting nurse/tech to confirm with the subject upon arrival that they have complied with each of the instructions in 4.1.)

Consolidated Statement – Upon arrival 1) confirmation of subject compliance with pre-procedure instructions and 2) the occurrence of potentially confounding events should be documented on the appropriate case report forms.  The documentation might include some or all of the following:

- timing, character, and amount of the most recent previous oral and/or intravenous intake of fluid and nutrients

- timing and dosages of relevant non-prescription and prescription medications taken prior to the PET/CT scan (e.g., the last cycle of chemotherapy or non-cytotoxic pharmacotherapy, administration of growth factors, cytokines, steroids, beta blockers, etc.)

- extent of physical activity and most recent exposure to cold temperature for the preceding 24 hours

- timing and description of medical procedures performed prior to the PET/CT scan (e.g., radiation therapy, biopsy, surgery)

- timing and description of relevant medical tests performed prior to the PET/CT scan (e.g., invasive tests and/or tests that involve the administration of exogenous substances and/or tests that involve vigorous physical activities)

- confirmation that the subject has completed the trial Informed Consent Document.

The FDG-PET/CT procedure should be explained to the subject and exam-specific consent should be obtained if that is the standard of care for the site or the standard established for the specific clinical trial.  There should be documentation of subject-specific risk factors including, but not limited to, previous contrast reactions (if iodinated contrast is to be used).

    - Ancillary Testing To Be Performed Upon Arrival

(e.g. blood draws, weight/blood pressure measurement, etc. associated with the imaging and downstream actions relative to such testing)

Consolidated Statement: - Subject height and body weight must be measured precisely with standardized measurement devices and with the subject in a gown or light clothing and recorded as the minimum acceptable standard (EU, Neth, NCI, ACRIN).  The target standard would add that for serial studies in the same subject, weight should be measured directly prior to each PET study since body weight often changes during the course of the study (EU, SNM GHS).

Blood glucose monitoring, measurement and documentation and the appropriate management/disposition of hyperglycemic/ diabetic subjects are addressed by all references and should be included as a minimum acceptable standard of performance.

- It is important to measure and document subject blood glucose level shortly prior to and target within the 2 hours prior to (ideally within 1 hour for all subjects and target within 1 hour for insulin-requiring diabetic subjects) FDG administration (all, SNM GHS).

- The ideal is blood glucose level < 120 mg/dL (all).

- Subjects with blood glucose level > 200 mg/dL should be rescheduled (all).

- For subjects with blood glucose measurements between 120 mg/dL and 200mg/dL, there are varying actions suggested by the different references.

  - Two references (EU, ACRIN) indicate that the subject must be rescheduled, and adjustments to diet and medications made if necessary, so that the fasting blood glucose concentration can be brought down to the acceptable range at the time of FDG injection, or excluded depending on the subject circumstances and the trial being conducted.

  - One reference (NCI) indicates that for diabetic subjects, the allowable glucose reference range is up to 150-200mg/dL.  The handling of non-diabetic subjects is not specifically addressed.

  - One reference (Neth) indirectly indicates that for diabetics (and non-diabetics?) with glucose levels between 120-200 mg/dl, PET scanning can proceed

  - There is no consensus from these references for diabetic or non-diabetic subject management in the glucose range of 120-200 mg/dL.  The imaging protocol for each individual clinical trial should indicate the glucose cut-off thresholds and the exact management for diabetic and non-diabetic subjects with plasma glucose levels between 120-200 mg/dl,  especially if the data from the FDG-PET/CT examination will be used towards a primary or secondary endpoint and/or will be compared in a serial manner over the course of the protocol.

- Secondary to recognized problems with administration of insulin (due to alteration of FDG biodistribution and diminished accuracy of SUV determination-NCI), insulin must not be given to reduce pre-FDG-administration glucose levels (EU, NCI), unless the interval between administration of insulin and FDG is more that 4 hours (EU).  One reference (ACRIN) suggests that if a hyperglycemic subject is to be rescheduled, the subject’s physician should be contacted to manage the blood glucose prior to the rescheduled study.

CONSENSUS:

Management of pre-scan fasting blood glucose results on the day of the FDG-PET/CT study is as follows:

1.  For known diabetics with FBG of <200 mg/dl, the baseline scan should be performed as scheduled, given the caveats described in sections “scheduling” and “inclusion/exclusion criteria.”  Specifically, the subject’s blood glucose of <200 mg/dl should not be related to administration of regular or short-acting insulin within the four hours preceding the PET study.

2.  For known diabetics with FBG of >200 mg/dl, the baseline scan and/or follow up scan(s) should be rescheduled only if the quantitative FDG-PET/CT results are to be used as a primary endpoint, except as provided for by the discussion in Section 3 consensus statement regarding baseline examination scheduling bullet point 2” and “inclusion/exclusion criteria, Section 1.7.2.”  However, if the FDG-PET/CT exam is contributing to secondary or exploratory endpoints and/or is being used as a qualitative measure, then the subject does not need to be rescheduled.

3.  For subjects not known to be diabetic with FBG >150 mg/dl, the baseline scan should be rescheduled and the referring physician should be contacted so that a diabetic work-up may be performed (unless it is clear that the subject merely did not follow dietary instructions properly, in which case the study should just be rescheduled).  If the subject is found to be diabetic, the subject should be managed as per the diabetic management protocol.  If the subject is found to be normal (not diabetic), then the subsequently performed FBG will be <150 mg/dl by definition.

4.  For subjects not known to be diabetic but in whom treatment, disease-progression effect, or physiology has caused hyperglycemia and/or a significant change in blood glucose from baseline, the scheduling and management of the FDG-PET/CT should take into account statements for Section 1.7.2 Consensus Statement regarding “On-study inclusion/exclusion criteria bullet points #1 and 2.”

    - Preparation for Exam

(e.g. empty bladder, removal of metal objects, etc.)

Consolidated Statement – In order to avoid artifactual distribution of the FDG, it is critical that subject preparation after arrival and prior to imaging are standardized among all sites and subjects throughout the conduct of the clinical trial. (EU, Neth, NCI, Hallet, ACRIN)

- The waiting and preparation rooms should be relaxing and warm (> 75° F or 22° C) during the entire uptake period (and for as long as reasonably practicable prior to injection, at least 15 minutes is suggested as acceptable).  Blankets should be provided if necessary. (SNM GHS)

- The subject should remain recumbent or may be comfortably seated; activity and conversation should be kept to an absolute minimum.  For example, the subject should be asked to refrain from speaking, chewing, or reading during the uptake period (SNM GHS).  For brain imaging the subject should be in a room that is dimly lit and quiet for FDG administration and subsequent uptake period (ACRIN).

- The subject may use the rest room, but if possible not for the 30 minutes immediately after injection of FDG.  The subject should void immediately (5 – 10 minutes) prior to the FDG-PET/CT image acquisition phase of the examination (SNM GHS).

- Bladder catheterization is not routinely necessary; but if necessary the catheter should be placed prior to injection of FDG.  Bladder catheterization may be important for the evaluation of pelvic tumors (e.g., cervix or prostate cancer).

- Following the administration of FDG, the subject should drink 500 ml (or 8 – 12 oz 237-354 ml per ACRIN) of water (or receive by intravenous administration 250 - 500 ml of non-glucose containing fluid).  Fluid intake may need to be modified for those subjects on fluid restriction.

- For specific areas of anatomic interest (e.g., tumors located in the lower abdomen, pelvis or kidney) intravenous diuretic agents may be used (e.g., 20 – 40 mg of furosemide given nearly contemporaneously (within 10 – 15 minutes) with the administration of FDG).  Per the SNM harmonization summit if bladder catheterization is performed IV diuretics should be administered as described herein so as to ensure that the concentration of activity in the renal collecting systems and bladder is relatively dilute.

- Sedation is not routinely required, but is not contraindicated provided that the sedative used does not interfere with the uptake of FDG.  If sedation might be used, the subject was be instructed in advance that operation of a motorized vehicle will be prohibited after the FDG-PET/CT test.  Sedation may have utility in specific clinical circumstances such as brain or head and neck tumors, claustrophobic subjects, or children.

- The amount of fluid intake and use of all medications (e.g., diuretic, sedative) must be documented on the appropriate case report form.

- Subjects undergoing a CT scan should empty their pockets and remove any clothing containing metal and any metallic jewelry from the body parts to be scanned, changing into a hospital gown if necessary (ACRIN SOP).

- Subjects undergoing a CT scan should be asked about the presence of implanted electronic devices (e.g. pacemakers, neural stimulators, cochlear implants) (ACRIN SOP).

- Imaging-related Substance Preparation and Administration

(e.g. Contrast agents, radiopharmaceuticals or stress agents intended to directly affect the imaging process.  Does not include therapeutic drugs/agents unless they are also imaging agents.)

IV and oral iodinated contrast is not discussed as part of this document as its utility is related to the diagnostic CT examination.

Per the SNM harmonization summit, the FDG must meet USP specifications or meet other current specifications as defined by the FDA or other appropriate regulatory agency.

  - Substance Description and Purpose

FDG is a glucose analogue.  Its use in oncology is based on the fact that most types of tumours need more glucose than most other types of normal tissue (Neth).

5.2.	Dose Calculation and/or Schedule

Consolidated Statement - The 18F - FDG dose is usually around 5mCi in Europe and between 10mCi (=370 MBq) (Hallet) and 20 mCi (=740 MBq) (ACRIN) in the United States.  Further FDG dose refinement and/or dose reduction can be achieved by taking into account: (1) patient weight, for example by applying a dose of 5 – 8 MBq/kg (NL, EANM, Hallet, NCI, ACRIN); (2) 2D versus 3D scanning mode (NL, EANM, Hallet, NCI); (3) acquisition time per bed position (NL, EANM, HALLET) and; (4) percentage bed overlap of subsequent bed positions (NL, EANM). The exact dose and the time at which dose is calibrated should be recorded (all, SNM GHS). Residual dose remaining in the tubing, syringe or automated administration system or any dose spilled during injection should be recorded (all).

- In the case of using a syringe for administration, residual activity within the syringe and administration system must be measured and the residual dose should be documented (EU, NL).

- In the case of using an automated system and procedures must assure a net administered FDG activity within 3% accuracy (this must be ensured by manufacturer and verified by the user); i.e., the actual administered activity may not deviate more than 3% from that indicated by the reading of that device or used dose calibrator following instructions given by the manufacturer of the automated administration system (EU).

- Residual activity as determined by the above methods should be used to correct the administered dose for any quantitative results reported.

Any upper dose limits recommended by the manufacturer to take dead time/count rate limitations into account should be considered (EANM, ACRIN). Moreover, (upper) dose limits may apply because of national or local legislation (EANM). In case upper dose limits apply, consistent image quality across sites should be accomplished by increasing scanning time (time per bed, NL, EANM, Hallet). For pediatric studies, other (existing) guidelines may apply, such as the EANM pediatric dose card (EANM).

  - Timing, Subject Activity Level, and Factors Relevant to Initiation of Image Data Acquisition

Consolidated Statement:

FDG uptake into both tumors and other body tissues is a dynamic process that peaks and plateaus at various time points dependent upon multiple variables.  Therefore, it is extremely important that (1) the time interval between FDG administration and the start of emission scan acquisition is consistent and (2) when repeating a scan on the same subject, it is essential to use the same interval between scans performed across different time points.

The suggested consensus time (from all references) between FDG administration and scan acquisition is 60 minutes based on historical use of this test; assuming this is the target window, an acceptable window is often cited as +/- 5 minutes (55-65 minutes).  Two references (NCI and ACRIN) allow the acceptable window to be +/- 10 minutes (50-70 minutes), which is considered the absolute minimum of acceptability.

However, on the basis of the SNM harmonization summit while the “target” tracer uptake time is 60 minutes, the “acceptable” window is from 55 to 75 minutes so as to ensure that imaging does not begin prematurely so as to allow adequate tumor uptake of FDG and to account for the practicality of work flow which often does not accommodate imaging at exactly 60 minutes after FDG injection.  The exact time of injection must be recorded; the time of injection initiation should be used as the time to be recorded per the SNM GHS.  Ideally, the injection and flush should be completed within one minute with the rate of injection appropriate to the quality of the vein accessed for FDG administration so as to avoid compromising the integrity of the injection vein.

Insert paragraph suggesting that more recent evidence might justify a target interval of greater than 60 minutes for a particular trial.  If a target time greater than 60 minutes is chosen for a specific trial, the imaging protocol should justify the specific time chosen as well as the acceptable window about this target time.  Furthermore, as routine clinical practice might not allow the use of pre-recruitment scan for the study, the protocol should include a plan for repeating the baseline scan if necessary to allow appropriate inter-time-point comparisons.

Multiple references (EU, ACRIN, and NCI) imply that the consistency of  timing (relative to baseline) for a follow-up time point should be prioritized for a case where the baseline scanning ideal/target was not achieved; rather than defaulting to a 55 – 75 minute (SNM harmonization summit) expectation at follow-up if baseline was performed outside this targeted window.  This issue should be specifically addressed for each individual protocol.

When repeating a scan on the same subject, especially in the context of therapy response assessment, it is essential to apply the same time interval with target window of +/- 10 minutes (with an acceptable window of +/- 15 minutes) provided that the scan must not begin prior to 55 minutes after the injection of FDG (SNM GHS).  If a limited or targeted scan is obtained at follow-up after a whole body scan was performed at baseline, one should consider adjusting the timing of the follow up scan to be congruent with the timing for the same anatomic region as achieved during the baseline study.

If, for scientific reasons, an alternate time (between dose administration and scan acquisition) is targeted for a specific protocol, then the rationale for this deviation should be stated; inter-time point consistency must still be followed (NCI, EU, Neth).

  - Administration Route

Consolidated Statement – FDG should be administered intravenously through a large bore (≥21 gauge) indwelling catheter placed anatomically remote (e.g., contralateral extremity to site of disease if at all possible) to any site(s) of suspected pathology, preferably in an antecubital vein (SNM GHS).  Intravenous ports should not be used, unless no other venous access is available.  If a port is used addition flush volume should be used (SNM GHS).  As reproducible and correct administration of FDG is required for quantitation purposes extravasation or paravenous administration should be avoided (EU, Neth, NCI, ACRIN).  If an infiltration is suspected, the fact should be recorded and if the study is quantitative, i.e. SUVs will be measured, then the infiltration site should be imaged and the approximate amount of infiltration should be calculated  If the infiltration is greater than 5% of the administered dose and the quantitative result from the FDG-PET/CT study is a primary or secondary endpoint, the data point might be censored from review or the subject might not be included in study (SNM GHS).  The injection site should be documented on the appropriate case report form (ACRIN).

Presuming that the IV access site is properly functioning, the same route of administration may be used for iodinated contrast as is used for FDG.

  - Rate, Delay and Related Parameters / Apparatus

Consolidated Statement – Either manual or automated injection systems may be used to administer the FDG.

- In the case of manual administration, a three-way valve system should be attached to the previously placed intravenous cannula (See Section 5.4) so as to allow at least a 10 cc normal (0.9% NaCl) saline flush following FDG injection.  Residual activity within the syringe, and as much of the administration system as is available (inclusive of the needle cap) must be measured and the residual dose should be documented; and if the residual dose is >1%, the remaining activity should be measured in a calibrated dose calibrator (See Section 5.2) (EU, NL, ACRIN)

- In the case of an automated administration system, the manufacturer’s instructions should be followed.  However, the automated system and administration procedures must be ensured by the manufacturer and verified by the user to perform within the characteristics specified in Section 5.2) (EU)

  - Required Visualization / Monitoring, if any – NA

  - Quality Control

See 12.2.

- Individual Subject Imaging-related Quality Control

See 12.3.

- Imaging Procedure

When the imaging procedure involves acquisition of multiple series, Section 7 may be repeated (7a, 7b) as necessary to describe each acquisition.

  - Required Characteristics of Resulting Data

Describe the characteristics of the imaging data resulting from acquisition and reconstruction that are relevant to its use in the clinical trial.  This description is generally independent of the vendor, platform, and version of the imaging equipment.

    - Data Content

Describe what the acquired images should contain/cover.

(e.g. Anatomic Coverage, Field of View)

Consolidated Statement

For most Oncology indications, anatomic coverage should include from the skull base (external auditory meatus (EU, ACRIN), skull base (Neth), angle of jaw (NCI)) to the proximal to mid-thigh, considered a ‘whole body’ scan.  Per the SNM GHS the acceptable / target / ideal field of coverage is from the external auditory meatus to mid-thigh.  However, other ranges could be used as appropriate for specific clinical trials.  However, the clinical trial should then provide specific instructions with justification.  Three references (EU, Neth, SNM GHS) specify that scanning direction should be caudiocranial to minimize effect from increasing bladder activity during the scan.  Scanning direction should be protocol specified.  It is critical that for a given subject, scanning direction on baseline scans be duplicated at follow-up time points (NCI, SNM GHS).

Two references (EU, NCI) note that extended ‘whole body’ anatomic coverage may be performed for tumors that show higher probability of metastasis or direct extension to other levels (e.g. head, skull, brain or extremities).

One reference (EU) recognizes an option of performing a two-step protocol for examination of head and neck tumors (head / neck portion and apex of lung through mid thigh portion).

One reference (EU) suggests that limited-area tumor imaging can be considered for follow-up examinations, if the disease is restricted to a defined region.  If this strategy is used, attention to scan timing is critical to provide comparison with earlier timepoint(s) – reference section 5.3.

Three references indicate that the subject should be fully included within the CT and PET FOV in order to avoid attenuation and scatter correction artifacts resulting in non-quantifiable data (NL, EANM, ACRIN)

Among various clinical trials and authoritative sources, there is currently no broadly accepted scanning workflow strategy for obtaining the PET and CT components of a FDG-PET/CT exam.  The following text describes two scanning strategies that can be used in a given trial.  For strategy 1, there is no intent to obtain a diagnostic CT scan at the FDG-PET imaging session while for strategy 2, a diagnostic CT scan is obtained.  There are further considerations which must be followed for each of the two strategies. The workflow chosen for a given protocol (1) should be described in the protocol and (2) should be tailored commensurate to the level of expectation of the obtained data (e.g. qualitative or quantitative SUV analysis).

For both strategies, there are several common issues specific to the CT exam that may have an impact on quantitative FDG-PET output, which need attention and protocol specification.  These include (1) contrast material administration, (2) respiratory motion compensation instructions (reference section7.2.2 for complete respiratory instructions) and (3) CT scanning technique (kVp, mAs and pitch).  An attempt is made below to identify ideal, target and acceptable levels of behavior/procedure for each of these three issues.

Understandably, there are practical issues, some of which are conflicting in regard to combining quantitative FDG-PET and diagnostic CT imaging, such that target or ideal behavior may not be achievable for both the FDG-PET and diagnostic CT aspects of certain imaging protocols in certain clinical trials, at least not in a single acquisition. By understanding these issues, imaging protocols should be written to achieve prioritized image quality to address specific protocol goals. At a minimum, all these issues should be addressed in the protocol, ideally with consistency across all sites and all subjects (both inter subject, and intra- and inter-facility) with the target of consistency across all time points for each given subject. The actual details of imaging for each subject at each time point should always be recorded.

Strategy 1:  For FDG-PET/CT in which the CT is used for attenuation correction and localization only (no diagnostic CT intent):

- CT scout topogram, followed by

- CT for anatomic localization and attenuation correction, followed by

- Emission scan acquisition

For the CT scan (Item 2 above), the following considerations apply:

- Contrast Material

The presence of a positive contrast agent (IV or oral), by affecting the CT attenuation map, can result in a small variability of quantitative SUV evaluation.  If this was the only consideration, then ideal would be to prohibit CT contrast administration.  However, in some clinical situations (dependent upon tumor type, tumor behavior or level of anatomic interest), the benefit of oral CT contrast may outweigh the small errors induced in SUV measurement which may include increased SUV variability.  Consequently, ideal and target approaches are grouped as below.  Each protocol should specify the desired approach for the given study.  Most importantly, for each subject, the same approach should be followed for all imaging time points.

  - Ideal / Target

i.	No positive contrast agent (IV or oral) for FDG-PET/CT studies with a predominant intent of quantitation at both baseline and follow-up

ii.	No IV contrast agent; negative or dilute positive oral contrast is allowed for FDG-PET/CT studies with primary quantitative intent with additional need for oral contrast to increase confidence of true positive disease detection and/or additional qualitative assessment.

  - Acceptable

No IV contrast; dilute positive oral contrast is acceptable.

- Respiratory Motion Compensation

Respiratory motion causes SUV errors by two mechanisms: motion blurring and attenuation correction mismatches between CT transmission map and emission data.

  - Ideal

Verbal instructions to subject for similar shallow breathing during both the PET and CT acquisitions; respiratory gating if called for given protocol specification; possibly with advanced methodologies for respiratory synchronization if offered by manufacturer and appropriate to the study.  Respiratory gating on PET may require several CT attenuation maps for optimal quantitation.

  - Target

Verbal instructions to subject for similar shallow breathing during both the PET and CT acquisitions; respiratory gating if called for a given protocol specification

  - Acceptable

Verbal instruction to the subject for shallow breathing during CT and PET.

- CT Technique

  - Ideal

Use of manufacturer recommended kVp and exposure CT Dose Index (CTDI) or Dose Length Product (DLP) settings for low dose exam in addition to the Target and Acceptable conditions stated below.  CT dose exposure should be appropriately reduced in smaller patients and children.

  - Target

Consistency in use of kVp and low exposure (CTDI, DLP) for all time points for a given subject in addition to the Acceptable conditions stated below.  CT dose exposure should be appropriately reduced in smaller patients and children.

  - Acceptable

Recording of actual kVp and exposure (CTDI, DLP) for each subject at each time point.  CT dose exposure should be appropriately reduced in smaller patients and children.

Regarding CT radiation exposure, rules of “As Low as Reasonably Achievable” (ALARA) should be followed.  For a given protocol, the purpose of performing the CT scan (attenuation correction only or attenuation correction and anatomic localization) should be determined.  The CT technique (mAs, pitch, collimation, kVp, and slice thickness) used should result in as low as reasonably achievable exposure needed to achieve the intended goal of imaging working with the scanner manufacturer (Appendix reference) to achieve this objective.  The technique used for an imaging session should be repeated for that subject for all subsequent time points assuming it was properly performed on the first study.

Strategy 2: For FDG-PET/CT in which a diagnostic CT is performed in conjunction with FDG-PET:  As there may be variability introduced into SUV calculations by the presence of even dilute intravascular iodinated contrast, each clinical trial should choose either the Target / Ideal strategy OR the Acceptable strategy as described below for use at all sites, for all time points, and for all subjects.  Any particular clinical trials should NOT allow some sites to implement one strategy and other sites to implement the alternative.

- Target / Ideal –

Follow Strategy 1 (steps 1-3 above) and then with no or minimal patient motion between the diagnostic CT and the PET/CT perform an additional IV contrast-enhanced diagnostic CT after the emission PET scan acquisition. Ensure that the diagnostic CT acquisition is performed consistently for a given subject across all time points. Note that for this case, use negative or dilute positive oral contrast for the non-attenuation CT scan.

In some instances, such as head and neck cancer, a separate dedicated PET and CT acquisition may be appropriate with the arms in a different position (down) than would be used for the remainder of the whole body study (see also Section 7.2.1 “Subject Positioning”).

- Acceptable –

Perform a contrast enhanced (IV and dilute or negative oral contrast) diagnostic CT before step 1 of Strategy 1, then with no or minimal patient motion between the diagnostic CT and the PET/CT complete steps 1-3 (including a separate tidal-breathing AC / localization CT) of Strategy 1 ensuring that the diagnostic CT acquisition is performed consistently for a given subject across all time points.  The IV contrast would then be in equilibrium phase during the emission scan acquisition and the AC / localization CT scan.  (note – since there are no data as to the magnitude of variance in SUV calculation between the IDEAL / Target strategy and the Acceptable strategy, perhaps QIBA should investigate if the Acceptable strategy is indeed truly acceptable for quantitative FDG-PET/CT in the conduct of a clinical trial.)

- Unacceptable -

Performance of a single diagnostic quality CT study prior to or after the emission scan for all purposes (i.e., anatomic localization, attenuation correction, and diagnostic CT information) is considered unacceptable for clinical trial use.

This potential strategy of performing a diagnostic contrast enhanced CT (and using this scan for attenuation correction and anatomic localization purposes as well) and emission FDG-PET scan was considered; but this is deemed not acceptable.  This decision is supported by two sources that state that the diagnostic CT cannot be used for attenuation correction purposes (EU, Neth) and by the SNM GHS meeting.  The major negatives for this strategy are due to misregistration and incorrect attenuation correction application (especially around the level of the diaphragm) due to differential diaphragmatic position between optimal diagnostic CT (typically full breath hold inspiration) and emission (tidal breathing) FDG-PET scan acquisitions.  This is believed to strongly outweigh the benefit of radiation dose reduction achieved by eliminating the low-dose CT for anatomic localization / attenuation correction map.   A dose reduction can be achieved in cases in which a diagnostic IV contrast CT is required, by limiting the CT with contrast to the most relevant regions of the body, which may be a smaller extent than the area imaged on PET.

    - Data Structure

Describe how the data should be organized/sampled.

(e.g. Spatial Resolution, Collimation Width, Slice Interval, Temporal Resolution, Framing Rates)

Ideal – Reconstructed voxel size (i.e., not achieved through post-processing) should be as small as possible without introducing artifacts and also so as to be consistent across all trial sites; with current technology 2 – 3 mm in all three dimensions is achievable.

Acceptable / Target – The matrix size, slice thickness, and reconstruction zoom will yield a target reconstructed voxel size of 3 – 4 mm in all three dimensions (i.e., not achieved through post-processing), although not necessarily isotropic. – for QC see section 12.1.1.  QIBA and/or SNM Image Reconstruction Committee?????

    - Data Quality

Describe any needed measurements of quality.

(e.g. Image Quality, Noise Levels, Motion Artifacts, Radiation Dose)

Image quality (as defined by SUV calibration, SUV Recovery Coefficient, and SNR) should be such that when applying the same acquisition and reconstruction protocol as used in subject scanning to the protocol specified phantom(s) the output should meet the QC standards as stated in Section 12.1.1.

Treatment response assessment and classification (based on criteria) require several quantitative and qualitative assessments.  For details see Sections 9 and 10.  In summary, however, the analysis and interpretation steps depend on several aspects including, but not limited to, assessment of lesion eligibility, percentage change in activity of specified lesions at each time point relative to baseline, and the appearance of new lesions that meet eligibility criteria.

For the first two aspects (lesion eligibility and measuring percentage change) harmonization of quantitative image quality, e.g., by means of harmonizing recovery coefficients measured in specific dedicated phantoms, will result in more uniform lesion selection and response assessments across institutes. Consequently, harmonizing quantitative performance of PET/CT systems coupled with defining some minimum and/or optimum performance metrics should be a strong consideration in the design of a multicenter trial.

For the assessment of progression related to the appearance of one or more new lesion(s), it is important to set a minimal threshold for image quality with respect to lesion detectability. As such, scanners need to have a minimal image quality performance/lesion detectability/SNR in order to be suitable to be used in trials.  It therefore is conceivable that two different sets of reconstruction algorithms and settings may be necessary to use in the trial; one for lesion detection and the other for lesion quantitation.

Both lesion detectability and quantitation must be carefully considered during study design so as to properly define minimum quality standards to be applied across all sites and scanner platforms (see Section 12.1.1).

  - Imaging Data Acquisition

All QC procedures should be followed and documented prior to the initiation of acquisition (SNM GHS).

For serial scan of the same subject, every attempt should be made to use the same scanner, and the same scanner model throughout the trial (ACRIN).

However, in some cases a different scanner that has been previously qualified and is the same platform as the scanner used at baseline can be used for a subject’s follow-up scan in the instance of equipment malfunction (ACRIN 6678).

Per the SNM GHS, the ideal level of performance is that all serial scans on a subject should be performed on the same scanner with the same software version; acceptable / target performance is that all serial scans on a subject should be performed on equivalent scanners (i.e., the same model) but also with the same software version).  (SNM GHS)

Additionally, all scan acquisitions for a given subject should include identical transmission and emission scan technique and emission scan duration per bed position (ACRIN).  There is no consensus provided on emission scan time range. The number of bed positions and the acquisition time per bed position will be scanner specific.  Typical parameters are 6 bed positions and an acquisition of 2 – 5 min per bed position (ACRIN 6678).  Per the SNM harmonization summit the minimum acceptable time per bed position should be between 2 and 4 minutes for a 3D acquisition with 2D acquisitions typically requiring at least 1.5 - 2x longer depending on the administered FDG dose; although the absolute impact on image quality by scan time per bed position is currently undefined it is dependant on several pertinent factors including, but not limited to, administered dose, body weight and habitus, bed overlap, and specific model / version of the imaging platform used.  In general, increased scan time per bed position will improve the SNR and thus it may be important to increase scan time when quantitative metrics are used towards a primary endpoint.  A more detailed set of preliminary recommendations for time per bed position relative to pertinent factors is provided in Appendix XYZ (or is this Table X in Section 13.2??).

As new technology becomes available, it is important that acquisition parameters are implemented to ensure at least equivalent, if not superior, measurable image quality and output metrics. (SNM GHS)

One reference (NCI) notes that whole body acquisitions can be in either 2- or 3- dimensional mode with attenuation correction, but a consistent method should be chosen for all serial scanning of an individual subject throughout the trial.

Two references (NL, EU) describe a relationship between applied FDG dose, acquisition time per bed position, percentage bed overlap and scanning mode (2D, 3D) in order to harmonize image quality (and avoid bias in quantification). Using this relationship these parameters are directly linked, e.g. a higher FDG dose can be offset by shorter acquisition times per bed position etc.  Please see the equations and Table X in Section 13.2 for examples relating dose, bed positions, scan time per bed position, scanning mode, and bed overlap, etc. (NL, EU).

Acceptable:

While there may be variance based on type of scanner, scanning algorithm, model, and software version, the following guidelines are meant to assist each site in achieving the desired data quality as specified in Sections 5.2, 7.1.3, xxxx, and yyyy.

For a dose of 5 MBq/kg or higher (370 MBq or more for a 75 kg patient) the minimal time per bed position using the manufacturers’ recommended bed overlap specifications.  The time per bed position should be at least 2 mins for 3D systems showing ≥50% bed overlap and at least 4 min for 3D systems showing <50%.  Time per bed position may be modified inversely proportional to alteration in injected dose per body weight within the limits of the scanner performance as determined by the manufacturer or an appropriately qualified independent standard-setting organization or peer-reviewed publication.

For 2D systems these times per bed should be at least 1.5 times longer for the same injected dose based on body weight.  Time per bed position may be modified inversely proportional to alteration in injected dose per body weight within the limits of the scanner performance as determined by the manufacturer or an appropriately qualified independent standard-setting organization or peer-reviewed publication.

In general, increased scan time per bed position will improve the SNR and thus it may be important to increase scan time when quantitative metrics are used towards a primary endpoint.

Target;

Use an anthropomorphic phantom (SNM, EANM) with a uniform area to assess ‘noise’ by means of the coefficient of variation of the pixel values within a specified region of interest (ROI) positioned in the uniform area. Phantom should be filled such that activity concentration in the uniform area is similar (within 10%) to the expected average normal tissue concentration seen in an average weight (70-80 kg) subject in combination with the intended FDG dosage. Phantom should than be scanned with various times per bed and reconstructed using the acceptable reconstruction methods and settings (e.g. minimal and/or harmonized resolution criteria etc etc). Next COV will be derived in each image following the SNM-provided protocol for the SNM phantom or the NEMA NU-2 protocol for EANM phantom (for details of these protocols see Appendix XXXX). The COV within the above-indicated ROI should be plotted as function of time per bed position. The minimal allowable time per bed position is then given by the time per bed that provides a COV of 10% or better.  Afterwards higher FDG doses may be offset by longer scan durations

Ideal: I guess this is yet unknown.

    - Subject Positioning

Consolidated Statement –

During PET-CT, subjects should be positioned in the center of the field of view (FOV), preferably with the subjects’ arms to be positioned over head (to minimize beam hardening and FOV truncation artifacts) (ACRIN, SNM GHS).  Alternatively, the arms can be positioned along the side for head and neck imaging (for two-step procedure – see section 7.1.1) (EU, SNM GHS).  One reference notes that subject may be unable to maintain arms above head for the examination, in which case protocol specific handling needs to be defined.  Arm positioning in a particular subject should be consistent as possible across all time points. (ACRIN)

One reference (EU) notes that if PET-CT data are used to radiation planning, the examination should be carried out in the radiation position using the same dedicated radio-opaque positioning devices as used in the radiotherapy department.  Support devices, under the back and/or the legs, may be used to enable the subject to comfortably maintain his/her position throughout the exam (ACRIN).

    - Instructions to Subject During Acquisition

(e.g., breath hold, etc.)

Consolidated Statement -  The diagnostic CT is usually performed in maximal inspiration breath-hold which could result in image artifacts due to mis-registration of the lung-liver interface between emission and CT images should the diagnostic CT is being used for attenuation correction (i.e., there is only one CT scan performed for both diagnosis and attenuation correction which is not the UPICT recommended method per section 7.1.1).  Therefore, two references (EU, Neth) indicate that during CT-AC, shallow breathing is recommended for multi-slice CT scanners; one reference (EU) notes that for single or two-slice CT scanners, breath hold acquisition should be acquired.  Per the SNM GHS, the CT acquisition for attenuation correction should be done with shallow breathing without regard to the CT technology used (acceptable / target / ideal).

One reference (ACRIN) notes that the subject should be monitored and coached (1) in the breathing protocol, if one is used and (2) to remain motionless throughout the scan

For Acceptable, Target, and Ideal performance levels please see Section 7.1.1,

    - Timing/Triggers

(e.g., relative to administration of imaging agents; inter-time point standardization)

    - Model-Specific Parameters

Appendix G.1 lists acquisition parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1 while also complying with the radiation dosimetry specified in Section 13.

    - Archival Requirements for Primary Source Imaging Data

See 11.3.

  - Imaging Data Reconstruction

Describe the data reconstruction process including any inherent data correction, smoothing, etc.  This may conceivably be performed on a different system than acquisition, and occasionally at a later time, particular for supplementary reconstructions.

Consolidated Statement (n.b., concepts to capture STILL A WIP, NOT COMPLETED)

- PET emission data must be corrected for geometrical response and detector efficiency (normalization), system dead time, random coincidences, scatter and attenuation (EU, Neth, ACRIN).

-  Data acquired in the 3D mode can be reconstructed directly using a 3D reconstruction algorithm or rebinned in 2D data and subsequently be reconstructed with a 2D reconstruction algorithm (EU).

- Iterative reconstruction algorithms are current standard for PET (rather than filtered back projection) (EU).

- (ACRIN, Must) Perform reconstructions with and without attenuation correction (EU, ACRIN).

-Scanners must be properly normalized and calibrated to ensure uniformity and accuracy of SUV measurements within the limits of the spatial resolution (ACRIN)

- Standardization of reconstruction settings is necessary to obtain comparable resolution and SUV recoveries across the same subject and inter-subject across sites < model specific issues to allow multicenter convergence - reference in template to 7.3.1 and Appendix??> (Netherlands, EU)

    - Model-Specific Parameters

Appendix G.2 lists reconstruction parameter values for specific models/versions that can be expected to produce data meeting the requirements of Section 7.1.

    - Archival Requirements for Reconstructed Imaging Data

See 11.4.

    - Quality Control

See 12.4.

- Image Post-processing

Describe subsequent modification of the reconstructed image pixels prior to analysis. This is often performed on a different platform than the acquisition system.

(e.g. spatial registration, spatial re-orientation, re-slicing, feature enhancement, 3D view generation)

  - Input Data to Be Used

Describe required input data and any necessary validation or adjustments which should be performed on it.

(e.g. particular image series or views; raw, processed, both)

Input data can be either Reconstructed Data, or Post-Processed Image Data as defined below.

8.1.1	Definitions

Raw Data: This is an ambiguous term as it can refer to scanner raw data (i.e., sinograms or list-mode) or image raw data. This term should not be used.

Raw Projection Data:  This term refers to the data as acquired by the scanner before reconstruction (i.e., sinograms or list-mode).  When this term is used, the user should specify the exact type of Raw Projection Data.

Reconstructed Image Data: This is the image data exactly as produced by the reconstruction process on the PET or PET/CT scanner. i.e., a stack of DICOM slices/files constituting a PET image volume with no processing other than that occurring during image reconstruction. This is always a stack of DICOM slices/files constituting a PET image volume that can be analyzed on one or more of the following: PET scanner console, PET image display workstation, PACS system, etc.

Post-Processed Image Data: An image that has been transformed in some manner, including but not limited to: smoothing, sharpening, image zoom, rotation/translation, resampling, interpolation, slice averaging, MIP, etc.. This is typically a stack of DICOM slices/files constituting a PET image volume that can still be analyzed on one or more of the following: PET scanner console, PET image display workstation, PACS system, etc.

Secondary Image: This is an ambiguous term as it can refer to either Post-Processed Image Data or a DICOM secondary capture image (akin to a photograph). This term should not be used.  Instead please see Post-Processed Image Data above.

  - Methods to Be Used

Describe how the analysis should be performed.

(e.g. algorithms to be used; where measurements should be taken; definition of key anatomical points or pathology boundaries; related annotations)

After data collection and image reconstruction as detailed in section x.x and y.y, respectively, Reconstructed Image Data (PET images) are generated that meet the image characteristics defined the by the trial (see sections XXXX).

For both visualization/interpretation and quantification, no unintended additional image processing (interpolation, re-binning, reorientation, zooming etc) should be applied to the originally reconstructed PET data.

8.2.1	Definitions

Image Processing: transformations applied to an entire image or a region of an image. These transformations include, but are not limited to: smoothing, resolution recovery, image zoom, rotation/translation, re-sampling, interpolation, slice averaging, de-identification, etc.. The output of this process is itself an image, often intended for visual or quantitative analysis.

8.2.2 Processing affecting quantification

Acceptable:

Image Post-Processing methods and parameters that are used should be recorded and applied to all images in a consistent manner following methods specified in the clinical trial. For example all images might be smoothed to the same overall resolution and/or reconstructed with the same voxel size (or in a defined range of voxel sizes).  Quantitation should be applied consistently across all time points and all subjects within a given site.

The originally reconstructed PET data set should always be preserved. In case processed PET datasets are required, they should be saved as separate secondary datasets.

Target:

No Image Post-Processing is used for quantitation and all analyses are applied to the Reconstructed Image Data.   Post-Processed Data may be used for visualization and to facilitate identifying the ROI / VOI.  However, the underlying Reconstructed Image Data should be used for all quantitative purposes.  The ROI / VOI derived from the Post-Processing should be transferred to the Reconstructed Image Data for quantitation.  Quantitation should be applied consistently across all time points and all subjects within a given site.

Ideal:

No Image Processing is used for quantitation. Instead the analysis software for ROIs and VOIs always applies the analysis to the Reconstructed Image Data, regardless of the appearance of the image on the display station (which may be Post-processed). This is also a component of the QIBA FDG-PET Profile.  NB –  the Ideal level of performance is equivalent to the Target level of performance, but in addition to being applied consistently across all time points and all subjects within a given site the consistency is also across all subjects, all time points, and all sites within a given trial.

8.2.3 Processing affecting visualization

Addition image processing may be performed for specific applications or use cases. For visualization most of the image viewing software or platforms will ‘automatically’ apply some kind of image interpolation (on screen) and image zoom to enhance visual image quality, i.e., almost all viewing and data analysis SW application will perform online image interpolation while displaying PET images on screen. Additional image processing may be applied upon user input, such as zooming, re-binning, reorientation, adjustment of slice thickness or summing of slices and image filtering.  When automatic interpolation is applied, it would be desirable that the user has accessibility to replicated zoomed image data at its original matrix size.

Acceptable/target/ideal:

For visual inspection/interpretation of PET/CT data the by the viewing software or platform default online interpolation and zooming may be used. In addition, so-called maximum intensity projections (MIP) may be generated as they may facilitate localization and detection of lesions. Additional processing, such as zooming, re-binning, reorientation and filtering may be applied upon user request only. User should be able to manipulate color scale settings (window/level and color table).. It should always be possible to revert to the default orientation, zoom and binsize (preferably a ‘revert to default’ button is available).

8.2.4 Image de-identification  (See also Section 11.x)

Acceptable:

If images are de-identified to remove PHI, no information that affects quantitation should be removed.

Target/Ideal:

Only the minimal required PHI should be removed; i.e., all information that is not required to be removed should be retained.

  - Required Characteristics of Resulting Data

  - Platform-specific Instructions

Appendix G.3 lists parameter values and/or instructions for specific models/versions that can be expected to produce data meeting the requirements of Section 8.3.

  - Archival Requirements

See 11.5.

  - Quality Control

See 12.5.

- Image Analysis

Describe the generation of new data/information based on the images. This may be performed by an automated program, a semi-automated program, or entirely by a human.

(e.g. organ segmentation, size or volume measurement, flow rate calculation, tissue characterization, observation of the presence/absence/degree of features such as edema)

For quantitation to be most robustly applied, images must meet the image acquisition guidelines as stated within the UPICT Protocol, including, but not limited to, similar tracer uptake times (see Section 5.3), same scanner and reconstruction algorithm (see Sections XX and YY) and similar injected dose (see Section 5.2).  Additionally, the same software and workstation model and version should be used for a given subject across all time points (and for central analysis for all sites and all subjects and all time points) for the analyses described in this section.  Stability and acceptability guidelines have been articulated in the PERCIST 1.0 guidelines (ref).

Image analysis and interpretation also presumes that the image datasets to be used are reconstructed and attenuation corrected as per Sections AA and BB of this UPICT Protocol.

  - Input Data to Be Used and Covariates Necessary for Analysis

Describe required input data and any necessary validation or adjustments which should be performed on it.

(e.g. particular image series or views; raw, processed, both)

Image quantitation is typically performed by determining a Standardized Uptake Value (SUV) in tumor and, ideally, in a reference normal organ.  The SUV measure to be utilized needs to be specified for each protocol and needs to be used consistently at all sites and across all subjects and all time points for all lesion measurements.

    - - The SUV Statistic

Nomenclature relevant to the SUV statistic shall be defined to address the (1) subject relevant versus (2) statistical sampling relevant issues.

      - - Subject indices (bw,lbm,bsa,other)

The subject relevant issue is whether to use body weight (bw),lean body mass (lbm) or body surface area (bsa).

SUL = SUV-lbm = reference to lean body mass

SUV = SUVbw = reference to body weight

SUVbsa = reference to body surface area (rarely used)

From the SNM GHS*, there was consensus that SUV normalized to lean body mass (SUL) is considered ideal (preferably with SUV reported as well) and should be reported.  It was acknowledged that the requirement of SUL may be limiting at this time due to either vendor platform limitation or staff workflow behavior.  Target/acceptable is SUV reporting with inclusion of measurement and reporting of subject height and weight (see separate section 4.2.2) and reporting to allow for other normalizations.

    - Statistical sampling – including report-out values

      - single voxel

      - multiple voxel

Each of the SUV statistics defined above may be measured by one of three statistical sampling methods.  That is the SUL, SUV, and SUVbsa may each be measured using a single voxel measure (max) or multi-voxel measures (mean or peak).  The following discussion (and the remainder of Sections 9 and 10) will use SUV as the generic example.  However, the discussions are generally applicable to SUL and SUVbsa (when appropriate and necessary discussion differentiating among these statistics will be included in various sections of this document).

SUVmax = single voxel (hottest voxel in tumor ROI)

SUVmean=  mean SUV value for ROI with more than one voxel

SUVpeak = subcategory of SUVmean where volume (SUVpeak-3D) or area (SUVpeak-2D) is defined specifically.  The SUVpeak-3D, is obtained from a 1 cc volume sphere (measuring approximately 1.2 cm in diameter) and defines the most metabolically active 1 cc volume in a tumor.  The SUVpeak-2D is the value obtained by measuring the SUVpeak of an area which is 1.2 cm in diameter and which usually subtends only a single slice, but which might also be defined on multiple (most usually three) slices (for further discussion on the methods to be used for defining the 3D volume and the 2D area please see Section 9.2).  There are several approaches for determining the region to be used for the SUVpeak metric.  One involves locating the SUVmax and then centering the SUVpeak VOI/ROI on the SUVmax pixel.  Another involves moving the VOI/ROI throughout the tumor and measuring multiple SUVpeaks (one for each VOI/ROI) until the highest intratumoral SUVpeak measurement  is located.  An automated search mechanism to find the hottest SUVpeak has been developed as a computer code in some systems. It is often, though not always, the case that SUVpeak is centered on the SUVmax pixel in a tumor.

All references indicate that SUVmax (maximum voxel value or hottest voxel) is required for each lesion that is reported as specified in the study protocol and/or considered clinically relevant.

Multiple references (NCI, ACRIN, EU) also indicate that SUVmean of the VOI/ROI obtained be reported.  The SUVpeak equals the SUVmean only when the VOI is a sphere with a specified diameter which is also indicated as a reportable statistic (EU, ACRIN) and the SUVpeak is the most intense region of the tumor.  PERCIST recommends the use of SULpeak. (PERCIST article, Wahl).

The optimal method of assessing tumor response may vary depending on the tumor type and is not yet fully resolved. Furthermore, the underlying tasks of choosing and prioritizing the optimal statistical metric to use and the optimal methodology to define lesion VOI/ROI (section 9.2.a) is challenging given the lack of rigorous comparative studies to date on which to rely.  It is clear that the differing metrics are strongly correlated with one another.  Methods with a single voxel are statistically more variable than those with slightly larger numbers of voxels included; meaning that changes in single voxel SUV measure (i.e., SUL, SUV, SUVbsa) between studies may have to be larger to be statistically different.  Intuitively, the most accurate representation of a lesions cellular tumor burden should include a combination of tumor burden volume and the metabolic activity of that burden (e.g. Total Lesion Glycolysis - TLG) see section XX.  For very small tumors, the SUVpeak values may include some tissue which is non-tumor, lowering apparent tumor activity (see Section YY).

Note that by combining strategies i (Subject indices) and ii (Statistical Sampling) above, using the PERCIST example of SULpeak, this is an SUV measurement using lbm as subject distribution ‘unit’ and mean value of specific size (1.2cm diameter sphere) VOI/ROI as statistical sampling method.  Furthermore, SUVpeak can be provided which uses bw as subject distribution “unit” and mean value of specific size VOI/ROI as statistical sampling method.

IDEAL*,# – In addition to recording the Target metrics, the SULpeak (SULpeak-3D more desirable than SULpeak-2D) and SULmax- both in hottest region of each particular target tumor should be captured;

TARGET*,# - SUVpeak or SUVbsapeak and SUVmax (must specify and should be the same across all subjects and time points)

ACCEPTABLE# - SUVmax or SULmax - single voxel (must specify and should be the same across all subjects and time points); ideally size of single voxel should be known and recorded (although not fully resolved.

*For discussion of how partial or fractional pixel / voxel data could and should be managed, see Section 9.2.2.

#EXPLORATORY –it is recommended to supplement Ideal, Target, and Acceptable performance with an exploratory measures of Total Lesion Glycolytic (TLG) activity and Metabolic Tumor Volume (MTV), see Section XX for discussion of procedures to determine these exploratory measures), suggested, but not mandated, for all three performance levels.

n.b. for discussion on prescribing the Measurement Object (see Section 9.2.1)

    - Covariate inputs (e.g. glucose uptake time, height, weight, FDG-dose)

Please see Section 4.2.2 on obtaining and recording covariate inputs and Section 10.2.1 regarding glucose correction

  - Methods to Be Used

Describe how the analysis should be performed.

(e.g. algorithms to be used; where measurements should be taken; definition of key anatomical points or pathology boundaries; scoring scales and criteria, related annotations)

9.2.1   Methodology to use

ROI (or VOI) tool to be utilized to define either fixed symmetrical size object or lesion constraint condition and strategy to define edge detection needs to be prescribed.  To follow is a catalogue of potential strategies, but the UPICT Protocol does not stipulate any one as preferred.  However, the trial design should stipulate which of the strategies is to be used uniformly across all subjects and time points during the course of the trial.  These strategies can be summarized as below:

      - Regular (fixed symmetrical size and shape) VOI/ROI (SUVpeak-3D, 2D)

        - Automatic

        - Manual

9.2.1.2 Region (Seed) Grow tool strategy

9.2.1.2.1 Thresholding (absolute vs. relative) and

9.2.1.2.2 Constraint or gradient (or tendril)

9.2.1.3 Constraint definition (“region shrinking strategy”)

9.2.1.3.1 Relative threshold (e.g., 70 % SUVmax, 70% SUVpeak)(rw5)

9.2.1.3.2 Absolute threshold (e.g., SUV liver mean + 2SD, SUV = 2.5)

By way of an example, the threshold for definition of an evaluable lesion for tumor volume articulated by PERCIST is mean liver SUL in a 3 cm. diameter sphere in the right lobe of the liver + 2 SD of liver noise.  This threshold is defined at baseline so that lesions can be "hot enough" to have a measurable decline in F18 activity on subsequent studies with therapy.  For relative threshold as the constraint definition, SNM GHS notes that tumor ROI's reflecting the metabolic volume of the tumors are desirable. For simplicity, volumes based on a 70% threshold of the peak tumor SUV should be produced.  This(ese) are viewed as exploratory reports but recognize the tumor volume may provide data beyond that of the peak or max SUV in a tumor.

        - Edge detection (e.g., gradient tool)

9.2.2 Geometric issues (e.g. handling partial pixel/voxel)

The SNM GHS suggested that appropriate use of partial pixel values to secure a 1.2cm diameter (≈1 cc volume) ROI was appropriate and desirable, since standard pixel sizes would not allow selection of a 1 cc volume precisely in most cases.

Acceptable:  Any regular 3D volume for peak activity measurement (e.g., SUVpeak-3D) VOI would be defined as an isotropic spherical VOI with a diameter of 1.2 cm within ±X%.  Any regular 2D area for peak activity measurement (e.g., SUVpeak-2D) ROI would be defined as a circular ROI on a single axial slice with a diameter of 1.2 cm within ±X%.

Target:  Any regular 3D volume for peak activity measurement (e.g., SUVpeak-3D) VOI would be defined as an isotropic spherical VOI with a diameter of 1.2 cm within ±Y%.  Any regular 2D area for peak activity measurement (e.g., SUVpeak-2D) ROI would be defined as a circular ROI on a single axial slice with a diameter of 1.2 cm within ±Y%.

Ideal:  Any regular 3D volume for peak activity measurement(e.g., SUVpeak-3D) VOI would be defined as an isotropic spherical VOI with a diameter of 1.2 cm within ±1%.  Any regular 2D area for peak activity measurement (e.g., SUVpeak-2D) ROI would be defined as a circular ROI on a single axial slice with a diameter of 1.2 cm within ±1%.

Exploratory:  For irregular VOI (TLG, MTV) no single method is specified as Ideal or Target.  However, Acceptable performance of this Exploratory metric is defined as specifying which method is used and using the same method consistently across all time points for a all subjects and sites, and providing the data as stated in Section 9.1.

  - Required Characteristics of Resulting Data

9.3.1	Tumor assessment – See Sections 9.1, 9.2, and 10.

9.3.2	Internal normalization / Comparator tissue(s)

9.3.2.1  liver - Check noise (see multi-extract 12.3.2)

9.3.2.2  blood pool

The stability of normal tissue SUV in tests performed at differing times in the same patient is considered to be a reasonable and practical indicator of the use of similar techniques of performance of PET. Such stability can suggest it appropriate to use the tumor SUV data for response assessment.  Measurement of the normal liver mean was suggested as easy, ideally using a 3 cm diameter spherical VOI (PERCIST standard) that should be reported at each time point.  An alternate method (PERCIST standard) is use of blood pool activity (especially if the liver SULmean is adversely affected by metastatic disease) (as described separately -reference section10.2.1.1.1).

It is suggested that measurement of these parameters and not adjustment by the parameters is most appropriate.  PERCIST proposed the following as acceptable:  Normal liver SUL must be within 20% (and 0.3 SUL mean units) for baseline and follow-up study to be assessable.  If liver is abnormal, blood pool SUL must be within 20% (and 0.3 SUL mean units) for baseline and follow-up study to be assessable.

While not documented to date, it is at least theoretically possible that a subject’s liver (blood pool) SUL may change during the course of the trial (perhaps as a consequence of disease progression or the therapeutic intervention).  The study protocol should specify how quantitative measurements in subjects with “out of range” liver (blood pool) SUL measurements will be managed.  One potential mechanism would be to analyze the data both including and excluding subjects with “out of range” liver (blood pool) SUL measurements.

Acceptable:  SUL of the liver and/or blood pool should be reported for all subjects and all time points.

Target:  SUL of the liver and/or blood pool should be within 20% or 0.3 SUL units of the comparator (either baseline or immediate previous as dictated by the study protocol) study with study data reported both including and excluding outliers when quantitative FDG-PET/CT is used as a primary or secondary endpoint. If using SUV the range of variability should still be within 20% but the absolute value for the range may be greater than 0.3 SUV units.

Ideal:  Unknown

Exploratory:  The ratio of tumor SULpeak to liver (blood pool) SULmean could be reported as an exploratory metric.

Liver (or blood pool) SULmean and SD are important to report, but not a full substitute for quality control (see Section 9.6).  Liver (or blood pool if liver is replaced with disease) noise per PERCIST is considered a reasonable method to assess noise, although acceptable noise level in PET has not yet been determined.

Acceptable:  SD noted at baseline and all subsequent time points

Target:  based on liver SD; threshold (TBD) for baseline and all subsequent scans as well as inter-time point COV (TBD)

Ideal:  Currently undefined.

  - Platform-specific Instructions

Appendix G.4 lists parameter values and/or instructions for specific models/versions that can be expected to produce data meeting the requirements of Section 9.3.

TO BE AUTHORED

  - Archival Requirements

See 11.6.

  - Quality Control

See 12.6.

9.6.1 Statistical Quality of measurement(s) (e.g. noise)

Quality control of the required inputs (imaging data acquisition and reconstruction and covariates) has been described elsewhere in this document and must be satisfied prior to analysis and interpretation.  Additional QC metrics should include:

9.6.1.1.  Subjective assessment of image quality.  For example, movement or mis-registration can lead to invalid AC, poor quality / unreliable quantitative data.  Some images may be too poor in quality (e.g., inadequate counts per field) to quantify.  All necessary data available to determine if quality is acceptable or not; (e.g., both AC and non-AC images should be generated routinely and must be available).  Specific sources of degradation in quality that should be assessed include, but are not limited to:

      - Artifacts secondary to implants in area of concern

      - Patient motion

      - Extraneous activity (e.g., IV tubing or urine) in field.

      - Extravasation of FDG

The output of this subjective QC assessment must include the judgments to whether the study, despite artifacts, still has utility in analysis (e.g., quantitative, semi-quantitative, and/or qualitative).

9.6.1.2. Objective Assessment

Ideal:  Use of a digital reference object is necessary to assess the performance characteristics (e.g., accuracy, precision, etc.) of the software tool, the user interface, and the “user” during the SUL determination workflow including, but not limited to, the determination of the “hottest” pixel / voxel and the creation of the standardized ROI / VOI.

Acceptable / Target:  Document the workstation and software models and versions used and ensure that for each subject the same workstation and software model and version is used across all time points; should hardware and/or software upgrades occur during the course of the trial, testing should verify the comparability of quantitative metrics used in the trial (with comparability defined by the specifications in the clinical trial documentation) also see Section 12.1.1.  Document that the selected parameters used for analysis were achieved in actual practice.  All workstations and software tools should have gone through validation by the manufacturer with approval by the appropriate regulatory body(ies) or the validation should be publically and transparently available.  The trial should include specific QC tasks to ensure QC of the users with documentation at the time of site qualification and periodically during the trial.

- Image Interpretation

Describe the diagnostic conclusions of interest to be drawn from the images.

(e.g. progression of disease, presence/absence/degree of pathology, viable tumor vs. necrotic)

While Analysis is primarily about computation; Interpretation is primarily about judgment.  Interpretation may be performed at both the lesional / target level and in the aggregate at the subject level (e.g., in an oncology study each index lesion may be measured in longest diameter during the analysis phase, but in this phase a judgment may be made as to whether there is a new “non-index” lesion; the aggregation of the measured lesions with comparison to previous studies coupled with the judgment as to the presence or absence of a new lesion will result in the RECIST classification at the subject level).

  - Input Data to Be Used

Describe required input data and any necessary validation or adjustments which should be performed on it. May also specify data which should not be used until after the clinical trial interpretation is recorded.

(e.g. particular image series or views; before and after processing versions of images to evaluate/validate the effects of processing; analysis results)

Input Data- reference section 9.3

  - Methods to Be Used

Describe how the interpretation should be performed.

(e.g. definition of key anatomical points or pathology boundaries; scoring scales and criteria such as BIRADS, interpretation schema such as RECIST, related annotations)

The points listed serve to take the input data and then:

(a) discriminate - qualify as lesion? - if so, then target or non-target

(b) compare - to baseline

(c) derive- use combination of target / non-target / presence/absence of new disease to

(d) describe, stratify, and potentially classify or categorize into discrete classifications- into response assessment category (responder, stable, progressive disease)to obtain Output data (which could also include SUL data of each lesion) from which an Interpretation (Section 10.3- Required Characteristics of Resulting Data) can be rendered (with incorporation of QC check).There are overlap issues (to Baseline and On-study time points), but there are also time-point specific issues which discriminate Baseline from On-study.

Section 10.2.1 -Baseline Time Point Evaluation

10.2.1.1 Qualification of Target Lesions

While PERCIST using one lesion indicates the "hottest": lesion will be taken, if this lesion cannot be reliably be measured on PET due to, for example, artifacts from nearby intense F18 containing structures (like the bladder), then an alternative nearly as hot lesion can be quantified. Similarly, if the 'hottest" lesion is in a region where the quality of quantitation is suspect perhaps due to motion or attenuation artifacts (e.g. at the diaphragm/liver interface, or in the neck under the circumstance that the head has moved) then (an) alternative lesion(s) can be chosen, ideally nearly as intense in activity.  The less easily measurable lesion would be a non-target lesion and would still be assessed for disappearance in the case of possible PR or clear increase in activity in the case of PD.   While PERCIST does not require a lesion to be measurable by CT or anatomic measures when choosing (a) target lesion(s), if two lesions are of similar FDG avidity (i.e., within 10-15% of one another), then the lesion which is more easily measurable anatomically might be preferable for analysis.  Details are enumerated below.

10.2.1.1.1    Minimum metabolic threshold

From the SNM GHS*, there was consensus that a lesion had to be sufficiently  “hot” to quantify, so that a reduction in tumor metabolism could be detected above background levels.

The ideal is taken from the PERCIST 1.0 recommendation.  The minimum threshold metabolic activity of a qualified target lesion on a baseline scan is 1.5 x liver mean SUL + 2 x SD of liver noise.  The measurement of normal liver activity and thereby liver noise  (which is used as a surrogate for the noise level in the study) is made using a 3-cm diameter spherical ROI placed in the right lobe of the liver midway between the dome and inferior liver margin and equidistant between the portahepatis and lateral liver margin.  If the liver is not in the field of view or is abnormal to a degree that normal liver cannot be assessed, then the alternate comparator is to use a minimum threshold level of 2 times SUL mean of blood pool in a 3D object defined as a 1-cm diameter ROI in descending thoracic aorta extended over 2-cms tracking the long axis of the aorta; or by making this measurement in multiple 2D 1-cm diameter ROIs extending sequentially over 2-cm of the descending aorta. If the descending aorta is not evaluable a VOI of the same volume should be measured from elsewhere in the thoracic aorta.

The target level is to use the minimum threshold level of 1.7 times liver mean SUL.  If the liver is not evaluable, the minimum threshold level is 2 times the blood pool as measured above.  This was suggested for simplicity especially as not all software packages may currently offer measurements of variance of counts in a region.

Consensus recommendations:

Ideal:  Minimum threshold for baseline quantitative assessment of a tumor focus is 1.5 x liver* mean SUL  (3 cm diameter spherical ROI in the right lobe of normal  liver as specified in PERCIST) + 2SD of variability in this liver ROI in the right lobe of the liver activity within the PERCIST defined baseline ROI** (liver preferred, 2x blood mean SUL + 2SD of variability in this blood ROI if necessary per modified PERCIST*)

Acceptable / Target:  1.7x liver* mean SUL based on the measurement within the PERCIST defined baseline ROI** (liver preferred, 2x blood ROI if necessary per modified PERCIST*)

*for liver severely affected by lesions, blood ROI mean values may be used as surrogate

**PERCIST defined liver and blood pool ROI/VOI defined above.

        - Influence of anatomic measurability of lesion size; including reportablity of lesion anatomic size

In PERCIST 1.0, lesions selected as target lesions on the basis of meeting minimum metabolic activity thresholds as defined above (Section 10.2.1.1.1) need not meet minimum size requirements; although if multiple lesions with similar FDG activity are present, the hottest anatomically measurable lesion(s) are preferable to FDG-avid lesion(s) that are not anatomically measurable. This may be more valid for lesions that are markedly FDG-avid than for lesions that show relatively low-level FDG activity. Therefore by extension for lesions that have less FDG avidity, it may be reasonable to include a minimum lesion size threshold (or guideline) in addition to other minimum criteria for target lesion qualification.  This is especially important for small lesions in anatomic areas subject to artifact from motion (e.g., lung base or hepatic dome) or for lesions difficult to separate from contiguous normal tissues showing metabolic activity (e.g. urinary bladder).  The SNM GHS* suggests that tumors should typically be over 2 cm in diameter for target lesion inclusion at baseline, although a lesion meeting the appropriate FDG activity metrics need not meet this anatomic measurement threshold as a mandatory minimum. Practically, evaluation of lesion size (e.g., longest diameter) may be difficult especially if no dedicated CT was performed either in conjunction with or within an allowable temporal association with the FDG-PET scan.  This may be due to intrinsic lesion characteristics (e.g., infiltrative or CT lesion isodensity to surrounding tissue) or due to the anatomic location of tumor (e.g., bone marrow site).  For lesions subject to partial volume effect of SUV measurement, notably due to anatomic location (e.g., peri-diaghragmatic lesions at either lung base or hepatic dome), a minimum size requirement may also be reasonable.

If multiple candidate target lesions of similar FDG intensity are present, then the chosen target (or targets depending upon response assessment paradigm being used) should be the larger of the lesion(s) also taking into account the reproducibility of lesion measurement based on subjective factors described below (Section 10.2.1.1.3).

These issues should be addressed prospectively in the clinical trial protocol and protocol-specific guidelines should document whether or not minimum size criteria for target lesion qualifications are used and if so how such size criteria will be used.

        - Subjective assessment on reproducibility of measurement (e.g., contiguous structures, conglomerate lesions, hypometabolic lesions, fluid collections, etc.)

Given multiple lesions that qualify on the basis of threshold activity and minimum size, priority should be given to those lesions that are measurable in an accurate and reproducible way.  Therefore, lesions with a problematic anatomic location or configuration might not be chosen for measurement if there are other lesions that may be measured with more accuracy and reproducibility.  If a lesion is not chosen at baseline secondary to difficulty in accurate measurement, but on subsequent scans the lesion is assessed as dominant or progressive then hindsight review may be appropriate.  The analysis and interpretation should explain the interscan discrepancy (see section 10.3) and such a lesion may have to be assessed as a “non-target” lesion.

      - Use of Non-target lesions

Non-target lesions can be considered as disease that is quantifiable but exceeds the maximum target lesion count or disease that is assessable qualitatively but does not meet threshold requirements for target disease. The presence of non-target lesions should be noted; this can be done either by noting the presence/absence of non-target disease or by identifying sites of non-target disease by organ or anatomic location (e.g., liver or abdominal nodes).  Non-target disease should be qualitatively evaluated at each time point.  Furthermore, changes in the status of the non-target lesions may be noted if only in a qualitative manner (see section 10.2.1.3).  However, if a non-target lesion becomes a target lesion on a later scan, hindsight quantitative review may be appropriate.  The analysis and interpretation should explain the interscan discrepancy (see section 10.3).  Note that in PERCIST, non-target lesion(s) can become target if the lesion increases in intensity beyond the original target lesion, such that the previously defined non-target lesion is the hottest lesion on the subsequent scan performed on-study. This would typically be considered disease progression if PERCIST criteria are met.

      - Use of Qualitative lesion assessment

Incorporation of a visual assessment in the analysis and interpretation with documentation in the CRF may have utility especially in certain oncologic conditions (e.g., Cheson criteria in lymphoma).

      - Other Observations and reporting methods

The assessment should include commentary related to false positive and false negative (e.g. disease mimics/variants/QC) activity as not all foci that meet the preceding criteria may be indicative of disease (e.g., infection, inflammation, fracture, post-radiation changes).  Similarly, there may be artifacts that mimic or obscure reportable disease (e.g., metallic orthopaedic and/or dental implants).  The trial case report forms should include a mechanism for ensuring the capture of these data.

      - Covariate & Normalization Strategies

        - What to use and what not to use (e.g., glucose correction)

Glucose normalization (both for SUV and SUL):  not discussed at SNMGHS, but discussion needs to be included in UPICT protocol.  Proposal for discussion:  Acceptable – collect glucose data on everyone shortly before radiotracer is injected Target – use properly specified glucometer and collect glucose data; Ideal – It is not clear yet if corrections for glucose levels enhance the ability of PET to predict treatment response.   It is suggested this can be explored prospectively to help determine if the actual corrections of SUL are appropriate / necessary / possible.  It is possible the "corrections" may add additional errors to assessments so it is not viewed as appropriate to routinely apply "corrections" in this setting.

Correction for the timing of image acquisition relative to the time of FDG injection outside the prescribed window has been suggested by some references.  However, this is not universally accepted and considered exploratory at this time.

Section 10.2.2 - On-study Evaluation

10.2.2.1 Strategy dependent upon the analysis and interpretation paradigm

The workflow for the analysis and interpretation of the non-baseline imaging examinations (i.e., “on-study” evaluations) is based on the response assessment paradigm that has been chosen for the specific clinical trial; and therefore the baseline requirements.

For example in the paradigm that depends on using the hottest single lesion, but not necessarily the same lesion from time point to time point, the workflow for the on-study evaluations is based on determining the hottest tumor lesion on each individual study independent of the baseline or any previous studies and performing the analysis and interpretation of the hottest single lesion; thereafter finding the non-target lesions (lesions other than the hottest lesion) and performing the analysis and interpretation on those that are pertinent, if any; and finally performing the summary statistical interpretation on the per subject basis (as opposed to the per lesion basis).

Alternatively, in the paradigm that depends on using the five hottest lesions as defined on the baseline examination (with no more than two per organ and all lesions meeting the defined metabolic threshold), the workflow for the on-study evaluations begins with finding the same lesions that were chosen as the target lesions on the baseline examination and performing the analysis and interpretation on each of them; thereafter finding the non-target lesions from the baseline examination and performing the analysis and interpretation on each of them; and thereafter finding any new lesions that meet the minimum threshold requirements and performing the analysis and interpretation on each of them; and finally performing the summary statistical interpretation on the per subject basis (as opposed to the per lesion basis).

The preceding workflow is contrasted with the workflow in the paradigm that depends on using the five hottest lesions as defined on each examination independently from one another (with no more than two per organ and all lesions meeting the defined minimum threshold), the workflow for the on-study evaluations begins with defining the five hottest lesions as previously defined without regard to the lesions chosen at baseline or any preceding studies and performing the analysis and interpretation of those five lesions; thereafter finding any pertinent non-target lesions (lesions other than the five hottest lesions) and performing the analysis and interpretation on those that are pertinent, if any; and finally performing the summary statistical interpretation on the per subject basis (as opposed to the per lesion basis).

The details for response assessment within each of these paradigms are specified in the subsequent Section 10.3.  The definition of “the target lesion” should be based on the preceding criteria that include SUV measurement, reproducibility, measurability, motion, etc.  The use of the response assessment paradigms is categorized by performance level as:

Acceptable -	Option 1:  Single target lesion (generally the hottest single lesion but potentially a different lesion from time point to time point provided that the lesions were both present on both studies – i.e., not a new lesion on the subsequent study(ies)).

Option 2:  Single target lesion at baseline followed over all subsequent studies (i.e., generally the hottest single lesion but defined as the same lesion from time point to time point).

Whichever option is chosen as the primary metric for the specific clinical trial, it is strongly suggested that data derived by both methods would be archived to allow post-hoc analysis of the clinical trial data.

Target - 	Option 1:  In addition to the acceptable performance, sum of the hottest five target lesions with no more than two per organ (potentially different lesions from time point to time point) with all lesions meeting the PERCIST minimum threshold.

Option 2:  Hottest five target lesions at baseline followed over all subsequent studies (i.e., defined as the same lesions from time point to time point).  This option may have utility when lesion selection is performed in the context of RECIST 1.1 anatomic response assessment criteria.

Whichever option is chosen as the primary metric for the specific clinical trial, it is strongly suggested that data derived by both methods would be archived to allow post-hoc analysis of the clinical trial data.

Ideal (exploratory) - 	In addition to the acceptable and target (either Option 1 or Option 2) level of performance one would also determine the TLG activity across lesions included in the paradigm’s dataset meeting the PERCIST minimum threshold (either only the five target lesions or all lesions, to be specified in the protocol).  The use of TLG activity has not yet been validated across multiple tumor types in a multi-institutional setting.  Hence, while this level of performance may be categorized as ideal, it is at this point in time exploratory in nature.

There may be alternative trial designs for specific clinical trial endpoints (e.g., targeting specific lesions based on local-regional therapies or correlation with biopsy).

      - Definition and Management of “New Lesions”

A new lesion is defined as either 1) an anatomic area that had no evidence of disease at baseline by FDG activity but with FDG activity on the follow up study AND a confirmatory anatomic lesion that is not related to a false positive cause (e.g., infection, treatment effect) or 2) an anatomic area that had no evidence of disease at baseline by FDG activity but with FDG activity on follow up study BUT WITHOUT a confirmatory anatomic lesion that is not related to a false positive cause (e.g., infection, treatment effect) that is confirmed as persistent on at one-month follow up (by FDG and/or CT and/or biopsy).  In the case of the latter definition, the dating of the new lesion should be the time of first appearance that met the previously defined minimum FDG-activity threshold.  Some tumors might be anatomically new lesions without FDG activity.  Non-FDG avid lesions should be assessed by RECIST 1.1 criteria.  For non-target lesions please see Section 10.2.1.2.

  - Required Characteristics of Resulting Data – Summary Output Data (Response Assessment)

Objective response

Description of response should preserve the intrinsically continuous and quantitative nature of PET SUV.  Determination if a response has occurred at all (i.e., if the quantitative alteration is greater than expected due to intrinsic biological variability and measurement error) is critical.  It may also be convenient to further classify or categorize response  (e.g., CMR, PMR, SMD, PD). Quantitative response metrics should be determined with consideration of multiple factors including, but not limited to, the purpose of the trial, the precise timing of the PET/CT scans within the imaging schedule (including the allowable window around each time point), the tumor type, the treatment paradigm employed, and the type(s) of decision(s) that will be based on the response assessment. There are a number of proposed schemas (EORTC, PERCIST) available to guide the categorization of quantitative response metrics (as derived by methods described previously in Section 10 of this document), which are otherwise a continuous variable.

Should the proposed schema include confirmatory imaging studies, the type and timing of such confirmatory imaging should be specified in the protocol.

The proposed response assessment schema references two comparator imaging timepoint scans: baseline scan and “best response” scan.  The baseline scan timepoint is defined as the scan timepoint performed prior to initiation of the focused intervention under investigation.  The best response scan timepoint is defined as the scan timepoint at which the lowest level of disease (or maximal response to the therapeutic intervention) is identified.  The best response timepoint may be the same as the baseline timepoint if there is no interval (on-study) timepoint that shows improvement.

For assessment of a responder (CMR or PMR), comparison is made to the baseline timepoint.  For assessment of progression (PMD), comparison can be made to either the baseline timepoint or the nadir timepoint.  See section below on PMD for further discussion. If the nadir timepoint is used as the comparator for PMD and time to progression is being evaluated as a reportable value, then time zero should be defined as the time of the baseline timepoint.  This calculation would then capture the time interval between initiation of focused intervention and time of progression.

One potential categorization schema is presented for consideration in this document (PERCIST).  This schema also does capture the essence of the EORTC criteria.

Objective response reporting should be provided based on the following performance thresholds:

Acceptable:

The categorization schema used for a particular clinical trial should be clearly outlined in the clinical trial protocol prior to activation and data analysis.  The rationale for the categorization schema used should be provided in the clinical trial design (which may be accomplished by reference to a societal standard or a publication in the peer-reviewed literature).  Whichever categorization schema is used, the continuous un-categorized quantitative data as derived by methods described previously in Section 10 of this document should be retained and made available for post hoc analysis.  Furthermore in cases of disease progression and/or response, data should be retained and made available regarding the quantitative and qualitative behavior of target, non-target, and new lesions including both PET and concomitant / follow-up CT-derived information.

Target and Ideal:  TBD.

The example categorization schema follows.

PMD (Progressive Metabolic Disease):

In a clinical trial that includes only a pre-intervention scan and a post-intervention scan……

In a clinical trial that includes multiple post-intervention scans (perhaps in trials with longer term follow up post completion of therapy)……..but too early to dictate whether the baseline or best response scan should be used as the comparator.

Based on PERCIST criteria, assessment for PMD is based on comparison to the baseline time point.  However, the PERCIST criteria (see table 7 footnote) allow determination of progression from the best response time point.  It is acknowledged that progression from the baseline is a very conservative approach which may undercall the date of PMD.  If the best response timepoint is prospectively defined as  the comparator for PMD assessment in a protocol, then it is strongly suggested that a confirmatory followup timepoint be performed at least when progression is defined ONLY in terms of a rise in SUV (and not new lesions).

Progressive disease can be assigned based on progression of target lesions, identification of one or more new lesions or unequivocal progression of non-target lesions as further defined:

1) Target Lesion Assessment:  At least a 30% increase in 18F-FDG SUL peak, with 0.8 SUL unit

increase in tumor SUL peak . If more than one target lesion option is chosen, then this increase will be calculated as sum change of all qualifying target lesions identified, not based on any one of the target lesions; and/or

2) Non-target Lesion Assessment:  Unequivocal progression of 18F-FDG-avid non-target lesions. There is currently no literature-based threshold defined to qualify the unequivocal requirement.  Intuitively, the level of increase should probably be larger than that required for target lesion PMD to avoid overweighting of non-target assessment in PMD categorization.  If PMD is based on non-target lesion assessment ONLY or primarily, then progression should be verified by confirmatory contemporaneous and/or follow-up imaging (which should be performed within 1 month) and/or biopsy unless PMD also is clearly associated with progressive disease by RECIST1.1; and/or

3) New Lesion Assessment:  One or more new 18F-FDG-avid lesion(s) that are typical of cancer and not related to treatment effect, infection or inflammation; this typification may also require confirmatory studies in some circumstances. (See Section 10.2.2.2).

PMD should be reported to include percentage change in SUV peak, (including, time after treatment, in weeks) and whether new lesion(s) are present/absent and their number (i.e., PMD, +35, 4, new: 5). Because SUL is continuous variable, dividing response criteria into limited number of somewhat arbitrary response categories may result in loss of data. For this reason, PERCIST preserves percentage changes in SUV peak in each reported category. Because rapidity with which scan normalizes may be important (faster appears to be better), PERCIST asks for time from start of treatment as part of reporting. For example, CMR -90, 1, is probably superior to CMR -90, 10, especially if the latter

subject were SMD 020, 1.

If TLG volume is being analyzed as an “ideal” exploratory endpoint, then at least 75% increase in TLG volume is required for PMD and should be reported.

CMR:

1.  Complete resolution of 18F-FDG uptake within measurable target lesion(s) so that the uptake is less than or indistinguishable from blood-pool levels (When liver activity is available for evaluation, this implies that the lesion uptake would be less than mean liver activity).

2.  Disappearance of all other (i.e., non-target lesions) lesions to background blood pool levels.

3.  Percentage change in SUL should be recorded from measurable region, as well as the time in weeks after treatment was begun (i.e., CMR -90, 4).

4.  No new 18F-FDG–avid lesions in pattern typical of cancer.

5.  If progression is noted by RECIST (anatomic measurement), but not by metabolic activity, verify with follow-up.

6.  There may be “faint” activity in certain lesions that is greater than immediate background but that is less than or indistinguishable from blood-pool levels.  The presence of such lesions and the absolute SUV measurement should be noted; however, their presence should not dissuade classification as CMR provided those lesions meet the aforementioned criteria.

PMR:

1)  Reduction of minimum of 30% in target measurable tumor 18F-FDG SUL peak.

2)  Absolute drop in SUL must be at least 0.8 SUL units, as well. Measurement is commonly in same lesion(s) as baseline but can be (an)other lesion(s) if the lesion(s) was previously present and is currently the most active lesion after treatment.  ROI/VOI does not have to be in precisely same area as baseline scan, though typically it is.

3)  No increase equal to or greater than 30% in SUL (must be 0.8 SUL units) or size of target or non-target lesion(s) (i.e., no PD by RECIST 1.1 or IWC) (if PD anatomically, must verify with follow-up). Reduction in extent of tumor 18F-FDG uptake is not requirement for PMR. Percentage change in SUL should be recorded, as well as the time in weeks after treatment was begun (i.e., PMR -40, 3).

4)  No new lesions.

SMD:

- Not CMR, PMR, or PMD.

- SUL peak in metabolic target lesion(s) should be recorded, change in SUL peak of the target since the prior study , as well as the time from start of most recent therapy, in weeks (i.e., SMD -15, 7).

Overall Response (summation of time point determinations):

1. Best response recorded in measurable disease from treatment

start to disease progression or recurrence.

2. Non-PMD in measurable or nonmeasurable-nontarget lesions will

reduce CR in target lesion to overall PMR.

3. Non-PMD in nontarget lesions will not reduce PR

Duration of Response (summation of time point determinations):

1. Overall CMR: from date CMR criteria are first met; to date

recurrent disease is first noted.  Note, CMR by RECIST 1.1 is not required.

2. Overall response: from date CMR or PMR criteria are first met

(whichever status came first); to date recurrent disease is first noted.

3. SMD: from date of treatment start to date PMD is first noted.  If PMD must be confirmed on a follow up scan, PMD would be timed to the date when PMD was FIRST noted, not the date of confirmation.

  - Platform-specific Instructions

Appendix G.5 provides instructions for specific models/versions that can be expected to produce data meeting the requirements of Section 10.3.

  - Reader Training

  - Archival Requirements

See 11.7.

  - Quality Control

See 12.7.

- Archival and Distribution of Data

Describe the required data formats, transmission methods, acceptable media, retention periods, …

(e.g. Is the site required to keep local copies in addition to transmitting to the trial repository?  Must all intermediate data be archived, or just final results? At what point may various data be discarded?)

  - Central Management of Imaging Data

Two sources (EANM, ACRIN) mention use of DICOM formatted data. One source (EANM) indicates that data should be stored in DICOM format Part 10: Media Storage and File Format for Media Interchange.  DICOM format should meet the Conformance Statement written by manufacturer of the PET/CT system (EU).

Acceptable:  Data should be stored and transmitted in compliance with pertinent DICOM standards (which for CD and DVD storage and transmission is DICOM format Part 10:  Media Storage and File Format for Media Interchange).  When data are transmitted using ftp or other Internet-based systems.  The archival and transfer method used must allow transmission of all data necessary for qualitative and quantitative assessments without alteration of the data from the acquisition state.  All data transfer should be secure and HIPAA-compliant.

  - De-identification / Anonymization Schema(s) to Be Used

Two sources (EU, ACRIN) indicate that DICOM image data need to be de-identified/anonymized. The header of the DICOM formatted images may contain information that identifies the patient and these tags should be scrubbed or these tags may be replaced by information about study ID, randomnisation or case IDs as indicated upon by the image core lab.  De-identification must be performed prior to transmittal of the data from the local site to the image core lab. Both sources indicate use of (s)FTP as means of transmittal. One source (EU) indicate storing de-identified DICOM formatted images on media (CD, DVD) and sending it by regular mail.

Acceptable:  Data de-identification / anonymization is performed on a third-party or PACS workstation in a manner that is HIPAA-compliant and compliant with the directions of the clinical trial.  However, all data necessary to perform qualitative and quantitative assessments must remain available and unaltered.  Hence, removal of PHI should not affect the underlying imaging data.  There needs to be a mechanism to perform quality control to ensure that the de-identified  / anonymized imaging data correctly correspond to a specific subject ID.

Target / Ideal:  In addition to the acceptable performance level, data de-identification / anonymization is performed on the image acquisition platform in a manner that is HIPAA-compliant and compliant with the directions of the clinical trial.  There is no admixture of PHI and imaging data within the same DICOM fields.  There should be no PHI in private tags.  There should be no imaging data necessary for qualitative or quantitative assessments in private tags.

  - Primary Source Imaging Data

Acceptable / Target / Ideal:  See section 11.4.

Not mentioned in either of the sources

  - Reconstructed Imaging Data

Two sources (EU, ACRIN) mention use and storage of DICOM formatted data. One source (EU) indicates that data should be stored in DICOM format Part 10: Media Storage and File Format for Media Interchange.  DICOM format should meet the Conformance Statement written by manufacturer of the PET/CT system.

Acceptable:  Archival of reconstructed image data either by DICOM format Part 10-compatible media storage or local PACS / server-based storage by both the sites and the central review entity (if any).  Archival of raw projection data is optional.  All archival should be secure and should include disaster recovery.

Target / Ideal:  In addition to the acceptable level of performance, archival of raw projection data is also mandated in a secure and redundant manner.

  - Post-Processed Image Data

Not mentioned in either of the sources

Acceptable:  If post-processed image data is included in the clinical trial imaging protocol or is used during the analysis and interpretation steps whether specified in the trial protocol or not, such post-processed image data should be archived at the time and by the site at which the post-processing is performed, inclusive of all data that was used in the post-processing.

  - Analysis Results

Not mentioned in either of the sources

Acceptable:  Archival of the analysis is performed at the time and by the site at which the analysis is performed by use of a clinical trial-specific case report form that references the specific slices and lesions and provides all pertinent qualitative and quantitative data as required by the clinical trial protocol.  DICOM secondary image capture may be optionally included for clarification.

Target:  In addition to the acceptable level of performance, archival of the analysis is performed at the time and by the site at which the analysis is performed by use of annotations and/or mark-ups on the reconstructed (or post-processed) image data and saved as a new series so that the original reconstructed (or post-processed) image data are retained without alteration.  Theses annotations and/or mark-ups may be archived either as a “screen save” or DICOM secondary image capture.

Ideal:  As per Target, except the ROI / VOI data are captured as true primary data in DICOM format rather than as a representation of the ROI / VOI data captured as an image.

  - Interpretation Results

Not mentioned in either of the sources

Acceptable:

- Quality Control

  - QC Associated with the Site

    - Quality Control Procedures

Describe required procedures and documentation for routine and periodic QC for the site and various pieces of equipment.

Draft Consolidated Statement: - PET and PET-CT systems (and dose calibrator, scales, glucose monitoring equipment, etc.) should routinely be assessed for quantitative integrity and stability by being tested using various imaging protocols on a standard phantom (NL, EANM, Hallet, ACRIN).  n.b., no consensus from extractions but for consideration:  Currently there is no single “standard phantom” agreed upon by all references or extracted studies.  Phantoms might need to be specified for certain types of cancers or anatomic locations and therefore might vary from trial to trial based on diagnosis, treatment, and/or anatomic location.  Options that might be considered on a per protocol basis including, but are not limited to 1) each site would use the same phantom for the duration of the trial (but the phantoms might not be exactly the same among all sites), 2) all sites would use the same general type of phantom for the duration of the trial, 3) all sites would use phantoms that are precisely specified for the duration of the trial, 4) all sites would share the exact same phantom for the duration of the trial.  For SUV measurements, this assessment should include a comparison against a dose calibrator to ensure accuracy; that is, a comparison of the absolute activity measured, versus the measured injected, should be performed (5% NL, 10% EANM, ACRIN). This comparison is particularly important after software or hardware upgrades (NL, EANM, ACRIN).  In case absolute besides relative (longitudinal, response assessment) quantitation is required, it should be considered to include an image quality and/or contrast recovery QC assessment as part of the routine QC procedures and/or scanner validation process, see Appendix E.  (NL, EANM).

CONSENSUS:

If exceptions to any of the performance standards stated below occur, the site should promptly communicate the issue to the coordinating center / core lab for advice as to how the irregularity should be managed; if possible this communication should occur prior to acquisition of any subject data.

Acceptable:

Calibration:

Auxiliary Equipment -

All auxiliary equipment (e.g., clocks, scales, stadiometer, dose calibrators) are calibrated and/or synchronized and/or periodically monitored and documented as part of an ongoing QC program.

Clocks:  Weekly checks against a reference standard for internal consistency.  Dose calibrator and scanner are synchronized within +/- 60 seconds.

Scales and Stadiometer:  Checked on a regular basis by assigned institutional staff.

Dose calibrator:  All calibration tests are performed per the manufacturer’s directions and as defined by the applicable regional and national regulatory bodies.  The most recent manufacturer-specific F18 gain settings are used during these calibration tests.

Scanner -

Scanner is calibrated with same dose calibrator used to assay patient injections.

The same scanner with the same acquisition/reconstruction protocol, software and settings should be used for each subject study.  Only if the primary scanner is unavailable, a scanner demonstrated as having equivalent output (as predefined by the clinical trial site qualification and QC documentation and supported by accepted international standards) and qualified through the protocol’s site qualification process may be used (ideally the second scanner should be of the same make, model, and software version as the primary scanner).  The same scanner acquisition and reconstruction parameters should be used for QC as are being used for subject image acquisition (except for scan duration which may be extended for QC purposes).

Scanner calibration factors (as defined by each manufacturer specific to each scanner model) should be recorded and monitored. Variances of more than 3-5% are potentially due to mis-calibration and therefore should result in verification of correct calibration and/or recalibration as necessary.

At a minimum annual phantom calibration standard when using protocols for clinical trials, acceptable standards are as below.  The same method should be used by each site for the duration of the trial (not necessary for every site to use the same method).

A) ACRIN / EANM criteria for uniform cylinder

1. overall Mean Bkgd. SUV = 1.0 ± 0.1

Ref Scheurman J Nucl Med 2009; 50:1187–1193

B) ACR phantom criteria

1. Mean Bkgd SUV: 0.85 – 1.15

[C) SNM CTN criteria

1. SUV = 1.0 ± 0.1 as assessed in the standard uniform portion of the standard CTN oncology phantom.

Resolution:

At a minimum annually, each site shall perform and document a qualitative resolution QC test by using the manufacturer’s settings and demonstrating resolution of normal gross anatomic features within clinical images of the brain, heart, and abdomen.

Uniformity:

In addition during the calibration methods outlined above, regional uniformity should be assessed qualitatively (i.e., by visual inspection) to ensure that there are no artifactual variations within or between axial slices.

Noise:

During the calibration, uniformity, and resolution testing, when the site uses the trial-specific acquisition parameters (e.g., time per bed position, dose, etc.) and the manufacturer’s settings for acquisition, filtering, and reconstruction, the images should be assessed qualitatively to be of consistent and acceptable quality for trial-specific purposes.

Target (In addition to the Acceptable metrics above):

Auxiliary Equipment:

Clocks:  Checked weekly against and external reference standard (e.g., NTP or equivalent appropriate standard at the site of acquisition).

Dose calibrator:  QC procedures should use of traceable NIST (or equivalent) F18-calibration source to verify the calibration with deviation <+/-3%.

Scanner (Calibration, resolution recovery / detectability, and uniformity should be assessed at least quarterly and in conjunction with any major service or upgrades that may affect quantitative accuracy):

Syringes are assayed post-injection and information is recorded into image DICOM fields.

Calibration:

Same phantoms, calibration method, and quantitative criteria used at all sites chosen from among the acceptable methods as enumerated above.

Resolution Recovery / Detectability:

Scanner reconstruction protocols are adjusted to provide at least appropriate resolution and detectability properties as defined for the specific trial (i.e., RC) for a standard test object that contains specific “hot spot” objects with a pre-specified target to background ratio while maintaining no worse than a pre-specified noise level (e.g., modified NEMA IQ phantom) Ref Boellaard 2008. Include specifications for the phantom characteristics here).

For the CTN PET Oncology Phantom, all lesions 10mm or greater should be visually detectable.

The ACR criteria for resolution are:

The lower portion of the cylinder contains six sets of acrylic rods arranged in a pie-shaped pattern with the following diameters: 4.8, 6.4, 7.9, 9.5, 11.1, and 12.7 mm.  At this target level, the 9.5, 11.1, and 12.7 mm diameter rods must be visible.

By ACR criteria resolution should be achieved as measured by a 25 mm cylinder is >1.8 and <2.8 3 or by a 16/25 mm cylinder ratio: >.7 Ref ACR PET phantom test guidelines (revised 2/22/10). Using the CTN PET Oncology Phantom resolution is accessed by accurate visual lesion detectability (all lesions >10mm) and lesion SUVs and ratios as described in Appendix ZZZ.

For specifications per the EANM guidelines please see Appendix XXX.

For specifications per the SNM/CTN guidelines please see Appendix XXX.

Uniformity:

By ACRIN/EANM criteria axial slice uniformity does not vary more than 10% from one end of the axial FOV to the other as defined by ACRIN Standard Operative Procedures. By SNM CTN criteria, phantom sections of uniformity do not vary more than 10% from one another.  The ROI employed should conform with the use instructions for the particular phantom employed.  Phantom quantitative measurements with overall mean SUV = 1.0 ± 0.10 should be made with an ROI (approximately 3 cm or greater but not including portions subject to partial volume effects) appropriate to the use instructions for the particular phantom employed.

Images are reconstructed with a voxel size of 3-4 mm all three dimensions, but not necessarily isotropic.

Ideal:

Auxiliary Equipment:

Clocks:  Scanner and dose calibrator clocks are driven by the reference standard and are checked weekly.

Dose calibrator:  The NIST-traceable (or equivalent) F18-simulation source is used to calibrate the dose calibrator.

Scanner:

Standardized phantom which could either be the SNM CTN phantom or a uniform cylinder with overall mean SUV = 1.0 ± 0.05. Phantom quantitative measurements with overall mean SUV of each slice = 1.0 ± 0.05 should be made in a series of two-dimensional single slice ROIs (approximately 3 cm or greater but not including portions subject to partial volume effects) appropriate to the use instructions for the particular phantom employed (see Appendix XXX).

Vendors implement an acquisition and reconstruction protocol on each of their machines that ensures pre-defined image RC characteristics are met.  This implementation has two components.  The first component is that every site in a particular trial and preferably across all trials would use the same calibration methods / phantom as prescribed in the industry and multi-societal standard (either the same methods and phantom or the same methods coupled with a defined set of phantoms that have equivalent performance characteristics – i.e., an industry and multi-societal standard).  The second component is that the vendors would provide an automated image assessment tool to ensure that the acquisition and reconstruction protocols produce the desired results.

Each site shall perform and document the full range of the above QC tests using the automated, standardized methods and phantoms described above to document compliance.  This should be part of site qualification and then should be repeated periodically, at least annually and after any major service and after any scanner recalibration related to software upgrades.

See Boellaard et al. The Netherlands protocol for standardisation and quantification of FDG whole body …. European journal of … (2008)

    - Baseline Metrics Submitted Prior to Subject Accrual

See section 12.1.1.+ appendix E

    - Metrics Submitted Periodically During the Trial

See section 12.1.3. + appendix E

  - QC Associated with Imaging-related Substance Preparation and Administration

  - QC Associated with Individual Subject Imaging

    - Phantom Imaging and/or Calibration

<Do we want to place reference here to Appendix D?>

    - Quality Control of the Subject Image and Image Data

Consolidated Statement – The integrity of DICOM image headers should be reviewed and confirmed for regulatory compliance (HIPAA), protocol compliance, and consistency with source data such as CRFs. In some cases, internal references such as the liver can be used for quality control to confirm acceptable ranges of SUVs (ACRIN 6678).  For details of performance see Appendix D.

  - QC Associated with Image Reconstruction

Consolidated Statement - CT images should be reviewed for potential artifacts such as beam hardening, metal objects, and motion. PET images should be compared to the CT images for proper image registration and potential attenuation correction artifacts. (ACRIN 6678).  For details of performance see Appendix D.

  - QC Associated with Image Processing

  - QC Associated with Image Analysis

  - QC Associated with Interpretation

- Imaging-associated Risks and Risk Management

  - Radiation Dose and Safety Considerations

The radiation dose of the PET/CT study results from the radiation exposure from the injection of FDG and the CT study (EANM, ACRIN, HALLET). One source (EANM) indicates that CT scans can be performed as low dose CT to be used for attenuation correction purposes to minimize radiation dose. Two sources (EANM, HALLET) indicate that radiation dose from the CT scans should be estimated specific to the system and imaging protocol used (EANM) or by means of standard estimates (HALLET). These standard estimates can be utilized within the framework of local regulatory requirements for risk analysis (HALLET), which will also depend on patient populations and life expectancy (HALLET) and particular considerations to reduce radiation exposure should be given for paediatric applications (EANM).  The protocol should contain language estimating the total radiation dose and its comparative radiation risk.  The exact estimates of radiation and risk will be protocol-specific and based on factors such as the number and frequency of studies.

  - Imaging Agent Dose and Safety Considerations

(ACRIN) Approximately 1 person in 1000 may have an allergic reaction from the iodinated contrast drugs. These reactions are temporary and treatable. Allergic reactions may include: mild itching or hives (small bumps on the skin), and shortness of breath and swelling of the throat or other parts of the body. The subject should be instructed to tell the technologist immediately if s/he experience any of these symptoms so s/he can be treated promptly.

(ACRIN) The placement of intravenous catheters has the associated risk of making the patient temporarily uncomfortable and a small bruise may form.  A slight bruise may form where the needle has been in a vessel.  There is a slight risk of infection at the site, but sterile technique reduces this risk nearly completely.  The patient may also experience claustrophobia from the imaging ring apparatus or discomfort from lying on the scanner table for 60-120 minutes

  - Imaging Hardware-specific Safety Considerations
Not mentioned in either source

  - Management and Reporting of Adverse Events Associated with Imaging Agent and Enhancer Administration
Not mentioned in either source

  - Management and Reporting of Adverse Events Associated with Image Data Acquisition

Not mentioned in either source

# Comments

Need consistent attributions throughout document and need to include attributions to SNM Global Harmonization Summit (SNM GHS).  Editorial.

Bullet from R Wahl.  Need reference or retain as editorial.

Italicized text in this bullet from R Wahl.  Need reference or retain as editorial.

Italicized text in this bullet from R. Wahl.  Need reference or  retain as editorial.

Per R. Wahl:  End of therapy assessments can allow a greater variance in timing than intratherapy monitoring studies.  Consolidated statement must contain rationale as to why and what the variances might be for intratherapy and end of therapy PET/CT studies.

Per R. Wahl:  It is probably too long for aggressive tumors to wait 8 weeks. The 21 days suggested before is probably more acceptable. Note also that qualitative PET for staging may be less of an issue than quantitative. For quantitative, the same scanner model, ideally scanner etc. must/should be used. For staging, probably less of an issue.

Per R. Wahl:  typically qualitative PET imaging for staging will require less stringency for repeat than for quantitative.

ICF must include description of FDG-PET/CT and radiation risks.  SNM GHS meeting.

Should we call this section Imaging Exclusion Criteria and should we merge 1.7.1 and 1.7.2 since relative contraindications become absolute when they cannot be remediated?  And furthermore when the imaging endpoint is a trial endpoint should there be a statement that the potential subject should be excluded from the trial?  Consider for UPICT V1.1

PERCIST criteria specify SUV peak at baseline for evaluable lesions must be >1.5 liver mean + 2SD of liver mean (3 cm diameter liver ROI), to be included.

SNM GHS - must have personnel trained in the conduct of the use of PET/CT imaging in clinical trials.  Add Ronald’s comments in rewrite.

SNM mtg:  Sites with PET only systems may be acceptable for small studies, but PET/CT is necessary for any large multi-site clinical trials (n.b., whole body oncologic PET).

QIBA TC as convener???

SNM mtg: All sites need to have an Institutional Review Board (IRB), or equivalent group, that oversees and can approve the conduct of experimental studies of human subject.  Also data management capabilities from SNM GHS document.

Should this statement go here??

Additional education/history emphasis made by SNM mtg:  (some of this is covered in section 4.2.1, 13.1. and  13.2 currently) When subjects are scheduled for the study, the scheduler should be provided with a list of all the necessary questions that need to be asked.  These questions should include: weight and height (they will be checked at time of study); can they lie still for the require time (both uptake and imaging); do they have claustrophobia; are they pregnant; are they diabetic.

The subjects should be given written, as well as oral, instructions on necessary preparation prior to the study.  The patient education should include why these requirements are important, what the scan experience is like, how long it takes, and that the CT may not be the same as a usual CT.  It is important to instruct the subjects that their behavior (medications, diet, exercise, sleep) should be the same on follow-up studies as it is at the time of the initial study

CONSENSUS point – to section 4.1 (Fluid Intake).

I do not recall this discussion point; this is currently not captured in consolidated statement. Needs a home.

GSD move to scheduling section 3.

SNM mtg:  for all, less than 150 mg/dL is target. For non-diabetics, (a) quantitative study  MUST achieve target (b) qualitative study acceptable is less than 200. For diabetics, see separate discussion.

This section is cut/paste from FBG in diabetics and non-diabetics for integration into UPICT protocol and suffix ep08Jan added.

Similar to a previous comment – needs to be incorporated in section 12.1?

QIBA – need to address the issue of DICOM field and standardization among vendors for recording dose information.

Ensure congruity with already extracted material in section 7.2.1.

Note to QIBA TC for text regarding voxel size (matrix, slice thickness, reconstruction zoom) in reconstructed image must be within certain parameters (3 – 4 mm perhaps) without regard to manufacturer or model.  See also Sections 7.1.3 and 12.1.1.

Could this be the “acceptable” level??

For consensus reference 12.1.1

Added per phone discussion on 12/8/2010.

Should the “ideal” be identical dose and identical timing per bed position while “target” should be that the timing per bed position could be corrected on the basis of administered dose based on the equations and tables as cited in the following paragraph; but only within a pre-specified limit / range so as to maintain an acceptable SNR.

Refer to QIBA TC noise related to bed position, dose, smoothing, etc.

Insert reference paper.

Should “ideal” be that this should be the same for all subjects and all sites; “target” should be all subjects at a given site; “acceptable” would remain as written.

This material may be moved from Section 13.2 to an Appendix in the final version, tbd.

Include reference to Section 12.1.1

Link from 5.2

Question of utility for ACR phantom – check with QIBA PET/CT TC.

Ref.

For discussion and consensus through QIBA TC; discuss supplementation with some human data as well.

Needs further review.  ACR Phantom use???

To QIBA TC for input from vendors to provide model and version specific settings for an appendix.

Add text in consensus statement to include use of positioning aids and immobilization / stabilization devices for specific anatomic areas, e.g., H&N.  Include text suggesting the ability to maintain position should be part of pre-exam screening (see Section X).

See section XXXX.

Could vendors take first pass at providing text for Appendix G for QIBA review.  Boilerplate language – SNM and QIBA facilitated activity????

??? Is this covered by the SNM evening meeting work group led by John Sunderland???  Ronald to work with John Sunderland to ensure more recent text is inserted in UPICT>

Needs to include consideration of PET and CT alignment / registration.

To be further discussed during the SNM Reconstruction workshop and potentially QIBA thereafter.

Suggest that the vendors would provide the model specific parameters to achieve the outputs specified in 7.1 and 12.1.1;  for “post-UPICT / QIBA” scanners, there should be a setting that guarantees the specified outputs.  Refer to QIBA TC.

This is important for DICOM, vendors, QIBA to address and standardize to allow compatibility among all workstations for PET datasets.  Therefore refer to QIBA TC.

- ???Is this section also to come from the Sunderland reconstruction group or QIBA???  Tracking of software version (compare sw version of all scans performed in a longitudinal patient study (acceptable), for all scans in a trial (target) )

Acceptable: visually checking dicom headers

Target: automatically checking SW version within a tumor tracking or response assessment analysis tool

NEED to put LINGS QIBA OUTPUT HERE

Check on DailyQC results (once available in dicom header)  This note applies to Sections 7 and 9.

For discussion purposes only:  If one chooses to use postprocessing to harmonize image quality and characteristics (i.e. smooth down everything to the worst performing scanner/site) than that should be here….it was decided that nobody wants this.

Maybe some wording on partial volume corrections could be put here, but maybe not for version 1.0 – see analysis.

GSD to contact Jeff Yap and John Sunderland; GSD to author some strawman summary text from discussion .

Must be explicitly stated in clinical trial design as to whether post-processing (filtering, smoothing, etc.) is to be allowed.  Must be more clearly defined as to difference between reconstruction vs. post-processing.  Free-standing workstation / display station functionality.  Any software or platform used must not introduce any unintended post-processing artifacts including, but not limited to, rebinning, filtering, smoothing, etc.

Need to differentiate post-processing for visual inspections (which needs interpolation) as opposed to quantitation (which should be performed without interpolation).

Raw data should be provided to the analysis and interpretation steps without any post-processing including, but not limited to, xxxxxxxx.  Same workstations to be used for any given subject across all time points.  There needs to be some QC and cross-calibration when changes are made (perhaps using a digital reference object per QIBA, NIST).

Define “raw data.”  Ronald to provide first draft.

Perhaps the eventual Ideal is unknown and should be further defined by QIBA PET/CT TC.

Call out to QIBA PET/CT TC - all manufacturers must agree that they will not co-mingle PHI and data necessary for quantitation within any DICOM fields.

Perhaps reword for Ideal, Target, Acceptable

QIBA TC:  For SUV by lbm, need to consult RB and RW (review the equation in most recent paper by Wahl, et al regarding which equation to use.  Need to achieve consensus to include all relevant formulae or pick one across all platforms.

A report by Chien et al at SNM 2011 demonstrates that the SUL as deployed by most scanner systems tends to overcorrect downward the SUL in very obese patients, especially women (Chien abstract SNM). This is because of a squared weight term in the formula. A more correct alternative is under evaluation.

| Reference #: | 1059718 |
| --- | --- |
| Publication No.: | 535 |
| Abstract Title: | Relationship of age, gender and body morphometry to normal hepatic SUL: Paradoxically low hepatic SUL in morbidly obese women. David Chien, Richard Wahl; Johns Hopkins University, Baltimore, MD |
| Accepted To: | Scientific Paper Session (Oral) |
| Session Info: | Technical Advances & Quantification III Wednesday, June 8 8:00 AM - 9:15 AM Room 217B |
| Reference #: | 1059718 |
| Publication No.: | 535 |
| Abstract Title: | Relationship of age, gender and body morphometry to normal hepatic SUL: Paradoxically low hepatic SUL in morbidly obese women. David Chien, Richard Wahl; Johns Hopkins University, Baltimore, MD |
| Accepted To: | Scientific Paper Session (Oral) |
| Session Info: | Technical Advances & Quantification III Wednesday, June 8 8:00 AM - 9:15 AM Room 217B |

Note this paragraph - in black - is taken as extraction, so should this be altered (e.g. SUV to SUL?) Also, the final words - in purple - are added by rw; not part of original extraction.

PERCIST also suggests that up to the 5 most intense tumor lesions be quantified on an exploratory basis for SUL peak.  PERCIST also suggests that the total glycolytic volume of all one, five or (? all) tumors can be explored as a metric of total tumor burden.  See Interpretation Section.

need to ensure collection/recording of covariates -height and weight (section 9.1.ii.b

Ronald please justify the need for this text.

Rathan to provide definition of TLG and MTV please.

Refer to QIBA TC.  Vendors need to implement standardized tools to accomplish these ROI/VOI creations.

Note to QIBA that most software takes 70% of the MAX SUV, which is biased; if we are recommending 70% SUVpeak the vendors need to be able to produce this metric.    Alt is to do a statistically based definition of tumor ROI, which can include liver + 2SD of liver SUL (as threshold)... PERCIST

Insert reference for this type of tool

Refer to QIBA TC to establish standardized mechanism among vendors.

To QIBA TC:  vendors may have varying approaches, so somewhat of a black box at present. Would suggest that consistent work station application be used.

Section 10.2.1.1.1 includes text on how to measure liver noise.  However, there is no information on the minimum threshold for what level of noise is acceptable.

To QIBA PET/CT TC and QIBA TC coefficient of variance (COV) but not as a result of filtering / smoothing; as well as baseline comparison;  15%, 25%?

To QIBA TC;  What is necessary to reach this Ideal state.  Also how do we QC the User.  The current QC is only focused on the software tool and not the user.

Target:  In addition to the acceptable QC, document adherence to the manufacturer’s QC processes and the clinical trial parameters including, but not limited to the pixel/voxel size, matrix, zoom, ROI/VOI size and geometry were achieved for each analysis.  Deleted this text as it is not applicable to this section of the template.

To QIBA PET/CT TC – how does one determine a SD from multiple 2D ROIs??

To QIBA TC – why can’t all software vendors do this??

While glucose correction has been variably applied, evidence supporting its routine application are absent. Indeed, some data suggest it is not a useful method.

| Reference #: Chien et al. 2011 SNM abst | 1062393 |
| --- | --- |
| Publication No.: | 1771 |
| Abstract Title: | Reproducibility of “glucose corrected” hepatic SUL compared to uncorrected hepatic SUL in baseline and follow-up studies. David Chien, Richard Wahl; Nuclear Medicine, Johns Hopkins University, Baltimore, MD |
| Accepted To: | Scientific Poster Session |
| Poster Section: | Technical Advances & Quantification Posters |

Include pertinent reference(s).  PERCIST, SNM GHS

inserted to address Jeff's point regarding scanning at a timepoint prior to next drug administration primarily to evaluate for 'appropriate' glycolysis rebound; although this might be more appropriately covered in first section of this document.

I inserted this sentence for consideration.

While I agree with this; the statement  may open up box of questions regarding whether the anatomic 'size' for targets/new lesions definitely needs to be reported at each timepoint.

We need to address how SUV as opposed to SUL change should be triggered - also 30%? what about absolute SUV threshold trigger - something greater than 0.8 SUV units?

We should insert a sample formula of how to do this since there are two mathematical ways to get to this SUVmean change number yielding different numbers - I do not know which 'way' is statistically correct.  That is, sum of percent changes divided by number of lesions, or sum the SUV's of all lesions at each of the timepoints and then calculate change and divide by number of lesions.  I used the former.

This text is captured in reword above within PMD section - ?delete???

Captured this concept here; rw, please review - I left this sentence highlighted below.

What does this mean – at what time point; by what means; etc.

We may not want a 30% increase in a non target lesion which has low activity to be PD.  Should discuss this.

note if multiple studies, do we want to use a new baseline, or refer back to original baseline?

It may be worth clearly stating that the Baseline for initial response is the immediate pre Therapy scan. However, for progression, it is suggested that if early detection of progression is desired, then a new baseline scan be established from which progression can be judged. This is the new BASELINE , post Rx, scan.  Progression from this NEW Baseline by SUV increase only should ideally be confirmed by a follow up scan, or by clear tumor progression on RECIST 1.1.

This section could benefit from QIBA PET/CT TC input.

As multi-center clinical trials might use data from various manufacturers, private tags that are vital to image data analysis could be “anonymized” thereby rendering the study quantitatively useless.  QIBA should find some mechanism to overcome this issue.

Information necessary for post-processing, reconstruction, quantification interpretation, and analysis should not be contaminated by PHI to allow de-identification without loss of prerequisite data.

If raw data are of interest for a particular trial, the trial protocol should state explicitly the standards for the format and storage of such data.

Refer to QIBA TC for standardization among vendors.

Refer to 7.1.3 when reviewing this section to ensure QC is appropriate.

From my notes of SNM mtg: One issue which arose was using the same dose calibrator and the dose calibrator that was used to calibrate a specific scanner.  See Consensus statement.

Does this introductory text need to remain???  And if so, should it be edited??

Must cross-check with Section 2.2 , imaging equipment.

QIBA TC or someone else to define.

reference appropriate Appendix.

Reference to dose calibration section of the UPICT protocol – also include in profile document (QIBA TC)

QIBA TC needs to work on language regarding noise level.

QIBA TC / SNM Reconstruction Harmonization working group – Ronald and PC to draft text with input from SNM Working Group and QIBA TC.

Check Section 7.1.x to ensure inclusion.

Suggest this as a goal for the SNM CTN and/or QIBA TC and/or some combination of the two.

Suggest this as a goal for the QIBA TC.

SNM mtg: They need a reliable source of FDG; check for extractable material to this section.

Insert from multi-extract into Appendix D.

Insert from multi-extract into Appendix D.